WAGENINGEN EVALUATING PROGRAMMES FOR ANALYTICAL LABORATORIES

DE17  Microplastics

Report 2022.2

Version no. 1 (21-12-2022)

# Introduction

Dear QUASIMEME-participants,

We are happy to present the final 2022 round 1 reports. We have noticed that many laboratories are still having problems continuing their regular activities due to Covid-19. If your laboratory has not been able to submit the data on time, we can still calculate z-scores for your laboratory based on the NDA mean values published in the report. Please get in contact with wepalquasimeme@wur.nl if you want to receive z-scores like that. Those z-score lists comes with a disclaimer of course.

Since 2020 round 1, the reports have included the data from the homogeneity study where relevant.

The z’-score reports (lab specific reports) are published separately and can be found in the report section on the participant site.

We would like to draw your attention to two new developments within WEPAL - QUASIMEME. WEPAL - QUASIMEME has decided to combine their separate email boxes and from now on they can be reached on one joint email box: wepalquasimeme@wur.nl

WEPAL - QUASIMEME has also worked on the creation of a discussion platform on LinkedIn. If you become a group member, interesting topics like ocean acidification and analytical problems can be discussed with other experts in the field. This is by far the most powerful tool to bridge your knowledge for cross collaboration.

See what WEPAL - QUASIMEME can do for you and become a member of our discussion group. Link WEPAL - QUASIMEME discussion platform: https://www.linkedin.com/groups/12544425/

QUASIMEME is always open to the possibility of introducing new exercises and or determinands.  Please contact the WEPAL - QUASIMEME Project Office with any complaints, requests, comments or feedback you may have at wepalquasimeme@wur.nl

Yours sincerely,

Steven Crum

Coordinator QUASIMEME Programme

Calculated  13-Dec-2022

Approved by Steven Crum, Co-ordinator Quasimeme Programme

Template: version 2021.1

Calculated with Matlab NDA version: QuasiNDAstat_V22_0

Figures version: WepalQuasi_figures_V22_1

# Content

Introduction	2

Content	3

General Information	4

Accreditation	4

Homogeneity of the distributed samples	5

Homogeneity tests	5

Check of results	5

The periodic report	5

Statistics	5

Normal Distribution Approximation (NDA)	5

The NDA-mean (assigned value)	6

The Indicative Values	6

Traceability of the assigned value	7

Uncertainty of the assigned value	7

Median and MAD	8

The z’-score Assessment	8

Evaluation of results	10

Rounding of results	10

Materials Analysed	10

Method Indicating Code (MIC)	10

References and related literature	13

Used abbreviations and symbols	14

Ranked z'-scores  DE17  2022.2	15

Total number of particles Analysis DE17  2022.2	17

Total number of particles Summary Statistics	18

Total number of particles Data and Statistics	21

Total number of particles Combined z'-scores	36

Total number of particles Histogram+PDFs and Ranked overview	60

Total mass of plastic particles Analysis DE17  2022.2	88

Total mass of plastic particles Summary Statistics	89

Total mass of plastic particles Data and Statistics	92

Total mass of plastic particles Combined z'-scores	97

Total mass of plastic particles Histogram+PDFs and Ranked overview	102

# General Information

## Accreditation

QUASIMEME programme is part of WEPAL (Wageningen Evaluating Programmes for Analytical Laboratories). WEPAL is accredited for the organization of Interlaboratory Studies by the Dutch Accreditation Council RvA since April 26, 2000. WEPAL’s registration number is R002. WEPAL has 5 different proficiency testing programmes: IPE, ISE, SETOC, MARSEP and BIMEP, which covers the plant, soil, sediment, manure and biomass samples. QUASIMEME has been added as a separate programme into the WEPAL programme and accreditation was obtained a few years ago for a selection of the QUASIMEME exercises as well. Information about the scope can be found at: http://www.wepal.nl/website/about_wepal/Scope.htm. Accreditation is based on the new standard ISO/IEC 17043:2010

Within these determinand groups only the determinands printed in bold meet the criteria for accreditation. These criteria are based on information about the homogeneity and stability of the samples. This information is available when sufficient participants have reported results for a determinand in the past 3 years. Determinands which are not reported regularly in sufficient numbers to have a statistical evaluation are not (yet) included in the scope of the accreditation.

Scope of the QUASIMEME programmes. 	(Determinands not in bold are not in the scope of the accreditation)

| DE17 | Determinand |
| --- | --- |
| Total number of particles | PC (50-299 µm); PC (300-5000 µm); Polycarbonate total; PE (50-299 µm); PE (300-5000 µm); Polyethylene total (HD+LD); PET (50-299 µm); PET (300-5000 µm); Polyethylene Terephat. Total; PP (50-299 µm); PP (300-5000 µm); Polypropylene total; PS (50-299 µm); PS (300-5000 µm); Polystyrene total; PVC (50-299 µm); PVC (300-5000 µm); Polyvinyl Chloride total; Other polymers (50-299 µm); Other polymers (300-5000µm); Other polymers total; |
| Total mass of plastic particles | PC (50-299 µm); PC (300-5000 µm); Polycarbonate total; PE (50-299 µm); PE (300-5000 µm); Polyethylene total (HD+LD); PET (50-299 µm); PET (300-5000 µm); Polyethylene Terephat. Total; PP (50-299 µm); PP (300-5000 µm); Polypropylene total; PS (50-299 µm); PS (300-5000 µm); Polystyrene total; PVC (50-299 µm); PVC (300-5000 µm); Polyvinyl Chloride total; Other polymers (50-299 µm); Other polymers (300-5000µm); Other polymers total; |

Subcontracting

Some aspects of the proficiency testing scheme are subcontracted, like the preparation of nutrients and DOC in different types of seawater (AQ1; AQ2 and AQ14) and the preparation of all biota samples and shellfish toxin samples. When subcontracting occurs, it is placed with a competent subcontractor. WEPAL-QUASIMEME is responsible to the scheme participants for the subcontractor's work. The analysis for the homogeneity tests of the samples used in this proficiency test are carried out by a subcontractor where applicable, with the exception of AQ11 as the homogeneity of these  samples  were checked by ourselves.

Confidentiality of results

The confidentiality of the results is extremely important in the WEPAL-QUASIMEME programs. The participants may opt for a code name that indicates their laboratory, or one that ensures their anonymity. In the reports, only the code names will be mentioned. When an accrediting body or a regulatory authority requires the proficiency test results to be provided by WEPAL-QUASIMEME the participants shall be notified and asked for permission.

Participants may not use or report individual data from other laboratories. Assigned values, means and standard deviations of the interlaboratory studies published in this report may be used

Complaints and or remarks

The reports of WEPAL-QUASIMEME are assembled with the utmost care. Please contact us on

wepalquasimeme@wur.nl if you feel that the reports are not at a satisfactory standard, if you encountered errors in your results or if you want to appeal against the evaluation of your performance. Also feel free to contact us if you have any other complaints, remarks or suggestions.

# Homogeneity of the distributed samples

## Homogeneity tests

WEPAL-QUASIMEME has developed special equipment for the production of representative subsamples (Houba, 1993) from sediment bulk material. The proper functioning of this equipment is tested by a homogeneity test in the final subsamples. To perform this test, samples are collected at regular intervals during the preparation of the samples. The collected samples, with a minimum of 10, are analysed in duplicate measurements under repeatability conditions. A selection of critical determinands is chosen for the tests. All samples used in this round of the proficiency test have passed the homogeneity test. Samples prepared by our subcontractors are tested for homogeneity in the same way. Finally, the WEPAL-QUASIMEME staff is judging the homogeneity for these subcontracted  samples. The samples used for AQ3, AQ4, AQ5, AQ6, AQ7, AQ8, AQ12 and AQ13 are not tested for its homogeneity. These samples are prepared in bulk, and it is assumed that these samples are homogeneous as the determinands are very good soluble in the matrix, either samples are spiked by the laboratories themselves.

Summary homogeneity test samples

Although the homogeneity determination was carried out by experienced laboratories, the values found for the mean of the parameters in the upper table, used for homogeneity testing, do not have to be the true value or assigned value. Deviations from these values can arise due to e.g. differences in calibration. If this occurs, it normally has no negative effect on the determination of the homogeneity itself. Sa and Ss are calculated with the F-test, while Sigma is equal to the target error (See page10).

## Check of results

Before distribution of the periodic reports to the participants, a final check is made based on the results found by the participants. This check is made for all reported determinands. The variations between laboratories and concentrations are compared with the patterns as found in previous  years and with the indicated ranges in the QUASIMEME brochure. Significant deviations from regular patterns and variations will be investigated. If necessary comments will be added to the report on these issues.

# The periodic report

In order to evaluate the accuracy and precision of the analytical procedures used, several testing programs have been established. At this moment the WEPAL-QUASIMEME comprises approximately 300 laboratories in many countries. The participating laboratories receive samples according to the QUASIMEME brochure and analyse the samples according to their own procedures. The results of the determinations are collected and processed at Wageningen University and published according to the time schedule given in the QUASIMEME brochure. The participating laboratories are informed of the results within one month after the dead-line mentioned in the QUASIMEME brochure. Occasionally, the dead-line can be extended and if so participants will be informed about it. Each participant can compare his results with those of all the other members of the exchange program. WEPAL-QUASIMEME will not comment on results unless asked to do so.

### Statistics

## Normal Distribution Approximation (NDA)

Interlaboratory studies like those of WEPAL-QUASIMEME frequently give rise to datasets that have complex distributions including excessive tailing and multiple modes. Consequently, sophisticated statistical methods are required to obtain meaningful assessments. A methodology is needed that does not rely on arbitrary outlier removal or subjective manual interpretations. The model that is chosen calculates population characteristics (mean and standard deviation) from experimental datasets as described by Cofino et al. (2000) and Molenaar, Cofino and Torfs (2018).

The statistical principles of the model that we use to assess the data are outlined in two steps. Firstly, the full model is described, thereafter a description is given of the way the model is implemented for the assessment of the data in WEPAL and Quasimeme.

We assume that each laboratory i submits a result given by a probability density function qi.  We start thus from a set of probability density function qi. I=1,....,N. We set ourselves to establish the average probably density function q  that best describes the set.

It is insightful to make at this point an analogy with the calculation of the arithmetic mean of a set of data ai, i=1,…,N.   The average  a can be defined as the point that minimises the sum of the squared Euclidean distances da,ai to the given data.  This can be accomplished by equating the first derivative of i=1Nd2a,ai=i=1N(a-ai)2 with respect to a to zero. One readily finds the well known expression a=1Ni=1Nai

In a similar manner, we construct the average probability density function  q  of the set of probability density functions qi, i=1,…,N.  We define a measure d(p,q) for the distance between two probability density functions p and q.  We obtain  q  by minimising the sum of the square distances from each probability density function qi to q,thus by equating the first derivative of i=1nd(q,qi)2 with respect to q  to zero. The calculation itself is extensive and not given here. The mean and standard deviation of the population are calculated using the first and second moments of the probability density function  q . The variance obtained from the second moment comprises both a within-laboratory and between-laboratory component.

In WEPAL and Quasimeme, laboratories report single data, we have no information about the underlying probability function. To cope with this problem we use a specific implementation of the model: the so-called Normal Distribution Approximation (NDA). The NDA approach is parametrised to reproduce the population characteristics of truly normal distributions, and is a robust method to evaluate interlaboratory studies.

The NDA approach has been devised using a set of normal distributions qi=Nμi,σ, i=1,…,N. We assume thus that all normal distributions have the same standard deviation σ.  The expected values μi are also taken to be normally distributed: μi=Nμ,S. It appears that the mean μ and the standard deviation S of the normal distribution describing the population can be exactly reproduced when σ=0.78*S. In the NDA method, the standard deviation S is calculated directly from the total variance, no distinction between within-laboratory and between-laboratory components is made.

In practice, we have N laboratories each reporting a single value. This gives rise to a dataset xi, i=1,…,N.   We calculate the population standard deviation from this dataset using the robust estimate S=1.4826*MAD  (MAD: median of absolute standard deviations). The normal distributions associated with the data xi are estimated by qi=Nxi, 0.78S=Nxi, 1.156*MAD. We calculate the average probability density function  of the set of normal distributions qi as described above. The mean and standard deviation of the interlaboratory study are obtained using the first and second moments of the  average probability density function .

## The NDA-mean (assigned value)

The NDA mean is centered around the highest density of values. Unless otherwise stated, the assigned value represents the consensus value of all data. Although all data are included in the assessment, those values that lie some distance from the NDA mean contribute less to the mean than values which occur at or near the mean.

With the NDA model, mean and standard deviation are calculated (under special conditions, see indicative values) using all reported data when at least 4 results are left after removal of reported ‘lower than’ (<) and 0 (=zero) values. No outliers are removed

## The Indicative Values

In some instances it is not possible to set an assigned value, and an indicative value is given.  No assessment of laboratory performance is given where an indicative value is set.  An overview of the assessment, with explanation, decision flowcharts and examples, is given in the paper Assessment Rules for the evaluation of the QUASIMEME Laboratory Performance Studies Data, available on the WEPAL-QUASIMEME website, www.wepal.nl.  Please note that this procedure can be used as QUASIMEME does not use the between-laboratory standard deviation as the standard deviation for proficiency assessment, but obtains the latter parameter independently from the dataset (see section on z’-score assessment)..

Category 1

For data with the number of numerical observations ≥ 7

An assigned value is based on the model mean when ≥ 50% of values have a z’-score of |z’| < 2 and at least 5 values have a |z’| < 3. The value is indicative when these conditions are not met.

Category 2

For data with the number of numerical observations > 3 and < 7

An assigned value is based on the model mean when ≥ 70% of values have a z’-score of |z’| < 3 and a minimum of 4 observations have |z’| < 2.  Otherwise the value is indicative.  This implies that for the small datasets under consideration, there needs to be very good agreement and a maximum of one extreme value before an assigned value can be given.

Category 3

For data with the number of numerical observations < 4

No assigned value is given.  Normally the median value is given as an indicative value.

Category 4

For data with a Total Error% >100%, meaning that there is a poor agreement among the participants, no assigned value is given. The median value is given as an indicative value.

### Traceability of the assigned value

The aim of this proficiency testing scheme is to establish comparability among laboratories. ”Real life” samples are collected, processed and distributed so that the true concentrations of measurands are not known. Assigned values are based on consensus values, obtained from the results of the participants using their routine methods. WEPAL is confident that the assigned values have an acceptable degree of traceability based on the following considerations:

1.	The large majority of participants in the programmes have an extensive experience with the analyses and are accredited, implying the use of validated methods, certified reference materials and internal quality control procedures in well equipped, maintained and managed laboratories.

2.	The programmes have an international participation, participants use standards that may differ according to country and/or methods that are based on different measurement principles;

3.	The data submitted by laboratories are analysed with robust statistics that associates a weighing factor to each individual result. The probability distributions of the raw and analysed data are graphically depicted and carefully examined. Consensus is assumed and an assigned value is established when the probability function of the weighed dataset (in terms of the model the mean pdf of the population of pdfs submitted by the laboratories) is near normal.

4.	The fact that consensus is achieved with data from experienced laboratories working with well-developed quality systems that employ a variety of methods is interpreted as evidence that possible biases arising from laboratories and/or methods is averaged out so that the consensus represents a reliable, traceable value. See Thompson (2016, 2018).

### Uncertainty of the assigned value

According to ISO 13528: 2015 (Cor. 2016-10), the uncertainty in the assigned value is calculated as:

ux = 1.25 * s / N

s = robust standard deviation

N = number of results

Depending on the NDA standard deviation calculated and the number of determinations observed, the uncertainty in the assigned value may influence the evaluation of the results (calculated Z-scores).  Therefore, the uncertainty is included in the Total Error which is used to evaluate the results from the participating laboratories.

## Median and MAD

For each determinand a median value and a median of absolute deviations (MAD) are calculated using all reported data except the reported '<' values. Deviating results like stragglers and outliers are not removed. The median is the middle observation of the sorted observations. In the case of an even number of observations it is the mean of the two middle observations. Using the median instead of mean, extreme data have less influence. MAD is the median of the absolute values of the observations minus their median.

## The z’-score Assessment

A z’-score  is calculated for each participant’s data for each matrix / determinand combination which is given an assigned value.

z'-score= Mean from Laboratory-Consensus Value (NDA mean)Total Error

It is emphasised that in many interlaboratory studies the between-laboratory standard deviation obtained from the statistical evaluation of the study is used as  the standard deviation for proficiency assessment in the formula above. In Quasimeme, the standard deviation for proficiency assessment is estimated independently taking the needs of present-day international monitoring programmes  as starting point. For each determinand in a particular matrix, a proportional error (PE) and a constant error (CE) have been defined by the Scientific Advisory Board of Quasimeme. Up to 2017, the standard deviation for proficiency assessment was calculated as the target error given by the following formula:

Target Error= Consensus Value*Proportional Error (%)100+0.5*Constant Error

From 2017 (Round 2016.2) onwards, the uncertainty of the assigned value is taken into consideration when calculating the z’-scores. To this end, the standard deviation for proficiency assessment is given by the total error defined as:

Total Error= UX2+(Target Error)2

The values for the Proportional Error (PE) and Constant Error (CE) are set by the Scientific Advisory Board and are monitored annually by the Quasimeme Project Office.  The values are based on the following criteria:

- Consistency of the required standard of performance to enable participating laboratories to monitor their assessment over time.

- Achievable targets in relation to the current state of the art and the level of performance needed for national and international monitoring programmes.

The assessment is based on ISO/IEC 17043:2010 as z’-scores. The QUASIMEME model is designed to provide a consistent interpretation over the whole range of concentration of analytes provided, including an assessment where Left Censored Values (LCVs) are reported.

The proportional error is set at 6% for nutrients and  standard solutions, and 12.5% for all other matrices.  This applies to all determinands, except salinity.  The constant error has been set for each determinand or determinand group (e.g. chlorinated biphenyls).  This value was initially set to reflect the limit of determination, but is at present more closely related to the overall laboratory performance.  The magnitude of the CE is set to provide a constant assessment in terms of z’-score regardless of concentration.  Therefore, at low concentrations the level of accuracy required to obtain a satisfactory z’-score is less stringent than at high concentrations.

The performance of the laboratories is examined in detail when the total error exceeds 100% of the consensus concentration.  If there is good agreement between the laboratories, i.e. the criteria to set an assigned value are met, the CE may be revised to a lower value reflecting the performance of laboratories for this measurement at lower concentrations.  These revisions are undertaken at the time of the assessment and ratified by the Scientific Advisory Board.  In making any adjustments to the CE an overall assessment of performance at these lower concentrations over a number of different rounds is reviewed.  This provides evidence of a long-term trend of improved performance rather than a single set of data. When the agreement is judged to be insufficient, no assigned value is established. In such cases an indicative value is given.

Following usual practices e.g. ISO/IEC 17043:2010, the z’-scores can be interpreted as follows for laboratories which take part in Quasimeme to assure the quality of their data for use in international marine monitoring programmes:

|z’| < 2 Satisfactory performance

2 <	|z’| < 3 Questionable performance

|z’| > 3 Unsatisfactory performance

The following figure illustrates the interpretation of the z’-scores:

Assigned

Value (AV)

AV

-

3*TE

AV

-

2*TE

AV+2*TE

AV+3*TE

Satisfactory

performance

Unsatisfactory

performance

performance

Questionable

performance

Assigned

Value (AV)

AV

-

3*TE

AV

AV+2*TE

AV+3*TE

Satisfactory

performance

performance

Unsatisfactory

performance

Questionable

performance

TE : total error

|z’| > 6 frequently points to gross errors (mistakes with units during reporting, calculation or dilution errors, and so on). Total error in the picture above is equal to the standard deviation for proficiency assessment.

It is not possible to calculate a z’-score for left censored values (LCV’s). Quasimeme provides a simple quality criterion:

AV‐2TE < LCV < AV+6TE : LCV consistent with assigned value.

LCV < AV – 2TE : inconsistent with assigned value, i.e. LCV reported would have been questionable or unsatisfactory, when reported as a numerical value.

LCV > AV + 6TE : inconsistent with assigned value, i.e. LCV reported by laboratory much higher than numerical values reported by other laboratories.

z’ score key: 	S – Satisfactory

Q – Questionable (*)

U – Unsatisfactory (**)

LCV key:	C – Consistent

I  –  Inconsistent

No data:	B – Blanc

## Evaluation of results

For the evaluation of results the absolute value of the z’-score is used. Questionable results 2 < | z’ | < 3  are marked as stragglers (*).  Deviating  results with | z’ | > 3  are marked as outliers (**).

For each participating laboratory, ranked z’-scores are calculated. The Possible % |z’| < 2 is calculated, based on the total number of |z’ | < 2 obtained by a laboratory, given as a percentage based on the total number of assigned values obtained in the study. The Actual % |z’| < 2 is calculated, based on the total number of |z’| < 2 obtained by a laboratory, given as a percentage based on the number of numerical values reported by the laboratory itself.

## Rounding of results

Rounding interval is set to have at least four significant digits for the statistical results. This is based on the value of the mean. If no mean value is available the median is used. In cases where between laboratory variation is small (based on the standard deviation) an extra digit is shown.

# Materials Analysed

Table 1 	Materials analysed in this period.

| Sample ID | Type | Origin |
| --- | --- | --- |
| QMP013MS | sand unspiked | VU Amsterdam |
| QMP014MS | sand spiked with PE, PET and PS | VU Amsterdam |
| QMP015MS | sediment Yerseke spiked with plastics | Westerscheldt, the Netherlands |

# Method Indicating Code (MIC)

In order to evaluate the analytical results for each reported element a Method Indicating Code (MIC) is used. Details of the analytical procedures used by the individual participants are indicated by a number of characters, added at the end of each row with results. With these MIC codes you can easily compare the results obtained by your laboratory with the results from other laboratories using the same results.

Method group	Total number of particles

Code	Sample filtration

-	None

QF	Quarz filter

SF	Synthetic filter

SS	Stainless steel filter

CF	Cellulose filter

PF	PFTE filter

AF	AIOx/Anodisk

GF	Glassfibre filter

CA	Combi cellulose / AIOx Anodisk

SP	Stainless steel and PTFE filter

AG	Ag filter 5 µm

SI	Silicon filter

PC	PC gold coated

PE	Polyester (PETG) gold-coated membrane

PA	PTFE (5µm) followed by anodisc (0.2 µm)

NY	Nylon net (20 µm) followed by Si membrane (1.2 µm)

Z	Others (Please inform quasimeme about details)

Code	Density separation

-	None

AA	ZnCl2 solution

NA	NaCl solution

NL	Nal solution

PT	Polytungstate

CC	CaCl2 solution

HC	HCOONa

AC	Combination of ZnCl2 and CaCl2

NN	Combination of NaCl an NaI

NB	NaBr solution

HK	HCOOK

OI	Oil separation

NP	Sodium Polytungstate, spiral conveyer technique

Z	Others (Please inform quasimeme about details)

Code	Digestion technique

-	None

HO	H2O2

HP	H2O2 (30%) at 50 °C for max 5 days

HN	Acid

AL	Alkaline

F	Fenton

ED	Enzymatic digestion (protease, lipase, amylase)

UR	1 urea+thiourea+NaOH; 2 H2O2 + NaOH

KH	KOH + H2O2

HF	H2O2 + enzymatic + fenton

KN	KOH and NAOCl

Z	Others (please inform quasimeme about details)

Code	Polymer identification and quantification

-	None

µF	µFTIR

AF	ATR-FTIR

FP	µFTIR-FPA

RA	Raman

PY	Pyr-GC/MS

µR	µ-Raman

MI	Optical Microscopy

GR	Gravimetric

TD	TED-GC-MS

NI	NIR

LD	LDIR

SE	SEM

FF	µFTIR + ATR-FTIR

µM	µFTIR + manual count FI microscope

FM	Fluorescent microcsopy based on Nile Red Staining and µFTIR

FA	Focal Plane Array

FI	FTIR imaging

MF	Microscopy + ATR-FTIR

SP	SPERO-QT

HN	Microscopy with hot needle test

Z	Others (please inform quasimeme about details)

Code	Stdev and technical aspects

Z	Please inform quasimeme about stdev of the method used and about smallest detectable particle size

Method group	Total mass of plastic particles

Code	Sample filtration

-	None

QF	Quarz filter

SF	Synthetic filter

SS	Stainless steel filter

CF	Cellulose filter

PF	PFTE filter

AF	AIOx/Anodisk

GF	Glassfibre filter

Z	Others (please inform QUASIMEME about the details)

Code	Density separation

-	None

AA	ZnCl2 solution i

NA	NaCl solution

NL	Nal solution

PT	Polytungstate

CC	CaCl2 solution

HC	HCOONa

NB	NaBr solution

Z	Others (Please inform QUASIMEME about the details)

Code	Digestion technique

-	None

HO	H2O2

HP	H2O2 (30%) qat 50 °C for max 5 days

HN	Acid

AL	Alkaline

F	Fenton

ED	Enzymatic digestion (protease, lipase, amylase)

AM	Ase extraction with methanol and CH2Cl2

AT	ASE extraction + MeOH + THF

AC	ASE extraction with CH2Cl2

AS	ASE

HE	H2O2 and enzymatic digestion

Z	Others (Please inform QUASIMEME about the details)

Code	Polymer identification and quantification

-	None

uF	µFTIR

AF	ATR-FTIR

RA	Raman

PY	Pyr-GC/MS

µR	µ-Raman

MI	Optical Microscopy

GR	Gravimetric

TD	TED-GC-MS

NI	NIR

SE	SEM

CD	Chloride determination by combustion-IC

Z	Others (Please inform QUASIMEME about the details)

Code	Stdev and technical aspects

Z	Please inform quasimeme about stdev of the method used and about smallest detectable particle size

# References and related literature

-	Cofino, W.P., I. van Stokkum, D.E. Wells, R.A.L. Peerboom, F. Ariese (2000). A new model for the inference of population characteristics from experimental data using uncertainties. Application to interlaboratory studies. Chemom. Intell. Lab. Syst. 53, 37-55.

-	Cofino, W.P., Molenaar, J. and Torfs, P. (2017) Evaluating Proficiency tests with Robust Statistics. Wiley StatsRef: Statistics Reference Online, 1-8.

- 	Dijk, D. van and V.J.G. Houba (2000). Homogeneity and Stability of Materials Distributed Within the Wageningen Evaluating Programmes for Analytical Laboratories. Commun. Soil Sci. Plant Anal. 31 (11-14), 1745 -1756.

-	Dijk, D. van, V.J.G. Houba and J.P.J. van Dalen (1996). Aspects of quality assurance within the Wageningen Evaluating Programmes for Analytical Laboratories (WEPAL). Commun. Soil SCi. Plant Anal. 27, 433 - 439.

-	Eurachem (2000). Selection, use and interpretation of proficiency testing (PT) schemes by laboratories. Eurachem Nederland, task group ‘proficiency testing schemes’ and Laboratory of the Government Chemist (LGC), United Kingdom.

-	Feinberg, M., E. Bugner, G. Theiller, V.J.G. Houba and F. Kadijk (1995). Expression of the reference value for proficiency tests. J. Chemometrics 9,197-209.

-	Houba, V.J.G. (1993). A device for automatic subsampling of soil, sediment and plant material for proficiency testing. Fresenius J. Anal. Chem. 345, 156 -157.

-	Houba, V.J.G., W.J. Chardon and K. Roelse (1993). Influence of grinding of soil on apparent chemical composition. Commun. Soil Sci. Plant Anal. 24, 1591 - 1602.

-	Houba, V.J.G., J.J. van der Lee and I. Novozamsky (1996). Evaluating the state-of-the-practice in soil measurements in relation to environmental regulations. Accred. Qual. Assur. 1, 92 - 98.

-	Houba, V.J.G. and I. Novozamsky (1998). Influence of storage time and temperature of air-dried soils on pH and extractable nutrients using 0.01 M CaCl2. Fresenius J. Anal. Chem. 360, 362 - 365.

-	Houba, V.J.G., I. Novozamsky and J.J. van der Lee (1994a). Status and future of soil and plant analysis. Commun. Soil Sci. Plant Anal. 25, 753 - 765.

-	Houba, V.J.G., I. Novozamsky and J.J. van der Lee (1994b). Standardization and validation of methods of soil and plant analysis as conditions for accreditation. Commun. Soil Sci. Plant Anal. 25, 827 - 841.

-	Houba, V.J.G., I. Novozamsky and J.J. van der Lee (1994c). Aspects of pre-treatment of soils for inorganic chemical analysis. QuRmica AnalitRca 13, 94 - 99.

-	Houba, V.J.G., I. Novozamsky and J.J. van der Lee (1995). Influence of storage of plant samples on their chemical composition. The Science of the Total Environment 176, 73 - 79.

-	Houba, V.J.G., I. Novozamsky and J.J. van der Lee (1996). Quality aspects in laboratories for soil and plant analysis. Commun. Soil Sci. Plant Anal. 27, 327 - 348.

-	Houba, V.J.G., J. Uittenbogaard and P. Pellen (1996). Wageningen Evaluating Programmes for Analytical Laboratories (WEPAL), organisation and purpose. Commun. Soil Sci. Plant Anal. 27, 421 - 431.

-  	ISO/IEC 17043:2010 (E). Conformity assessment – General requirements for proficiency testing. 39p.

-	ISO 13528: 2015 (Cor.2016-10), IDT Statistical methods for use in proficiency testing by interlaboratory comparison.

-	J. Molenaar  W. P. Cofino, P.J.J.F. Torfs (2018). Efficient and robust analysis of interlaboratory studies. Chemometrics and Intelligent Laboratory Systems 175 (2018) 65-73

-	Montfort, M. A.J. van (1996). Statistical remarks on laboratory-evaluating programs for comparing laboratories and methods. Commun. Soil Sci. Plant Anal. 27, 463 - 478.

-	Novozamsky, I., V.J.G. Houba, R.Ch. Daniel and the members of CII (1993). Certification of cabbage and carnation samples and their use in an international proficiency study. Fresenius J. Anal. Chem. 345, 198 - 201.

-	Sykes, M. and M. Thompson. Assessing the stability of a proficiency test material by participant-blind re-use after a period of storage. Analytical Methods, accepted manuscript royal society of chemistry.

-	Thompson, M. and R. Wood (1993). The international harmonized protocol for the proficiency testing of (chemical) analytical laboratories. Pure and Appl. Chem. 65, 2123 - 2144.

-	Thompson, M. (2011) Uncertainty functions, a compact way of summarising or specifying the behaviour of analytical systems. Trends in Analytical Chemistry 30-7, 1168-1175.

# Used abbreviations and symbols

Table 2	Used abbreviations and symbols

| Where | Abbreviation | Explanation |
| --- | --- | --- |
| Results | MIC | method indicating code |
| Results | NDA mean | Consensus value |
| Results | MAD | median absolute deviation |
| Results | NDA st dev. | standard deviation |
| Results | Coeff Var | NDA relative standard deviation (%) |
| Results, z’-scores | < | value smaller than |
| Results, z’-scores | * | straggler (questionable) |
| Results, z’-scores | ** | outlier (unsatisfactory) |
| Results, z’-scores | - | no result was submitted |
| Results statistical values | - | not calculated |
| z’-scores | # | less than 4 results, no mean and Sd calculated |
| z’-scores | & | no mean and Sd calculated |

Ranked z'-scores

DE17  2022.2

| Labcode | NObs ¦Z'¦<2 | Actual %¦Z'¦<2 |  | Labcode | NObs Submitted | Possible %¦Z'¦<2 |
| --- | --- | --- | --- | --- | --- | --- |
| Q4028 | 5 | 100 |  | Q3892 | 48 | 51 |
| Q114 | 5 | 100 |  | Q4049 | 43 | 48 |
| Q130 | 3 | 100 |  | Q3878 | 38 | 46 |
| Q4027 | 3 | 100 |  | Q4035 | 37 | 45 |
| Q3936 | 2 | 100 |  | Q122 | 31 | 45 |
| Q4026- 2 | 2 | 100 |  | Q110 | 47 | 43 |
| H221 | 1 | 100 |  | Q153 | 46 | 43 |
| Q3175 | 1 | 100 |  | Q101 | 45 | 43 |
| Q122 | 29 | 94 |  | Q131 | 34 | 40 |
| Q3013 | 17 | 81 |  | Q4001 | 41 | 37 |
| Q3887 | 12 | 80 |  | Q4014 | 31 | 37 |
| Q4039 | 4 | 80 |  | Q4037 | 32 | 35 |
| Q3878 | 30 | 79 |  | Q4018 | 28 | 32 |
| Q4035 | 29 | 78 |  | Q4040 | 41 | 29 |
| Q4014 | 24 | 77 |  | Q920 | 36 | 29 |
| Q131 | 26 | 76 |  | Q703 | 27 | 29 |
| Q4018 | 21 | 75 |  | Q4048 | 33 | 26 |
| Q3807 | 9 | 75 |  | Q3013 | 21 | 26 |
| Q4049 | 31 | 72 |  | Q4044 | 23 | 23 |
| Q4037 | 23 | 72 |  | Q3890 | 28 | 20 |
| Q703 | 19 | 70 |  | Q4026 | 19 | 20 |
| Q3892 | 33 | 69 |  | Q104 | 32 | 18 |
| Q4026 | 13 | 68 |  | Q3913 | 30 | 18 |
| Q4020 | 12 | 67 |  | Q4057 | 25 | 18 |
| Q4044 | 15 | 65 |  | Q4020 | 18 | 18 |
| Q3882 | 5 | 62 |  | Q3887 | 15 | 18 |
| Q101 | 28 | 62 |  | Q3807 | 12 | 14 |
| Q153 | 28 | 61 |  | Q3991 | 34 | 12 |
| Q110 | 28 | 60 |  | Q3277 | 37 | 11 |
| Q4001 | 24 | 59 |  | Q134 | 20 | 11 |
| Q920 | 19 | 53 |  | Q871 | 21 | 9 |
| Q4048 | 17 | 52 |  | Q3882 | 8 | 8 |
| Q102 | 4 | 50 |  | Q114 | 5 | 8 |
| Q3211 | 2 | 50 |  | Q4028 | 5 | 8 |
| Q980 | 1 | 50 |  | Q3946a | 9 | 6 |
| Q4057 | 12 | 48 |  | Q102 | 8 | 6 |
| Q3890 | 13 | 46 |  | Q4039 | 5 | 6 |
| Q4040 | 19 | 46 |  | Q3930 | 15 | 5 |
| Q3946a | 4 | 44 |  | Q130 | 3 | 5 |
| Q3913 | 12 | 40 |  | Q4027 | 3 | 5 |
| Q104 | 12 | 38 |  | Q3875 | 47 | 3 |
| Q134 | 7 | 35 |  | Q3211 | 4 | 3 |
| Q871 | 6 | 29 |  | Q4026- 2 | 2 | 3 |
| Q3991 | 8 | 24 |  | Q3936 | 2 | 3 |
| Q3930 | 3 | 20 |  | Q3940 | 9 | 2 |
| Q3277 | 7 | 19 |  | Q980 | 2 | 2 |
| Q3940 | 1 | 11 |  | Q3175 | 1 | 2 |
| Q3875 | 2 | 4 |  | H221 | 1 | 2 |
| Q4047 | 0 | 0 |  | Q3889 | 18 | 0 |
| Q4029 | 0 | 0 |  | Q4023 | 9 | 0 |
| Q4023 | 0 | 0 |  | Q3926 | 3 | 0 |
| Q3926 | 0 | 0 |  | Q3935 | 3 | 0 |
| Q3932 | 0 | 0 |  | Q3932 | 2 | 0 |
| Q3889 | 0 | 0 |  | Q4029 | 1 | 0 |
| Q3935 | 0 | 0 |  | Q4047 | 1 | 0 |

Total number of particles

Analysis DE17  2022.2

| Sample/ Determinand | Assigned Value | Units | Total Error | NDA st.dev | NDA rel. st.dev (%) | Nobs numerical | Nobs LCV | Median | MAD | Model Mean | Uncer- tainty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QMP013MS |  |  |  |  |  |  |  |  |  |  |  |
| PC (50-299 µm) | 69.7 | (No. p/kg) | 35.9 | 78.7 | 112.9 | 10 | 0 | 82.5 | 54.5 | 69.67 | 31.09 |
| PC (300-5000 µm) |  | (No. p/kg) |  |  |  | 1 | 0 | 1.00 |  |  |  |
| Polycarbonate total | 69.7 | (No. p/kg) | 35.9 | 78.7 | 112.9 | 10 | 0 | 82.5 | 54.5 | 69.67 | 31.09 |
| PE (50-299 µm) | 83.1 | (No. p/kg) | 34.2 | 100.2 | 120.5 | 22 | 0 | 81.0 | 65.0 | 83.15 | 26.71 |
| PE (300-5000 µm) |  | (No. p/kg) |  |  |  | 4 | 0 | 32.1 | 6.1 |  |  |
| Polyethylene total (HD+LD) | 99.7 | (No. p/kg) | 40.4 | 123.0 | 123.4 | 24 | 0 | 99.0 | 87.3 | 99.70 | 31.39 |
| PET (50-299 µm) | 42.6 | (No. p/kg) | 19.7 | 43.2 | 101.5 | 11 | 0 | 49.0 | 34.0 | 42.60 | 16.29 |
| PET (300-5000 µm) | 334 | (No. p/kg) | 268 | 408 | 122.0 | 4 | 0 | 345 | 289 | 334.3 | 255.0 |
| Polyethylene Terephat. Total | 63.0 | (No. p/kg) | 32.7 | 81.7 | 129.7 | 13 | 0 | 96.1 | 64.3 | 63.03 | 28.33 |
| PP (50-299 µm) | 168 | (No. p/kg) | 69 | 181 | 107.4 | 17 | 0 | 158 | 126 | 168.2 | 54.8 |
| PP (300-5000 µm) |  | (No. p/kg) |  |  |  | 3 | 0 | 20.0 | 4.0 |  |  |
| Polypropylene total | 155 | (No. p/kg) | 64 | 177 | 113.9 | 19 | 0 | 158 | 126 | 155.3 | 50.7 |
| PS (50-299 µm) | 30.6 | (No. p/kg) | 13.3 | 31.3 | 102.3 | 14 | 0 | 37.6 | 23.3 | 30.63 | 10.47 |
| Polystyrene total | 30.6 | (No. p/kg) | 13.3 | 31.3 | 102.3 | 14 | 0 | 37.6 | 23.3 | 30.63 | 10.47 |
| PVC (50-299 µm) |  | (No. p/kg) |  |  |  | 6 | 0 | 51.0 | 27.8 |  |  |
| PVC (300-5000 µm) |  | (No. p/kg) |  |  |  | 1 | 0 | 16.0 |  |  |  |
| Polyvinyl Chloride total |  | (No. p/kg) |  |  |  | 7 | 0 | 40.0 | 24.0 |  |  |
| Other polymers (50-299 µm) | 139 | (No. p/kg) | 60 | 159 | 114.2 | 17 | 0 | 157 | 117 | 138.8 | 48.1 |
| Other polymers (300-5000µm) | 161 | (No. p/kg) | 61 | 96 | 59.4 | 7 | 0 | 199 | 59 | 161.2 | 45.2 |
| Other polymers total | 185 | (No. p/kg) | 74 | 207 | 111.9 | 20 | 0 | 180 | 144 | 185.1 | 57.9 |
| Total polymers (50-299 µm) | 367 | (No. p/kg) | 143 | 487 | 132.7 | 31 | 0 | 410 | 346 | 367.2 | 109.4 |
| Total polymers (300-5000 µm) | 90.4 | (No. p/kg) | 57.3 | 173.0 | 191.5 | 17 | 0 | 140.0 | 124.0 | 90.36 | 52.45 |
| Total polymers (50-5000 µm) | 522 | (No. p/kg) | 194 | 676 | 129.5 | 35 | 0 | 600 | 466 | 521.7 | 142.8 |

| Sample/ Determinand | Assigned Value | Units | Total Error | NDA st.dev | NDA rel. st.dev (%) | Nobs numerical | Nobs LCV | Median | MAD | Model Mean | Uncer- tainty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QMP014MS |  |  |  |  |  |  |  |  |  |  |  |
| PC (50-299 µm) | 26.3 | (No. p/kg) | 14.5 | 28.6 | 108.9 | 8 | 0 | 27.9 | 19.0 | 26.26 | 12.64 |
| Polycarbonate total | 26.3 | (No. p/kg) | 14.5 | 28.6 | 108.9 | 8 | 0 | 27.9 | 19.0 | 26.26 | 12.64 |
| PE (50-299 µm) | 104 | (No. p/kg) | 42 | 126 | 121.8 | 24 | 0 | 103 | 87 | 103.8 | 32.2 |
| PE (300-5000 µm) | 69.2 | (No. p/kg) | 26.1 | 61.2 | 88.4 | 16 | 0 | 80.1 | 46.4 | 69.20 | 19.12 |
| Polyethylene total (HD+LD) | 161 | (No. p/kg) | 61 | 197 | 122.4 | 29 | 0 | 178 | 145 | 160.5 | 45.6 |
| PET (50-299 µm) | 31.8 | (No. p/kg) | 13.2 | 32.3 | 101.6 | 16 | 0 | 40.3 | 24.3 | 31.76 | 10.08 |
| PET (300-5000 µm) |  | (No. p/kg) |  |  |  | 5 | 0 | 48.0 | 32.0 |  |  |
| Polyethylene Terephat. Total | 43.8 | (No. p/kg) | 26.4 | 78.6 | 179.7 | 17 | 0 | 79.0 | 62.9 | 43.77 | 23.84 |
| PP (50-299 µm) | 116 | (No. p/kg) | 43 | 102 | 87.5 | 16 | 0 | 137 | 71 | 116.3 | 31.8 |
| PP (300-5000 µm) | 16.3 | (No. p/kg) | 4.6 | 1.3 | 8.0 | 7 | 0 | 17.0 | 1.1 | 16.27 | 0.62 |
| Polypropylene total | 92.3 | (No. p/kg) | 37.9 | 108.7 | 117.7 | 21 | 0 | 95.3 | 79.2 | 92.35 | 29.66 |
| PS (50-299 µm) | 162 | (No. p/kg) | 61 | 175 | 107.7 | 24 | 0 | 149 | 111 | 162.2 | 44.5 |
| PS (300-5000 µm) | 93.4 | (No. p/kg) | 34.3 | 71.0 | 76.0 | 13 | 0 | 95.2 | 49.8 | 93.36 | 24.61 |
| Polystyrene total | 278 | (No. p/kg) | 95 | 257 | 92.6 | 25 | 0 | 259 | 180 | 278.1 | 64.4 |
| PVC (50-299 µm) |  | (No. p/kg) |  |  |  | 5 | 0 | 20.0 | 5.0 |  |  |
| PVC (300-5000 µm) |  | (No. p/kg) |  |  |  | 1 | 0 | 49.0 |  |  |  |
| Polyvinyl Chloride total | 28.4 | (No. p/kg) | 13.8 | 22.4 | 78.9 | 6 | 0 | 34.0 | 16.6 | 28.45 | 11.45 |
| Other polymers (50-299 µm) | 54.6 | (No. p/kg) | 23.3 | 60.8 | 111.4 | 17 | 0 | 63.1 | 47.0 | 54.64 | 18.45 |
| Other polymers (300-5000µm) | 47.0 | (No. p/kg) | 25.9 | 51.7 | 110.0 | 8 | 0 | 55.7 | 39.3 | 47.00 | 22.85 |
| Other polymers total | 51.1 | (No. p/kg) | 22.1 | 63.1 | 123.6 | 20 | 0 | 63.0 | 46.9 | 51.08 | 17.65 |
| Total polymers (50-299 µm) | 367 | (No. p/kg) | 130 | 431 | 117.5 | 35 | 0 | 363 | 300 | 366.8 | 91.0 |
| Total polymers (300-5000 µm) | 187 | (No. p/kg) | 69 | 203 | 108.7 | 26 | 0 | 197 | 148 | 187.0 | 49.8 |
| Total polymers (50-5000 µm) | 658 | (No. p/kg) | 229 | 771 | 117.1 | 37 | 0 | 805 | 568 | 658.4 | 158.4 |

| Sample/ Determinand | Assigned Value | Units | Total Error | NDA st.dev | NDA rel. st.dev (%) | Nobs numerical | Nobs LCV | Median | MAD | Model Mean | Uncer- tainty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QMP015MS |  |  |  |  |  |  |  |  |  |  |  |
| PC (50-299 µm) | 74.5 | (No. p/kg) | 24.5 | 48.8 | 65.6 | 16 | 0 | 75.5 | 34.4 | 74.46 | 15.27 |
| PC (300-5000 µm) | 33.1 | (No. p/kg) | 16.1 | 34.1 | 103.1 | 10 | 0 | 35.0 | 26.0 | 33.05 | 13.47 |
| Polycarbonate total | 82.3 | (No. p/kg) | 24.6 | 44.0 | 53.4 | 19 | 0 | 87.5 | 30.5 | 82.31 | 12.61 |
| PE (50-299 µm) | 207 | (No. p/kg) | 90 | 299 | 144.2 | 26 | 0 | 222 | 205 | 207.1 | 73.2 |
| PE (300-5000 µm) | 79.3 | (No. p/kg) | 24.2 | 57.4 | 72.4 | 30 | 0 | 87.5 | 41.5 | 79.33 | 13.11 |
| Polyethylene total (HD+LD) | 265 | (No. p/kg) | 97 | 317 | 119.5 | 32 | 0 | 281 | 221 | 264.9 | 69.9 |
| PET (50-299 µm) | 51.9 | (No. p/kg) | 24.0 | 59.4 | 114.6 | 14 | 0 | 57.7 | 42.5 | 51.87 | 19.85 |
| PET (300-5000 µm) | 37.9 | (No. p/kg) | 31.8 | 54.0 | 142.5 | 5 | 0 | 61.0 | 43.6 | 37.86 | 30.16 |
| Polyethylene Terephat. Total | 40.8 | (No. p/kg) | 19.0 | 52.0 | 127.3 | 17 | 0 | 54.0 | 39.0 | 40.82 | 15.75 |
| PP (50-299 µm) | 67.2 | (No. p/kg) | 29.2 | 84.2 | 125.4 | 20 | 0 | 95.5 | 63.3 | 67.16 | 23.54 |
| PP (300-5000 µm) | 46.3 | (No. p/kg) | 24.3 | 53.2 | 114.8 | 10 | 0 | 53.5 | 39.0 | 46.33 | 21.03 |
| Polypropylene total | 156 | (No. p/kg) | 62 | 180 | 115.2 | 22 | 0 | 156 | 114 | 156.3 | 48.0 |
| PS (50-299 µm) | 46.1 | (No. p/kg) | 22.2 | 53.7 | 116.5 | 13 | 0 | 47.4 | 39.4 | 46.12 | 18.63 |
| PS (300-5000 µm) |  | (No. p/kg) |  |  |  | 1 | 0 | 8.00 |  |  |  |
| Polystyrene total | 46.9 | (No. p/kg) | 22.4 | 54.1 | 115.4 | 13 | 0 | 47.4 | 39.4 | 46.89 | 18.76 |
| PVC (50-299 µm) | 75.4 | (No. p/kg) | 26.2 | 59.8 | 79.3 | 18 | 0 | 80.4 | 44.5 | 75.36 | 17.61 |
| PVC (300-5000 µm) | 22.9 | (No. p/kg) | 11.7 | 26.3 | 114.4 | 11 | 0 | 23.0 | 18.0 | 22.94 | 9.89 |
| Polyvinyl Chloride total | 93.7 | (No. p/kg) | 36.4 | 98.1 | 104.6 | 20 | 0 | 103.5 | 72.5 | 93.75 | 27.41 |
| Other polymers (50-299 µm) | 63.8 | (No. p/kg) | 35.3 | 105.9 | 166.0 | 18 | 0 | 92.5 | 80.0 | 63.82 | 31.22 |
| Other polymers (300-5000µm) | 40.3 | (No. p/kg) | 21.2 | 52.9 | 131.3 | 13 | 0 | 53.6 | 38.6 | 40.27 | 18.34 |
| Other polymers total | 78.0 | (No. p/kg) | 37.8 | 120.2 | 154.2 | 22 | 0 | 102.5 | 90.0 | 77.97 | 32.04 |
| Total polymers (50-299 µm) | 463 | (No. p/kg) | 192 | 681 | 147.2 | 31 | 0 | 562 | 503 | 462.6 | 152.9 |
| Total polymers (300-5000 µm) | 148 | (No. p/kg) | 48 | 150 | 101.3 | 38 | 0 | 152 | 103 | 148.2 | 30.5 |
| Total polymers (50-5000 µm) | 596 | (No. p/kg) | 229 | 858 | 144.0 | 38 | 0 | 739 | 630 | 595.6 | 173.9 |

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PC (50-299 µm) ((No. p/kg))

Q101	213	**	81.0	**	75.0				PF|NL| -|FM

Q104	-  		-  		213	**			SI|AA|KN|µF|ZA

Q110	-  		-  		145	*			CF|AA|HP|µF

Q131	-  		-  		23.9	*			CF|NL| -|µF

Q153	-  		-  		50.0				AF|AA|HO|FP

Q3013	-  		-  		122

Q3875	312	**	941	**	-  				SS|AA|KH|LD

Q3878	111		15.7		100				SI|NL|HF|µF

Q3887	-  		-  		69.0				SS|AA| F|AF|ZD

Q3892	-  		-  		92.0

Q3913	114		-  		240	**			PF|NL|KH|FM

Q3991	-  		-  		121				ZB|AA| -|µF|ZB

Q4001	-  		-  		40.0				SI|NP|HO|µR|ZB

Q4014	-  		-  		30.0

Q4018	40.0		20.0		-  				CF|NA|HO|µF

Q4035	16.0		-  		-  				SF|OI| F|µF|ZB

Q4037	53.9		35.8		42.1				SS|CC| -|FA|ZB

Q4040	123		77.0	**	76.0				SS|CC| -|LD|ZA

Q4048	1.00		2.00		-  				PF|NL| F|µF|ZA

Q4057	15.7		15.8		-

Q703	-  		-  		61.0				SS|NL|HO|µR|ZB

=================     Statistical Results     =================

NDA mean	69.67		26.26		74.46

NDA st dev	78.66		28.59		48.85

Coeff Var (%)	112.9		108.9		65.6

N	10		8		16

Median	82.48		27.88		75.50

MAD	54.48		19.00		34.45

Total Error	35.89		14.48		24.46

====================================================

PC (300-5000 µm) ((No. p/kg))

Q110	-  		-  		30.0				CF|AA|HP|µF

Q131	-  		-  		63.6				CF|NL| -|µF

Q153	-  		-  		63.0				AF|AA|HO|FP

Q3013	-  		-  		8.00

Q3889	-  		-  		2600	**

Q3913	-  		-  		23.0				PF|NL|KH|FM

Q4001	-  		-  		40.0				SI|NP|HO|µR|ZB

Q4020	1.00		-  		-  				GF|NA|HO|AF

Q4049	-  		-  		57.0				 -|NL|HP|AF

Q703	-  		-  		8.00				SS|NL|HO|µR|ZB

Q920	-  		-  		10.0

=================     Statistical Results     =================

NDA mean	-  		-  		33.05

NDA st dev	-  		-  		34.07

Coeff Var (%)	-  		-  		103.1

N	1		-		10

Median	1.000		-  		35.00

MAD	-  		-  		26.00

Total Error	-  		-  		16.07

Category	3		-  		-

====================================================

Polycarbonate total ((No. p/kg))

Q101	213	**	81.0	**	75.0				PF|NL| -|FM

Q104	-  		-  		213	**			SI|AA|KN|µF|ZA

Q110	-  		-  		175	**			CF|AA|HP|µF

=================     Summary Statistics     ================

NDA mean	69.67		26.26		82.31

NDA st dev	78.66		28.59		43.98

Coeff Var (%)	112.9		108.9		53.4

N	10		8		19			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polycarbonate total ((No. p/kg))     (cont.)

Q131	-  		-  		87.5				CF|NL| -|µF

Q153	-  		-  		113				AF|AA|HO|FP

Q3013	-  		-  		130

Q3875	312	**	941	**	-  				SS|AA|KH|LD

Q3878	111		15.7		100				SI|NL|HF|µF

Q3882	-  		-  		4.00	**			SS|AA| -|µF|ZC

Q3887	-  		-  		69.0				SS|AA| F|AF|ZD

Q3889	-  		-  		2600	**

Q3892	-  		-  		92.0

Q3913	114		-  		263	**			PF|NL|KH|FM

Q3991	-  		-  		121				ZB|AA| -|µF|ZB

Q4001	-  		-  		80.0				SI|NP|HO|µR|ZB

Q4018	40.0		20.0		-  				CF|NA|HO|µF

Q4035	16.0		-  		-  				SF|OI| F|µF|ZB

Q4037	53.9		35.8		42.1

Q4040	123		77.0	**	76.0				SS|CC| -|LD|ZA

Q4048	1.00		2.00		-  				PF|NL| F|µF|ZA

Q4049	-  		-  		57.0

Q4057	15.7		15.8		-

Q703	-  		-  		69.0				SS|NL|HO|µR|ZB

Q920	-  		-  		10.0	*

=================     Statistical Results     =================

NDA mean	69.67		26.26		82.31

NDA st dev	78.66		28.59		43.98

Coeff Var (%)	112.9		108.9		53.4

N	10		8		19

Median	82.48		27.88		87.50

MAD	54.48		19.00		30.50

Total Error	35.89		14.48		24.56

====================================================

PE (50-299 µm) ((No. p/kg))

Q101	-  		113		33.0				PF|NL| -|FM

Q104	285	**	236	**	5242	**			SI|AA|KN|µF|ZA

Q110	48.0		95.0		1028	**			CF|AA|HP|µF

Q122	115		67.0		66.0				GF|NN| -|µF

Q131	-  		48.3		55.7				CF|NL| -|µF

Q134	33.6		16.7	*	10.0	*			GF|NA|HO|FF

Q153	83.0		333	**	825	**			AF|AA|HO|FP

Q3013	48.0		-  		297

Q3277	263	**	390	**	569	**			AF|PT|HO|µF

Q3875	312	**	-  		1220	**			SS|AA|KH|LD

Q3878	31.7		15.7	*	146				SI|NL|HF|µF

Q3887	79.0		-  		8.00	*			SS|AA| F|AF|ZD

Q3890	191	**	15.8	*	76.2				ZD|ZC|ZC|µR|ZD

Q3892	16.1		112		628	**

Q3913	225	**	219	*	139				PF|NL|KH|FM

Q3930	195	**	401	**	-  				SF|AA| -|RA|ZE

Q3932	-  		-  		11.0	*			GF|ZD| -|ZB|ZD

Q3991	-  		-  		969	**			ZB|AA| -|µF|ZB

Q4001	-  		275	**	350				SI|NP|HO|µR|ZB

Q4014	16.0		47.0		38.0				ZA|NA|ZA|ZA|ZA

Q4035	16.0		79.4		-  				SF|OI| F|µF|ZB

Q4037	38.5		28.6		5.26	*			SS|CC| -|FA|ZB

Q4040	262	**	92.0		417	*			SS|CC| -|LD|ZA

Q4044	-  		111		22.9	*			CF|NA| -|µF

Q4048	4.00	*	2.00	*	-  				PF|NL| F|µF|ZA

Q4049	47.0		-  		372				PF|NL|HP|µF

Q4057	1756	**	1579	**	-

=================     Summary Statistics     ================

NDA mean	83.15		103.8		207.1

NDA st dev	100.22		126.3		298.6

Coeff Var (%)	120.5		121.8		144.2

N	22		24		26			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PE (50-299 µm) ((No. p/kg))     (cont.)

Q703	-  		252	**	69.0				SS|NL|HO|µR|ZB

Q871	-  		36.0		450	*

Q920	1840	**	500	**	500	**

=================     Statistical Results     =================

NDA mean	83.15		103.8		207.1

NDA st dev	100.22		126.3		298.6

Coeff Var (%)	120.5		121.8		144.2

N	22		24		26

Median	81.00		103.0		221.6

MAD	65.00		86.8		204.6

Total Error	34.15		41.7		89.9

====================================================

PE (300-5000 µm) ((No. p/kg))

Q101	-  		65.0		33.0				PF|NL| -|FM

Q104	-  		-  		84.0				SI|AA|KN|µF|ZA

Q110	-  		126	*	327	**			CF|AA|HP|µF

Q122	-  		-  		99.0				GF|NN| -|µF

Q130	-  		-  		91.0				SS|ZC|UR|AF|ZB

Q131	-  		96.5		167	**			CF|NL| -|µF

Q134	-  		16.7	*	40.1				GF|NA|HO|FF

Q153	-  		33.0		63.0				AF|AA|HO|FP

Q3013	-  		-  		91.0

Q3277	-  		-  		68.0				AF|PT|HO|µF

Q3807	-  		-  		160	**			SI|AA| F|LD|ZA

Q3875	-  		314	**	-  				SS|AA|KH|LD

Q3878	-  		-  		192	**			SI|NL|HF|µF

Q3887	-  		-  		54.0				SS|AA| F|AF|ZD

Q3889	7675		5129	**	10401	**			ZC| -|ZA|µF

Q3890	-  		-  		91.5				ZD|ZC|ZC|µR|ZD

Q3892	32.2		-  		46.0

Q3930	-  		-  		61.0				SF|AA| F|RA|ZF

Q3991	-  		127	*	253	**			ZB|AA| -|µF|ZB

Q4001	-  		81.0		310	**			SI|NP|HO|µR|ZB

Q4014	-  		-  		76.0				CF|NA| -|AF|ZB

Q4018	20.0		20.0		129	*			CF|NA|HO|µF

Q4020	-  		32.3		46.2				GF|NA|HO|AF

Q4026	-  		-  		15.3	*			AF|AA|HO|µF|ZC

Q4035	32.0		95.2		69.5				SF|OI| F|µF|ZB

Q4037	-  		-  		42.1				SS|CC| -|AF

Q4040	-  		-  		76.0				SS|CC| -|LD|ZA

Q4044	-  		79.2		38.1				CF|NA| -|µF

Q4048	-  		8.00	*	-  				PF|NL| F|µF|ZA

Q4049	-  		63.0		129	*			 -|NL|HP|AF

Q703	-  		-  		99.0				SS|NL|HO|µR|ZB

Q920	-  		120		210	**

=================     Statistical Results     =================

NDA mean	-  		69.20		79.33

NDA st dev	-  		61.19		57.45

Coeff Var (%)	-  		88.4		72.4

N	4		16		30

Median	32.12		80.12		87.50

MAD	6.11		46.38		41.50

Total Error	-  		26.12		24.19

Category	2		-  		-

====================================================

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polyethylene total (HD+LD) ((No. p/kg))

Q101	-  		178		66.0	*			PF|NL| -|FM

Q104	285	**	236		5326	**			SI|AA|KN|µF|ZA

Q110	48.0		221		1355	**			CF|AA|HP|µF

Q122	115		67.0		166				GF|NN| -|µF

Q131	-  		145		223				CF|NL| -|µF

Q134	33.6		33.3	*	50.1	*			GF|NA|HO|FF

Q153	83.0		366	**	888	**			AF|AA|HO|FP

Q3013	-  		-  		389

Q3277	263	**	390	**	637	**			AF|PT|HO|µF

Q3807	-  		-  		160				SI|AA| F|LD|ZA

Q3875	312	**	313	*	1220	**			SS|AA|KH|LD

Q3878	31.7		15.7	*	338				SI|NL|HF|µF

Q3882	-  		-  		520	*			SS|AA| -|µF|ZC

Q3887	79.0		-  		62.0	*			SS|AA| F|AF|ZD

Q3889	7675	**	5129	**	10401	**

Q3890	191	*	15.8	*	168				ZD|ZC|ZC|µR|ZD

Q3892	48.3		112		674	**

Q3913	225	**	219		139				PF|NL|KH|FM

Q3930	195	*	401	**	-  				SF|AA| -|RA|ZE

Q3991	-  		-  		1222	**			ZB|AA| -|µF|ZB

Q4001	-  		356	**	660	**			SI|NP|HO|µR|ZB

Q4014	16.0	*	47.0		114				 -| -| -| -|ZC

Q4018	20.0		20.0	*	129				CF|NA|HO|µF

Q4020	-  		32.3	*	46.2	*			GF|NA|HO|AF

Q4026	-  		-  		15.3	*			AF|AA|HO|µF|ZC

Q4035	48.0		175		69.5	*			SF|OI| F|µF|ZB

Q4037	38.5		28.6	*	47.4	*

Q4040	262	**	92.0		492	*			SS|CC| -|LD|ZA

Q4044	-  		190		61.0	*			CF|NA| -|µF

Q4048	4.00	*	10.0	*	-  				PF|NL| F|µF|ZA

Q4049	47.0		63.0		502	*

Q4057	1756	**	1579	**	-

Q703	-  		252		168				SS|NL|HO|µR|ZB

Q871	1242	**	648	**	450

Q920	1840	**	620	**	710	**

=================     Statistical Results     =================

NDA mean	99.70		160.5		264.9

NDA st dev	123.04		196.5		316.5

Coeff Var (%)	123.4		122.4		119.5

N	24		29		32

Median	99.00		178.0		280.6

MAD	87.33		144.7		220.5

Total Error	40.40		61.1		96.7

====================================================

PET (50-299 µm) ((No. p/kg))

Q101	-  		97.0	**	-  				PF|NL| -|FM

Q104	-  		-  		198	**			SI|AA|KN|µF|ZA

Q110	-  		16.0		15.0				CF|AA|HP|µF

Q122	49.0		33.0		41.0				GF|NN| -|µF

Q131	-  		16.1		-  				CF|NL| -|µF

Q153	-  		17.0		25.0				AF|AA|HO|FP

Q3013	-  		-  		8.00

Q3277	482	**	1172	**	498	**			AF|PT|HO|µF

Q3875	2186	**	5334	**	5183	**			SS|AA|KH|LD

Q3878	31.7		-  		100	*			SI|NL|HF|µF

Q3887	127	**	-  		54.0				SS|AA| F|AF|ZD

Q3892	48.3		32.0		61.3

Q3991	516	**	253	**	-  				ZB|AA| -|µF|ZB

=================     Summary Statistics     ================

NDA mean	42.60		31.76		51.87

NDA st dev	43.22		32.27		59.42

Coeff Var (%)	101.5		101.6		114.6

N	11		16		14			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PET (50-299 µm) ((No. p/kg))     (cont.)

Q4001	-  		32.0		159	**			SI|NP|HO|µR|ZB

Q4018	-  		20.0		-  				CF|NA|HO|µF

Q4035	32.0		47.6		-  				SF|OI| F|µF|ZB

Q4037	-  		-  		15.8				SS|CC| -|FA|ZB

Q4040	46.0		154	**	106	*			SS|CC| -|LD|ZA

Q4048	15.0		20.0		-  				PF|NL| F|µF|ZA

Q4057	298	**	79.0	**	-

Q703	-  		363	**	30.0				SS|NL|HO|µR|ZB

=================     Statistical Results     =================

NDA mean	42.60		31.76		51.87

NDA st dev	43.22		32.27		59.42

Coeff Var (%)	101.5		101.6		114.6

N	11		16		14

Median	49.00		40.31		57.66

MAD	34.00		24.26		42.50

Total Error	19.74		13.15		23.99

====================================================

PET (300-5000 µm) ((No. p/kg))

Q101	643		466		61.0				PF|NL| -|FM

Q3875	625		-  		305	**			SS|AA|KH|LD

Q3991	-  		127		-  				ZB|AA| -|µF|ZB

Q4001	-  		48.0		64.0				SI|NP|HO|µR|ZB

Q4014	-  		16.0		-

Q4035	64.0		31.7		17.4				SF|OI| F|µF|ZB

Q4048	1.00		-  		-  				PF|NL| F|µF|ZA

Q4049	-  		-  		8.00				 -|NL|HP|AF

=================     Statistical Results     =================

NDA mean	334.3		-  		37.86

NDA st dev	408.0		-  		53.96

Coeff Var (%)	122.0		-  		142.5

N	4		5		5

Median	344.5		48.00		61.00

MAD	289.5		32.00		43.63

Total Error	268.5		-  		31.77

Category	-  		2		-

====================================================

Polyethylene Terephat. Total ((No. p/kg))

Q101	643	**	563	**	61.0				PF|NL| -|FM

Q104	-  		-  		198	**			SI|AA|KN|µF|ZA

Q110	-  		16.0		15.0				CF|AA|HP|µF

Q122	49.0		33.0		41.0				GF|NN| -|µF

Q131	-  		16.1		-  				CF|NL| -|µF

Q153	-  		17.0		25.0				AF|AA|HO|FP

Q3013	-  		-  		8.00

Q3277	482	**	1172	**	498	**			AF|PT|HO|µF

Q3875	2811	**	5334	**	5488	**			SS|AA|KH|LD

Q3878	31.7		-  		100	**			SI|NL|HF|µF

Q3882	67.0		-  		-  				SS|AA| -|µF|ZC

Q3887	127		-  		54.0				SS|AA| F|AF|ZD

Q3892	48.3		32.0		61.3

Q3991	516	**	380	**	-  				ZB|AA| -|µF|ZB

Q4001	-  		80.0		223	**			SI|NP|HO|µR|ZB

Q4014	-  		16.0		-

Q4018	-  		20.0		-  				CF|NA|HO|µF

Q4035	96.1		79.4		17.4				SF|OI| F|µF|ZB

Q4037	-  		-  		15.8

=================     Summary Statistics     ================

NDA mean	63.03		43.77		40.82

NDA st dev	81.72		78.64		51.97

Coeff Var (%)	129.7		179.7		127.3

N	13		17		17			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polyethylene Terephat. Total ((No. p/kg))     (cont.)

Q4040	46.0		154	**	106	**			SS|CC| -|LD|ZA

Q4048	16.0		20.0		-  				PF|NL| F|µF|ZA

Q4049	-  		-  		8.00

Q4057	298	**	79.0		-

Q703	-  		363	**	30.0				SS|NL|HO|µR|ZB

=================     Statistical Results     =================

NDA mean	63.03		43.77		40.82

NDA st dev	81.72		78.64		51.97

Coeff Var (%)	129.7		179.7		127.3

N	13		17		17

Median	96.05		79.00		54.00

MAD	64.30		62.90		39.00

Total Error	32.66		26.45		19.05

====================================================

PP (50-299 µm) ((No. p/kg))

Q101	-  		33.0		-  				PF|NL| -|FM

Q104	158		157		465	**			SI|AA|KN|µF|ZA

Q110	32.0		16.0	*	53.0				CF|AA|HP|µF

Q122	115		83.0		74.0				GF|NN| -|µF

Q131	-  		-  		39.8				CF|NL| -|µF

Q153	300		117		375	**			AF|AA|HO|FP

Q3013	-  		-  		53.0

Q3277	175		683	**	597	**			AF|PT|HO|µF

Q3875	312	*	628	**	610	**			SS|AA|KH|LD

Q3878	-  		-  		69.2				SI|NL|HF|µF

Q3887	63.0		-  		-  				SS|AA| F|AF|ZD

Q3890	-  		-  		152	*			ZD|ZC|ZC|µR|ZD

Q3892	32.2		80.1		337	**

Q3991	516	**	-  		545	**			ZB|AA| -|µF|ZB

Q4001	261		161		373	**			SI|NP|HO|µR|ZB

Q4014	158		268	**	91.0				ZA|NA|ZA|ZA|ZA

Q4026	-  		79.4		22.9				AF|AA|HO|µF|ZC

Q4037	-  		-  		26.3				SS|CC| -|FA|ZB

Q4040	508	**	200		432	**			SS|CC| -|LD|ZA

Q4048	3.00	*	2.00	*	-  				PF|NL| F|µF|ZA

Q4049	16.0	*	63.0		49.0				PF|NL|HP|µF

Q4057	251		205	*	-

Q703	-  		-  		38.0				SS|NL|HO|µR|ZB

Q871	18.0	*	-  		-

Q920	380	**	280	**	100

=================     Statistical Results     =================

NDA mean	168.2		116.3		67.16

NDA st dev	180.6		101.8		84.20

Coeff Var (%)	107.4		87.5		125.4

N	17		16		20

Median	158.0		137.0		95.50

MAD	126.0		71.2		63.34

Total Error	69.3		43.4		29.20

====================================================

PP (300-5000 µm) ((No. p/kg))

Q110	-  		-  		61.0				CF|AA|HP|µF

Q131	31.7		16.1		119	**			CF|NL| -|µF

Q153	-  		17.0		50.0				AF|AA|HO|FP

Q3889	-  		3205	**	-

Q3892	-  		-  		7.66

Q4001	16.0		-  		135	**			SI|NP|HO|µR|ZB

=================     Summary Statistics     ================

NDA mean	-  		16.27		46.33

NDA st dev	-  		1.31		53.21

Coeff Var (%)	-  		8.0		114.8

N	3		7		10			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PP (300-5000 µm) ((No. p/kg))     (cont.)

Q4018	-  		-  		14.0				CF|NA|HO|µF

Q4020	-  		16.1		-  				GF|NA|HO|AF

Q4026	-  		15.9		7.60				AF|AA|HO|µF|ZC

Q4035	-  		47.6	**	-  				SF|OI| F|µF|ZB

Q4049	-  		-  		57.0				 -|NL|HP|AF

Q703	-  		-  		15.0				SS|NL|HO|µR|ZB

Q920	20.0		40.0	**	60.0

=================     Statistical Results     =================

NDA mean	-  		16.27		46.33

NDA st dev	-  		1.31		53.21

Coeff Var (%)	-  		8.0		114.8

N	3		7		10

Median	20.00		17.00		53.50

MAD	4.00		1.10		39.00

Total Error	-  		4.61		24.25

Category	3		-  		-

====================================================

Polypropylene total ((No. p/kg))

Q101	-  		33.0		-  				PF|NL| -|FM

Q104	158		157		465	**			SI|AA|KN|µF|ZA

Q110	32.0		16.0	*	114				CF|AA|HP|µF

Q122	115		83.0		74.0				GF|NN| -|µF

Q131	31.7		16.1	*	159				CF|NL| -|µF

Q153	300	*	134		425	**			AF|AA|HO|FP

Q3013	-  		-  		53.0

Q3277	175		683	**	597	**			AF|PT|HO|µF

Q3875	312	*	628	**	610	**			SS|AA|KH|LD

Q3878	-  		-  		69.2				SI|NL|HF|µF

Q3882	67.0		-  		160				SS|AA| -|µF|ZC

Q3887	63.0		-  		-  				SS|AA| F|AF|ZD

Q3889	-  		3205	**	-

Q3890	-  		-  		152				ZD|ZC|ZC|µR|ZD

Q3892	32.2		80.1		345	**

Q3991	516	**	-  		545	**			ZB|AA| -|µF|ZB

Q4001	277		161		508	**			SI|NP|HO|µR|ZB

Q4014	158		268	**	91.0				 -| -| -| -|ZC

Q4018	-  		-  		14.0	*			CF|NA|HO|µF

Q4020	-  		16.1	*	-  				GF|NA|HO|AF

Q4026	-  		95.3		30.5	*			AF|AA|HO|µF|ZC

Q4035	-  		47.6		-  				SF|OI| F|µF|ZB

Q4037	-  		-  		26.3	*

Q4040	508	**	200	*	432	**			SS|CC| -|LD|ZA

Q4048	3.00	*	2.00	*	-  				PF|NL| F|µF|ZA

Q4049	16.0	*	63.0		105

Q4057	251		205	*	-

Q703	-  		-  		53.0				SS|NL|HO|µR|ZB

Q871	108		36.0		-

Q920	400	**	320	**	160

=================     Statistical Results     =================

NDA mean	155.3		92.35		156.3

NDA st dev	176.9		108.73		180.1

Coeff Var (%)	113.9		117.7		115.2

N	19		21		22

Median	158.0		95.30		155.8

MAD	125.8		79.17		114.0

Total Error	64.2		37.89		62.2

====================================================

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PS (50-299 µm) ((No. p/kg))

Q101	-  		226		-  				PF|NL| -|FM

Q104	364	**	441	**	84.0				SI|AA|KN|µF|ZA

Q110	32.0		300	*	30.0				CF|AA|HP|µF

Q131	-  		145		-  				CF|NL| -|µF

Q153	-  		483	**	13.0				AF|AA|HO|FP

Q3277	43.0		-  		-  				AF|PT|HO|µF

Q3875	-  		628	**	305	**			SS|AA|KH|LD

Q3878	-  		15.7	*	46.2				SI|NL|HF|µF

Q3890	143	**	15.8	*	-  				ZD|ZC|ZC|µR|ZD

Q3892	32.2		112		115	**

Q3913	86.0	**	492	**	-  				PF|NL|KH|FM

Q3930	977	**	502	**	-  				SF|AA| -|RA|ZE

Q3991	516	**	253		242	**			ZB|AA| -|µF|ZB

Q4001	16.0		533	**	79.0				SI|NP|HO|µR|ZB

Q4014	-  		32.0	*	-  				ZA|NA|ZA|ZA|ZA

Q4026	-  		222		-  				AF|AA|HO|µF|ZC

Q4035	-  		127		-  				SF|OI| F|µF|ZB

Q4037	61.5	*	78.6		47.4				SS|CC| -|FA|ZB

Q4040	15.0		154		8.00				SS|CC| -|LD|ZA

Q4044	15.9		47.5		-  				CF|NA| -|µF

Q4048	1.00	*	21.0	*	-  				PF|NL| F|µF|ZA

Q4049	-  		63.0		8.00				PF|NL|HP|µF

Q4057	31.4		94.7		-

Q703	-  		568	**	91.0	*			SS|NL|HO|µR|ZB

Q871	-  		-  		1.00	*

Q920	-  		100		-

=================     Statistical Results     =================

NDA mean	30.63		162.2		46.12

NDA st dev	31.35		174.6		53.72

Coeff Var (%)	102.3		107.7		116.5

N	14		24		13

Median	37.61		149.4		47.37

MAD	23.27		110.5		39.37

Total Error	13.28		60.6		22.17

====================================================

PS (300-5000 µm) ((No. p/kg))

Q101	-  		33.0		-  				PF|NL| -|FM

Q110	-  		252	**	-  				CF|AA|HP|µF

Q131	-  		80.4		-  				CF|NL| -|µF

Q153	-  		67.0		-  				AF|AA|HO|FP

Q3277	-  		390	**	-

Q3991	-  		127		-  				ZB|AA| -|µF|ZB

Q4001	-  		145		8.00				SI|NP|HO|µR|ZB

Q4026	-  		95.2		-  				AF|AA|HO|µF|ZC

Q4035	-  		206	**	-  				SF|OI| F|µF|ZB

Q4044	-  		111		-  				CF|NA| -|µF

Q4048	-  		2.00	*	-  				PF|NL| F|µF|ZA

Q4049	-  		95.0		-  				 -|NL|HP|AF

Q920	-  		40.0		-

=================     Statistical Results     =================

NDA mean	-  		93.36		-

NDA st dev	-  		70.99		-

Coeff Var (%)	-  		76.0		-

N	-		13		1

Median	-  		95.20		8.000

MAD	-  		49.80		-

Total Error	-  		34.26		-

Category	-  		-  		3

====================================================

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polystyrene total ((No. p/kg))

Q101	-  		259		-  				PF|NL| -|FM

Q104	364	**	441		84.0				SI|AA|KN|µF|ZA

Q110	32.0		552	*	30.0				CF|AA|HP|µF

Q131	-  		225		-  				CF|NL| -|µF

Q153	-  		550	*	13.0				AF|AA|HO|FP

Q3277	43.0		390		-  				AF|PT|HO|µF

Q3875	-  		628	**	305	**			SS|AA|KH|LD

Q3878	-  		15.7	*	46.2				SI|NL|HF|µF

Q3890	143	**	15.8	*	-  				ZD|ZC|ZC|µR|ZD

Q3892	32.2		112		115	**

Q3913	86.0	**	492	*	-  				PF|NL|KH|FM

Q3930	977	**	502	*	-  				SF|AA| -|RA|ZE

Q3991	516	**	380		242	**			ZB|AA| -|µF|ZB

Q4001	16.0		678	**	87.0				SI|NP|HO|µR|ZB

Q4014	-  		32.0	*	-  				 -| -| -| -|ZC

Q4026	-  		317		-  				AF|AA|HO|µF|ZC

Q4035	-  		333		-  				SF|OI| F|µF|ZB

Q4037	61.5	*	78.6	*	47.4

Q4040	15.0		154		8.00				SS|CC| -|LD|ZA

Q4044	15.9		158		-  				CF|NA| -|µF

Q4048	1.00	*	23.0	*	-  				PF|NL| F|µF|ZA

Q4049	-  		158		8.00

Q4057	31.4		94.7		-

Q703	-  		568	**	91.0				SS|NL|HO|µR|ZB

Q871	-  		-  		1.00	*

Q920	-  		140		-

=================     Statistical Results     =================

NDA mean	30.63		278.1		46.89

NDA st dev	31.35		257.5		54.12

Coeff Var (%)	102.3		92.6		115.4

N	14		25		13

Median	37.61		259.0		47.37

MAD	23.27		180.4		39.37

Total Error	13.28		95.1		22.39

====================================================

PVC (50-299 µm) ((No. p/kg))

Q101	350		-  		94.0				PF|NL| -|FM

Q104	-  		-  		670	**			SI|AA|KN|µF|ZA

Q110	-  		-  		129	*			CF|AA|HP|µF

Q131	-  		-  		71.6				CF|NL| -|µF

Q153	-  		-  		75.0				AF|AA|HO|FP

Q3013	-  		-  		8.00	*

Q3875	312		2510		1829	**			SS|AA|KH|LD

Q3878	-  		-  		84.6				SI|NL|HF|µF

Q3887	-  		-  		108				SS|AA| F|AF|ZD

Q3890	-  		-  		76.2				ZD|ZC|ZC|µR|ZD

Q3892	-  		-  		161	**

Q3991	-  		-  		121				ZB|AA| -|µF|ZB

Q4001	-  		48.0		64.0				SI|NP|HO|µR|ZB

Q4018	40.0		20.0		-  				CF|NA|HO|µF

Q4026	-  		-  		30.5				AF|AA|HO|µF|ZC

Q4037	30.8		-  		31.6				SS|CC| -|FA|ZB

Q4040	62.0		15.0		250	**			SS|CC| -|LD|ZA

Q4049	-  		-  		32.0				PF|NL|HP|µF

Q4057	15.7		15.8		-

Q703	-  		-  		46.0				SS|NL|HO|µR|ZB

=================     Summary Statistics     ================

NDA mean	-  		-  		75.36

NDA st dev	-  		-  		59.78

Coeff Var (%)	-  		-  		79.3

N	6		5		18			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PVC (50-299 µm) ((No. p/kg))     (cont.)

=================     Statistical Results     =================

NDA mean	-  		-  		75.36

NDA st dev	-  		-  		59.78

Coeff Var (%)	-  		-  		79.3

N	6		5		18

Median	51.00		20.00		80.42

MAD	27.77		5.00		44.50

Total Error	-  		-  		26.16

Category	2		2		-

====================================================

PVC (300-5000 µm) ((No. p/kg))

Q101	-  		49.0		5.00				PF|NL| -|FM

Q110	-  		-  		122	**			CF|AA|HP|µF

Q131	-  		-  		47.7	*			CF|NL| -|µF

Q153	-  		-  		13.0				AF|AA|HO|FP

Q3892	-  		-  		23.0

Q3991	-  		-  		121	**			ZB|AA| -|µF|ZB

Q4001	-  		-  		48.0	*			SI|NP|HO|µR|ZB

Q4018	-  		-  		14.0				CF|NA|HO|µF

Q4044	-  		-  		7.62				CF|NA| -|µF

Q4049	16.0		-  		81.0	**			 -|NL|HP|AF

Q703	-  		-  		15.0				SS|NL|HO|µR|ZB

=================     Statistical Results     =================

NDA mean	-  		-  		22.94

NDA st dev	-  		-  		26.25

Coeff Var (%)	-  		-  		114.4

N	1		1		11

Median	16.00		49.00		22.99

MAD	-  		-  		17.99

Total Error	-  		-  		11.70

Category	3		3		-

====================================================

Polyvinyl Chloride total ((No. p/kg))

Q101	350		49.0		99.0				PF|NL| -|FM

Q104	-  		-  		670	**			SI|AA|KN|µF|ZA

Q110	-  		-  		251	**			CF|AA|HP|µF

Q131	-  		-  		119				CF|NL| -|µF

Q153	-  		-  		88.0				AF|AA|HO|FP

Q3013	-  		-  		8.00	*

Q3875	312		2510	**	1829	**			SS|AA|KH|LD

Q3878	-  		-  		84.6				SI|NL|HF|µF

Q3887	-  		-  		108				SS|AA| F|AF|ZD

Q3890	-  		-  		76.2				ZD|ZC|ZC|µR|ZD

Q3892	-  		-  		184	*

Q3991	-  		-  		242	**			ZB|AA| -|µF|ZB

Q4001	-  		48.0		112				SI|NP|HO|µR|ZB

Q4018	40.0		20.0		14.0	*			CF|NA|HO|µF

Q4026	-  		-  		30.5				AF|AA|HO|µF|ZC

Q4037	30.8		-  		31.6

Q4040	62.0		15.0		250	**			SS|CC| -|LD|ZA

Q4044	-  		-  		7.62	*			CF|NA| -|µF

Q4049	16.0		-  		113

Q4057	15.7		15.8		-

Q703	-  		-  		61.0				SS|NL|HO|µR|ZB

=================     Summary Statistics     ================

NDA mean	-  		28.45		93.75

NDA st dev	-  		22.44		98.06

Coeff Var (%)	-  		78.9		104.6

N	7		6		20			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polyvinyl Chloride total ((No. p/kg))     (cont.)

=================     Statistical Results     =================

NDA mean	-  		28.45		93.75

NDA st dev	-  		22.44		98.06

Coeff Var (%)	-  		78.9		104.6

N	7		6		20

Median	40.00		34.00		103.50

MAD	24.00		16.60		72.46

Total Error	-  		13.75		36.39

Category	3		-  		-

====================================================

Other polymers (50-299 µm) ((No. p/kg))

Q101	157		113	*	43.0				PF|NL| -|FM

Q110	48.0		-  		61.0				CF|AA|HP|µF

Q122	131		150	**	124				GF|NN| -|µF

Q131	-  		16.1		55.7				CF|NL| -|µF

Q134	-  		-  		10.0				GF|NA|HO|FF

Q153	217		733	**	638	**			AF|AA|HO|FP

Q3013	16.0	*	-  		53.0

Q3277	702	**	1464	**	939	**			AF|PT|HO|µF

Q3807	-  		20.0		-  				SI|AA| F|LD|ZA

Q3875	7182	**	24474	**	11586	**			SS|AA|KH|LD

Q3878	143		63.0		15.4				SI|NL|HF|µF

Q3890	413	**	63.1		152	*			ZD|ZC|ZC|µR|ZD

Q3892	16.1	*	16.0		874	**

Q3913	333	**	925	**	-  				PF|NL|KH|FM

Q3991	-  		-  		242	**			ZB|AA| -|µF|ZB

Q4014	32.0		16.0		15.0				ZA|NA|ZA|ZA|ZA

Q4018	40.0		20.0		-  				CF|NA|HO|µF

Q4035	192		95.2		-  				SF|OI| F|µF|ZB

Q4044	-  		-  		7.62				CF|NA| -|µF

Q4048	245		79.0		-  				PF|NL| F|µF|ZA

Q4049	32.0		47.0		275	**			PF|NL|HP|µF

Q871	-  		-  		15318	**

Q920	220		60.0		20.0

=================     Statistical Results     =================

NDA mean	138.8		54.64		63.82

NDA st dev	158.5		60.84		105.95

Coeff Var (%)	114.2		111.4		166.0

N	17		17		18

Median	157.0		63.05		92.50

MAD	117.0		46.95		79.99

Total Error	59.6		23.25		35.29

====================================================

Other polymers (300-5000µm) ((No. p/kg))

Q101	257		97.0		38.0				PF|NL| -|FM

Q110	16.0	*	-  		16.0				CF|AA|HP|µF

Q130	-  		-  		8.00				SS|ZC|UR|AF|ZB

Q131	-  		-  		159	**			CF|NL| -|µF

Q134	-  		16.7		-  				GF|NA|HO|FF

Q153	-  		-  		63.0				AF|AA|HO|FP

Q3013	-  		-  		15.0

Q3807	140		-  		-  				SI|AA| F|LD|ZA

Q3875	-  		-  		1220	**			SS|AA|KH|LD

Q3889	463	**	732	**	202	**			 -|ZA|µF

Q3892	-  		32.0		53.6

Q3913	199		82.0		-  				PF|NL|KH|FM

Q4020	46.9		16.1		23.1				GF|NA|HO|AF

=================     Summary Statistics     ================

NDA mean	161.2		47.00		40.27

NDA st dev	95.7		51.71		52.89

Coeff Var (%)	59.4		110.0		131.3

N	7		8		13			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Other polymers (300-5000µm) ((No. p/kg))     (cont.)

Q4035	208		79.4		17.4				SF|OI| F|µF|ZB

Q4049	-  		16.0		97.0	*			 -|NL|HP|AF

Q871	-  		-  		360	**

=================     Statistical Results     =================

NDA mean	161.2		47.00		40.27

NDA st dev	95.7		51.71		52.89

Coeff Var (%)	59.4		110.0		131.3

N	7		8		13

Median	199.0		55.69		53.65

MAD	59.0		39.29		38.65

Total Error	60.9		25.93		21.16

====================================================

Other polymers total ((No. p/kg))

Q101	414	**	210	**	81.0				PF|NL| -|FM

Q110	64.0		-  		77.0				CF|AA|HP|µF

Q122	131		150	**	124				GF|NN| -|µF

Q131	-  		16.1		215	**			CF|NL| -|µF

Q134	-  		16.7		10.0				GF|NA|HO|FF

Q153	217		733	**	701	**			AF|AA|HO|FP

Q3013	-  		-  		69.0

Q3277	702	**	1464	**	939	**			AF|PT|HO|µF

Q3807	140		20.0		-  				SI|AA| F|LD|ZA

Q3875	7182	**	24474	**	12806	**			SS|AA|KH|LD

Q3878	143		63.0		15.4				SI|NL|HF|µF

Q3882	-  		-  		40.0				SS|AA| -|µF|ZC

Q3889	463	**	732	**	202	**			ZC| -|ZA|µF

Q3890	413	**	63.1		152				ZD|ZC|ZC|µR|ZD

Q3892	16.1	*	48.0		927	**

Q3913	532	**	1007	**	-  				PF|NL|KH|FM

Q3991	-  		-  		242	**			ZB|AA| -|µF|ZB

Q4014	32.0	*	16.0		15.0				 -| -| -| -|ZC

Q4018	40.0		20.0		-  				CF|NA|HO|µF

Q4020	46.9		16.1		23.1				GF|NA|HO|AF

Q4035	400	*	175	**	17.4				SF|OI| F|µF|ZB

Q4044	-  		-  		7.62				CF|NA| -|µF

Q4048	245		79.0		-  				PF|NL| F|µF|ZA

Q4049	32.0	*	63.0		372	**

Q871	90.0		-  		15678	**

Q920	220		60.0		20.0

=================     Statistical Results     =================

NDA mean	185.1		51.08		77.97

NDA st dev	207.1		63.14		120.23

Coeff Var (%)	111.9		123.6		154.2

N	20		20		22

Median	179.9		63.03		102.50

MAD	143.9		46.91		89.99

Total Error	74.4		22.08		37.77

====================================================

Total polymers (50-299 µm) ((No. p/kg))

Q101	720	*	663	*	245				PF|NL| -|FM

Q102	32.4	*	332		562				GF|NB|AL|MI

Q104	808	**	835	**	6872	**			SI|AA|KN|µF|ZA

Q110	160		427		1461	**			CF|AA|HP|µF

Q122	410		333		306				GF|NN| -|µF

Q131	-  		225		247				CF|NL| -|µF

Q134	33.6	*	16.7	*	20.0	*			GF|NA|HO|FF

=================     Summary Statistics     ================

NDA mean	367.2		366.8		462.6

NDA st dev	487.5		430.9		681.0

Coeff Var (%)	132.7		117.5		147.2

N	31		35		31			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Total polymers (50-299 µm) ((No. p/kg))     (cont.)

Q153	600		1683	**	2001	**			AF|AA|HO|FP

Q3013	64.0	*	-  		541

Q3211	-  		64.0	*	182				SF|NA| -|MI

Q3277	1665	**	3709	**	2603	**			AF|PT|HO|µF

Q3807	-  		20.0	*	-  				SI|AA| F|LD|ZA

Q3875	10616	**	34515	**	20733	**			SS|AA|KH|LD

Q3878	317		110		562				SI|NL|HF|µF

Q3890	747	*	94.6	*	457

Q3892	145		352		2269	**

Q3913	758	*	1636	**	379				PF|NL|KH|FM

Q3930	1173	**	903	**	-  				SF|AA| F|RA|ZG

Q3940	1700	**	1700	**	4100	**			CF|NA| F|MI|ZA

Q3946a	528		918	**	293				GF|NA| -| -|ZA

Q3991	1548	**	506		2240	**			ZB|AA| -|µF|ZB

Q4001	277		1051	**	1065	**			SI|NP|HO|µR|ZB

Q4014	206		363		175				ZA|NA|ZA|ZA|ZA

Q4018	120		80.0	*	-  				CF|NA|HO|µF

Q4023	32000	**	15000	**	164000	**			PF|NL|HP|MI

Q4026	-  		302		53.4	*			AF|AA|HO|µF|ZC

Q4029	-  		-  		47.7	*			 -| -| -|HN|ZB

Q4035	240		349		-  				SF|OI| F|µF|ZB

Q4037	185		143		168				SS|CC| -|FA|ZB

Q4040	1015	**	692	*	1288	**			SS|CC| -|LD|ZA

Q4044	15.9	*	158		30.5	*			CF|NA| -|µF

Q4048	269		126		-  				PF|NL| F|µF|ZA

Q4049	95.0		174		736

Q4057	2446	**	2084	**	-

Q703	-  		1183	**	335				SS|NL|HO|µR|ZB

Q871	18.0	*	36.0	*	15769	**

Q920	2220	**	880	**	620

=================     Statistical Results     =================

NDA mean	367.2		366.8		462.6

NDA st dev	487.5		430.9		681.0

Coeff Var (%)	132.7		117.5		147.2

N	31		35		31

Median	410.0		363.0		561.5

MAD	346.0		300.0		503.5

Total Error	143.2		129.6		192.0

====================================================

Total polymers (300-5000 µm) ((No. p/kg))

Q101	900	**	710	**	137				PF|NL| -|FM

Q102	-  		253		677	**			GF|NB|AL|MI

Q104	-  		-  		84.0				SI|AA|KN|µF|ZA

Q110	16.0		378	*	556	**			CF|AA|HP|µF

Q130	-  		-  		99.0				SS|ZC|UR|AF|ZB

Q131	-  		193		557	**			CF|NL| -|µF

Q134	-  		33.3	*	40.1	*			GF|NA|HO|FF

Q153	-  		117		252	*			AF|AA|HO|FP

Q3013	-  		-  		114

Q3211	-  		80.0		474	**			SF|NA| -|MI

Q3277	-  		390	*	68.0				AF|PT|HO|µF

Q3807	140		-  		160				SI|AA| F|LD|ZA

Q3875	625	**	314		1525	**			SS|AA|KH|LD

Q3878	-  		-  		192				SI|NL|HF|µF

Q3889	8138	**	9066	**	13203	**

Q3890	-  		-  		91.5

Q3892	32.2		32.0	*	130

Q3913	199		82.0		23.0	*			PF|NL|KH|FM

=================     Summary Statistics     ================

NDA mean	90.36		187.0		148.2

NDA st dev	173.01		203.3		150.2

Coeff Var (%)	191.5		108.7		101.3

N	17		26		38			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Total polymers (300-5000 µm) ((No. p/kg))     (cont.)

Q3930	-  		-  		249	*			SF|AA| F|RA|ZH

Q3936	-  		-  		69.0				SS|NA|HP|MI

Q3940	300	**	300		400	**			CF|NA| F|MI|ZA

Q3946a	528	**	734	**	147				GF|NA| -| -|ZA

Q3991	-  		380	*	374	**			ZB|AA| -|µF|ZB

Q4001	16.0		275		604	**			SI|NP|HO|µR|ZB

Q4014	-  		16.0	*	76.0				CF|NA| -|AF|ZB

Q4018	20.0		20.0	*	157				CF|NA|HO|µF

Q4020	46.9		64.5		69.2				GF|NA|HO|AF

Q4023	8000	**	5400	**	2000	**			PF|NL|HP|MI

Q4026	-  		111		22.9	*			AF|AA|HO|µF|ZC

Q4035	320	**	460	**	104				SF|OI| F|µF|ZB

Q4037	-  		-  		42.1	*			SS|CC| -|AF

Q4040	-  		-  		76.0				SS|CC| -|LD|ZA

Q4044	-  		190		45.7	*			CF|NA| -|µF

Q4047	-  		-  		320	**			GF|AA|HO|MI

Q4048	1.00		10.0	*	-  				PF|NL| F|µF|ZA

Q4049	16.0		174		429	**

Q703	-  		-  		137				SS|NL|HO|µR|ZB

Q871	-  		-  		360	**

Q920	20.0		200		280	*

=================     Statistical Results     =================

NDA mean	90.36		187.0		148.2

NDA st dev	173.01		203.3		150.2

Coeff Var (%)	191.5		108.7		101.3

N	17		26		38

Median	140.00		196.6		152.0

MAD	124.00		147.6		103.1

Total Error	57.31		68.7		48.4

====================================================

Total polymers (50-5000 µm) ((No. p/kg))

Q101	1620	**	1373	**	382				PF|NL| -|FM

Q102	32.4	*	585		1238	*			GF|NB|AL|MI

Q104	808		835		6956	**			SI|AA|KN|µF|ZA

Q110	176		805		2017	**			CF|AA|HP|µF

Q122	410		333		406				GF|NN| -|µF

Q131	-  		418		803				CF|NL| -|µF

Q134	33.6	*	50.0	*	60.1	*			GF|NA|HO|FF

Q153	600		1800	**	2253	**			AF|AA|HO|FP

Q3013	-  		-  		657

Q3277	1665	**	4099	**	3056	**			AF|PT|HO|µF

Q3807	140		20.0	*	160				SI|AA| F|LD|ZA

Q3875	11241	**	34829	**	22258	**			SS|AA|KH|LD

Q3878	317		110	*	754				SI|NL|HF|µF

Q3882	134	*	-  		724				SS|AA| -|µF|ZC

Q3890	747		94.6	*	549

Q3892	177		384		2399	**

Q3913	957	*	1717	**	403				PF|NL|KH|FM

Q3926	194341	**	339534	**	81567	**			ZB|NA|KH|µF

Q3930	1173	**	903		249				SF|AA| F|RA|ZH

Q3932	-  		-  		11.0	*

Q3935	24341	**	15003	**	8330	**			SS|NA|HO|MI

Q3940	2000	**	2000	**	4500	**			CF|NA| F|MI|ZA

Q3946a	1057	*	1652	**	440				GF|NA| -| -|ZA

Q3991	1548	**	886		2615	**			ZB|AA| -|µF|ZB

Q4001	293		1326	*	1670	**			SI|NP|HO|µR|ZB

Q4014	206		379		252				 -| -| -| -|ZC

Q4018	140		100	*	157				CF|NA|HO|µF

=================     Summary Statistics     ================

NDA mean	521.7		658.4		595.6

NDA st dev	675.6		770.8		857.7

Coeff Var (%)	129.5		117.1		144.0

N	35		37		38			(cont.)

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Total polymers (50-5000 µm) ((No. p/kg))     (cont.)

Q4020	46.9	*	64.5	*	69.2	*			GF|NA|HO|AF

Q4023	40000	**	20400	**	166000	**			PF|NL|HP|MI

Q4026	-  		413		76.3	*			AF|AA|HO|µF|ZC

Q4035	560		809		104	*			SF|OI| F|µF|ZB

Q4037	185		143	*	211

Q4040	1015	*	692		1364	**			SS|CC| -|LD|ZA

Q4044	15.9	*	349		76.2	*			CF|NA| -|µF

Q4048	270		136	*	-  				PF|NL| F|µF|ZA

Q4049	110	*	347		1165	*

Q4057	2446	**	2084	**	-

Q703	-  		1183	*	472				SS|NL|HO|µR|ZB

Q871	1440	**	684		16129	**

Q920	2240	**	1080		900

=================     Statistical Results     =================

NDA mean	521.7		658.4		595.6

NDA st dev	675.6		770.8		857.7

Coeff Var (%)	129.5		117.1		144.0

N	35		37		38

Median	600.0		805.0		738.9

MAD	466.0		568.0		629.9

Total Error	193.7		228.8		229.3

====================================================

Total mass of plastic particles

Analysis DE17  2022.2

| Sample/ Determinand | Assigned Value | Units | Total Error | NDA st.dev | NDA rel. st.dev (%) | Nobs numerical | Nobs LCV | Median | MAD | Model Mean | Uncer- tainty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QMP013MS |  |  |  |  |  |  |  |  |  |  |  |
| PC (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 1 | 0.651 | 0.649 |  |  |
| PC (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 0.400 |  |  |  |
| Polycarbonate total |  | (mg/kg) |  |  |  | 3 | 0 | 0.010 | 0.009 |  |  |
| PE (50-299 µm) |  | (mg/kg) |  |  |  | 0 | 3 |  |  |  |  |
| PE (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 2 |  |  |  |  |
| Polyethylene total (HD+LD) |  | (mg/kg) |  |  |  | 2 | 3 | 0.700 | 0.300 |  |  |
| PET (50-299 µm) |  | (mg/kg) |  |  |  | 3 | 2 | 0.224 | 0.206 |  |  |
| PET (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 2 |  |  |  |  |
| Polyethylene Terephat. Total |  | (mg/kg) |  |  |  | 4 | 2 | 0.262 | 0.141 |  |  |
| PP (50-299 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 0.009 |  |  |  |
| PP (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Polypropylene total |  | (mg/kg) |  |  |  | 3 | 1 | 0.320 | 0.311 |  |  |
| PS (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 2 | 2.96 | 2.95 |  |  |
| PS (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Polystyrene total |  | (mg/kg) |  |  |  | 2 | 3 | 2.96 | 2.95 |  |  |
| PVC (50-299 µm) |  | (mg/kg) |  |  |  | 3 | 2 | 0.032 | 0.013 |  |  |
| PVC (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 2 |  |  |  |  |
| Polyvinyl Chloride total |  | (mg/kg) |  |  |  | 5 | 2 | 0.280 | 0.248 |  |  |
| Other polymers (50-299 µm) |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Other polymers total |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Total Polymers (50-299 µm) |  | (mg/kg) |  |  |  | 4 | 0 | 0.788 | 0.631 |  |  |
| Total Polymers (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 0 | 0.400 |  |  |  |
| Total Polymers (50-5000µm) |  | (mg/kg) |  |  |  | 6 | 0 | 1.35 | 1.19 |  |  |

| Sample/ Determinand | Assigned Value | Units | Total Error | NDA st.dev | NDA rel. st.dev (%) | Nobs numerical | Nobs LCV | Median | MAD | Model Mean | Uncer- tainty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QMP014MS |  |  |  |  |  |  |  |  |  |  |  |
| PC (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 1 | 0.551 | 0.549 |  |  |
| PC (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 0.900 |  |  |  |
| Polycarbonate total |  | (mg/kg) |  |  |  | 3 | 0 | 0.010 | 0.009 |  |  |
| PE (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 2 | 0.205 | 0.119 |  |  |
| PE (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 1.10 |  |  |  |
| Polyethylene total (HD+LD) |  | (mg/kg) |  |  |  | 6 | 0 | 0.712 | 0.629 |  |  |
| PET (50-299 µm) |  | (mg/kg) |  |  |  | 1 | 3 | 6.93 |  |  |  |
| PET (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 2 |  |  |  |  |
| Polyethylene Terephat. Total |  | (mg/kg) |  |  |  | 2 | 3 | 3.79 | 3.15 |  |  |
| PP (50-299 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 0.008 |  |  |  |
| PP (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Polypropylene total |  | (mg/kg) |  |  |  | 3 | 1 | 0.090 | 0.082 |  |  |
| PS (50-299 µm) |  | (mg/kg) |  |  |  | 3 | 2 | 0.616 | 0.090 |  |  |
| PS (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 0.500 |  |  |  |
| Polystyrene total |  | (mg/kg) |  |  |  | 7 | 1 | 0.526 | 0.116 |  |  |
| PVC (50-299 µm) |  | (mg/kg) |  |  |  | 3 | 2 | 0.377 | 0.193 |  |  |
| PVC (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 2 |  |  |  |  |
| Polyvinyl Chloride total |  | (mg/kg) |  |  |  | 5 | 2 | 0.377 | 0.193 |  |  |
| Other polymers (50-299 µm) |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Other polymers total |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Total Polymers (50-299 µm) |  | (mg/kg) |  |  |  | 4 | 0 | 1.21 | 0.30 |  |  |
| Total Polymers (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 0 | 2.50 |  |  |  |
| Total Polymers (50-5000µm) |  | (mg/kg) |  |  |  | 7 | 0 | 3.60 | 2.89 |  |  |

| Sample/ Determinand | Assigned Value | Units | Total Error | NDA st.dev | NDA rel. st.dev (%) | Nobs numerical | Nobs LCV | Median | MAD | Model Mean | Uncer- tainty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QMP015MS |  |  |  |  |  |  |  |  |  |  |  |
| PC (50-299 µm) |  | (mg/kg) |  |  |  | 3 | 1 | 0.456 | 0.200 |  |  |
| PC (300-5000 µm) |  | (mg/kg) |  |  |  | 1 | 1 | 4.40 |  |  |  |
| Polycarbonate total |  | (mg/kg) |  |  |  | 5 | 0 | 0.420 | 0.164 |  |  |
| PE (50-299 µm) |  | (mg/kg) |  |  |  | 3 | 1 | 1.00 | 0.46 |  |  |
| PE (300-5000 µm) | 27.4 | (mg/kg) | 8.5 | 6.7 | 24.6 | 4 | 1 | 27.3 | 4.8 | 27.35 | 4.21 |
| Polyethylene total (HD+LD) | 19.9 | (mg/kg) | 8.7 | 15.3 | 77.0 | 8 | 0 | 22.6 | 11.5 | 19.92 | 6.78 |
| PET (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 2 | 0.242 | 0.229 |  |  |
| PET (300-5000 µm) |  | (mg/kg) |  |  |  | 2 | 1 | 1.09 | 0.52 |  |  |
| Polyethylene Terephat. Total |  | (mg/kg) |  |  |  | 5 | 2 | 0.570 | 0.557 |  |  |
| PP (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 1 | 0.339 | 0.199 |  |  |
| PP (300-5000 µm) |  | (mg/kg) |  |  |  | 0 | 1 |  |  |  |  |
| Polypropylene total |  | (mg/kg) |  |  |  | 5 | 1 | 0.210 | 0.070 |  |  |
| PS (50-299 µm) |  | (mg/kg) |  |  |  | 4 | 1 | 0.152 | 0.149 |  |  |
| PS (300-5000 µm) |  | (mg/kg) |  |  |  | 2 | 1 | 1.22 | 0.59 |  |  |
| Polystyrene total |  | (mg/kg) |  |  |  | 6 | 2 | 1.14 | 1.13 |  |  |
| PVC (50-299 µm) | 1.16 | (mg/kg) | 0.97 | 0.89 | 77.1 | 4 | 1 | 1.26 | 0.57 | 1.157 | 0.558 |
| PVC (300-5000 µm) |  | (mg/kg) |  |  |  | 2 | 1 | 12.2 | 12.0 |  |  |
| Polyvinyl Chloride total | 2.29 | (mg/kg) | 1.42 | 1.65 | 72.1 | 5 | 2 | 3.00 | 1.19 | 2.293 | 0.924 |
| Other polymers (50-299 µm) |  | (mg/kg) |  |  |  | 2 | 0 | 0.141 | 0.104 |  |  |
| Other polymers total |  | (mg/kg) |  |  |  | 2 | 0 | 0.141 | 0.104 |  |  |
| Total Polymers (50-299 µm) |  | (mg/kg) |  |  |  | 4 | 0 | 4.85 | 2.23 |  |  |
| Total Polymers (300-5000 µm) | 47.2 | (mg/kg) | 21.0 | 27.3 | 57.9 | 4 | 0 | 48.2 | 19.2 | 47.19 | 17.06 |
| Total Polymers (50-5000µm) |  | (mg/kg) |  |  |  | 7 | 0 | 34.1 | 18.3 |  |  |

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PC (50-299 µm) ((mg/kg))

Q4027	0.002		0.001		0.256				ZB|NB|HE|PY|ZA

Q4028	-  		-  		0.456				GF|NB|HE|PY

Q4032	0.009	< 	0.009	< 	0.009	<

Q4039	1.30		1.10		1.90				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	2		2		3

Median	0.6508		0.5507		0.4560

MAD	0.6493		0.5493		0.2000

Category	3		3		3

====================================================

PC (300-5000 µm) ((mg/kg))

Q4032	0.009	< 	0.009	< 	0.009	<

Q4039	0.400		0.900		4.40				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	1		1		1

Median	0.4000		0.9000		4.400

MAD	-  		-  		-

Category	3		3		3

====================================================

Polycarbonate total ((mg/kg))

Q4026-2	-  		-  		0.420				QF|AA|HO|PY

Q4027	0.002		0.001		0.256				ZB|NB|HE|PY|ZA

Q4028	-  		-  		0.456				GF|NB|HE|PY

Q4039	1.70		2.00		6.30				QF|AA| -|PY

Q980	0.010		0.010		0.070				 -| -|AS|PY|ZC

=============     Statistical Results (no NDA)     =============

N	3		3		5

Median	0.0100		0.0100		0.4200

MAD	0.0085		0.0086		0.1640

Category	3		3		4

====================================================

PE (50-299 µm) ((mg/kg))

Q4027	0.002	< 	0.086		0.217				ZB|NB|HE|PY|ZA

Q4028	-  		0.324		1.46				GF|NB|HE|PY

Q4032	0.020	< 	0.020	< 	0.020	<

Q4039	0.100	< 	0.100	< 	1.00				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	-		2		3

Median	-  		0.2050		1.000

MAD	-  		0.1190		0.458

Category	3		3		3

====================================================

PE (300-5000 µm) ((mg/kg))

Q114	-  		-  		21.0				 -| -|AT|PY

Q4027	-  		-  		24.0				ZB|NB|HE|PY|ZA

Q4028	-  		-  		30.7				GF|NB|HE|PY

Q4032	0.020	< 	0.020	< 	0.020	< I

Q4039	0.100	< 	1.10		33.7				QF|AA| -|PY

=================     Statistical Results     =================

NDA mean	-  		-  		27.35

NDA st dev	-  		-  		6.74

Coeff Var (%)	-  		-  		24.6

N	-		1		4

Median	-  		1.100		27.35

MAD	-  		-  		4.81

Total Error	-  		-  		8.46

Category	3		3		-

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polyethylene total (HD+LD) ((mg/kg))

H221	-  		-  		11.3				 -| -|AM|PY|ZB

Q114	-  		-  		21.0				 -| -|AT|PY

Q3175	1.70	< 	6.93		6.02				QF|AA|HO|TD|ZC

Q4026-2	0.400		0.080		11.0				QF|AA|HO|PY

Q4027	0.002	< 	0.086		24.3				ZB|NB|HE|PY|ZA

Q4028	-  		0.324		32.1				GF|NB|HE|PY

Q4039	0.100	< 	1.10		34.7				QF|AA| -|PY

Q980	1.00		3.50		92.0	**			 -| -|AS|PY|ZC

=================     Statistical Results     =================

NDA mean	-  		-  		19.92

NDA st dev	-  		-  		15.33

Coeff Var (%)	-  		-  		77.0

N	2		6		8

Median	0.7000		0.7120		22.65

MAD	0.3000		0.6290		11.54

Total Error	-  		-  		8.71

Category	3		4		-

====================================================

PET (50-299 µm) ((mg/kg))

Q114	10.6		6.93		-  				 -| -|AT|PY

Q4027	0.224		0.000	< 	0.470				ZB|NB|HE|PY|ZA

Q4028	0.018		-  		0.013				GF|NB|HE|PY

Q4032	0.000	< 	0.000	< 	0.000	<

Q4039	0.100	< 	0.100	< 	0.100	< 			QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	3		1		2

Median	0.2240		6.930		0.2415

MAD	0.2060		-  		0.2285

Category	3		3		3

====================================================

PET (300-5000 µm) ((mg/kg))

Q114	-  		-  		0.570				 -| -|AT|PY

Q4032	0.000	< 	0.000	< 	0.000	<

Q4039	0.100	< 	0.100	< 	1.60				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	-		-		2

Median	-  		-  		1.085

MAD	-  		-  		0.515

Category	3		3		3

====================================================

Polyethylene Terephat. Total ((mg/kg))

H221	-  		-  		2.17				 -| -|AM|PY|ZB

Q114	10.6		6.93		0.570				 -| -|AT|PY

Q3175	6.70	< 	6.70	< 	3.10	< 			QF|AA|HO|TD|ZC

Q4027	0.224		0.000	< 	0.470				ZB|NB|HE|PY|ZA

Q4028	0.018		-  		0.013				GF|NB|HE|PY

Q4039	0.100	< 	0.100	< 	1.60				QF|AA| -|PY

Q980	0.300		0.640		0.300	< 			 -| -|AS|PY|ZC

=============     Statistical Results (no NDA)     =============

N	4		2		5

Median	0.2620		3.785		0.5700

MAD	0.1410		3.145		0.5570

Category	4		3		4

====================================================

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

PP (50-299 µm) ((mg/kg))

Q4027	0.009		0.008		0.140				ZB|NB|HE|PY|ZA

Q4028	-  		-  		0.538				GF|NB|HE|PY

Q4032	0.040	< 	0.040	< 	0.040	<

=============     Statistical Results (no NDA)     =============

N	1		1		2

Median	0.0090		0.0080		0.3390

MAD	-  		-  		0.1990

Category	3		3		3

====================================================

PP (300-5000 µm) ((mg/kg))

Q4032	0.040	< 	0.040	< 	0.040	<

=============     Statistical Results (no NDA)     =============

N	-		-		-

Median	-  		-  		-

MAD	-  		-  		-

Category	3		3		3

====================================================

Polypropylene total ((mg/kg))

H221	-  		-  		0.203				 -| -|AM|PY|ZB

Q3175	1.70	< 	1.70	< 	0.800	< 			QF|AA|HO|TD|ZC

Q4026-2	0.320		0.090		0.210				QF|AA|HO|PY

Q4027	0.009		0.008		0.140				ZB|NB|HE|PY|ZA

Q4028	-  		-  		0.538				GF|NB|HE|PY

Q980	1.20		0.300		0.500				 -| -|AS|PY|ZC

=============     Statistical Results (no NDA)     =============

N	3		3		5

Median	0.3200		0.0900		0.2100

MAD	0.3110		0.0820		0.0700

Category	3		3		4

====================================================

PS (50-299 µm) ((mg/kg))

Q114	5.91		4.78		10.1				 -| -|AT|PY

Q4027	0.010		0.526		0.002				ZB|NB|HE|PY|ZA

Q4028	-  		0.616		0.003				GF|NB|HE|PY

Q4032	0.000	< 	0.000	< 	0.000	<

Q4039	0.100	< 	0.100	< 	0.300				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	2		3		4

Median	2.960		0.6160		0.1515

MAD	2.950		0.0900		0.1490

Category	3		3		4

====================================================

PS (300-5000 µm) ((mg/kg))

Q114	-  		-  		0.630				 -| -|AT|PY

Q4032	0.000	< 	0.000	< 	0.000	<

Q4039	-  		0.500		1.80				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	-		1		2

Median	-  		0.5000		1.215

MAD	-  		-  		0.585

Category	3		3		3

====================================================

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Polystyrene total ((mg/kg))

H221	-  		10.0		5.94				 -| -|AM|PY|ZB

Q114	5.91		4.78		10.7				 -| -|AT|PY

Q3175	1.70	< 	1.70	< 	0.800	< 			QF|AA|HO|TD|ZC

Q4026-2	-  		0.410		0.170				QF|AA|HO|PY

Q4027	0.010		0.526		0.002				ZB|NB|HE|PY|ZA

Q4028	-  		0.616		0.003				GF|NB|HE|PY

Q4039	0.100	< 	0.500		2.10				QF|AA| -|PY

Q980	0.010	< 	0.200		0.010	< 			 -| -|AS|PY|ZC

=============     Statistical Results (no NDA)     =============

N	2		7		6

Median	2.960		0.5260		1.135

MAD	2.950		0.1160		1.133

Category	3		4		4

====================================================

PVC (50-299 µm) ((mg/kg))

Q114	0.520		0.570		1.65				 -| -|AT|CD

Q4027	0.032		0.093		0.515				ZB|NB|HE|PY|ZA

Q4028	0.019		0.377		0.879				GF|NB|HE|PY

Q4032	0.200	< 	0.200	< 	0.200	< C

Q4039	1.00	< 	1.00	< 	2.90				QF|AA| -|PY

=================     Statistical Results     =================

NDA mean	-  		-  		1.157

NDA st dev	-  		-  		0.892

Coeff Var (%)	-  		-  		77.1

N	3		3		4

Median	0.0320		0.3770		1.265

MAD	0.0130		0.1930		0.568

Total Error	-  		-  		0.966

Category	3		3		-

====================================================

PVC (300-5000 µm) ((mg/kg))

Q114	-  		-  		0.160				 -| -|AT|CD

Q4032	0.200	< 	0.200	< 	0.200	<

Q4039	1.00	< 	1.00	< 	24.2				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	-		-		2

Median	-  		-  		12.18

MAD	-  		-  		12.02

Category	3		3		3

====================================================

Polyvinyl Chloride total ((mg/kg))

Q114	0.520		0.570		1.81				 -| -|AT|CD

Q3175	11.7	< 	11.7	< 	5.40	< C			QF|AA|HO|TD|ZC

Q4026-2	0.280		0.250		4.03				QF|AA|HO|PY

Q4027	0.032		0.093		0.515	< C			ZB|NB|HE|PY|ZA

Q4028	0.019		0.377		0.879				GF|NB|HE|PY

Q4039	1.00	< 	1.00	< 	27.1	**			QF|AA| -|PY

Q980	1.40		23.0		3.00				 -| -|AS|PY|ZC

=================     Statistical Results     =================

NDA mean	-  		-  		2.293

NDA st dev	-  		-  		1.654

Coeff Var (%)	-  		-  		72.1

N	5		5		5

Median	0.2800		0.3770		3.000

MAD	0.2480		0.1930		1.190

Total Error	-  		-  		1.417

Category	4		4		-

====================================================

Sample	QMP013MS	QMP014MS	QMP015MS		MIC

Other polymers (50-299 µm) ((mg/kg))

Q4027	0.002	< 	0.002	< 	0.037				ZB|NB|HE|PY|ZA

Q4028	-  		-  		0.245				GF|NB|HE|PY

=============     Statistical Results (no NDA)     =============

N	-		-		2

Median	-  		-  		0.1410

MAD	-  		-  		0.1040

Category	3		3		3

====================================================

Other polymers total ((mg/kg))

Q4027	0.002	< 	0.002	< 	0.037				ZB|NB|HE|PY|ZA

Q4028	-  		-  		0.245				GF|NB|HE|PY

=============     Statistical Results (no NDA)     =============

N	-		-		2

Median	-  		-  		0.1410

MAD	-  		-  		0.1040

Category	3		3		3

====================================================

Total Polymers (50-299 µm) ((mg/kg))

Q114	17.0		12.3		11.7				 -| -|AT|PY

Q4027	0.276		0.714		1.64				ZB|NB|HE|PY|ZA

Q4028	0.038		1.32		3.59				GF|NB|HE|PY

Q4039	1.30		1.10		6.10				QF|AA| -|PY

=============     Statistical Results (no NDA)     =============

N	4		4		4

Median	0.7880		1.209		4.847

MAD	0.6310		0.302		2.230

Category	4		2		2

====================================================

Total Polymers (300-5000 µm) ((mg/kg))

Q114	-  		-  		22.4				 -| -|AT|PY

Q3936	-  		-  		69.0				SS| -| -|MI

Q4028	-  		-  		30.7				GF|NB|HE|PY

Q4039	0.400		2.50		65.7				QF|AA| -|PY

=================     Statistical Results     =================

NDA mean	-  		-  		47.19

NDA st dev	-  		-  		27.30

Coeff Var (%)	-  		-  		57.9

N	1		1		4

Median	0.4000		2.500		48.18

MAD	-  		-  		19.17

Total Error	-  		-  		21.03

Category	3		3		-

====================================================

Total Polymers (50-5000µm) ((mg/kg))

Q114	17.0		12.3		34.1				 -| -|AT|PY

Q3175	-  		6.93		6.02				QF|AA|HO|TD|ZC

Q4026-2	1.00		0.830		15.8				QF|AA|HO|PY

Q4027	0.276		0.714		25.2				ZB|NB|HE|PY|ZA

Q4028	0.038		1.32		34.3				GF|NB|HE|PY

Q4039	1.70		3.60		71.8				QF|AA| -|PY

Q980	4.00		28.0		95.0				 -| -|AS|PY|ZC

=============     Statistical Results (no NDA)     =============

N	6		7		7

Median	1.350		3.600		34.13

MAD	1.193		2.886		18.34

Category	4		4		4

====================================================

CONTINUOUS LABORATORY-PERFORMANCE STUDIES ORGANIZED BY

|  | Microplastics Fee € 0.001,- (EUR) per year In this round 93 participants |
| --- | --- |

For more information and application, please contact:

WEPAL

PO BOX 8005

6700 EC  WAGENINGEN

THE NETHERLANDS

Tel.	:	+31 317 48 23 37 	no reply :  +31 317 48 3643

Fax.	:	+31 317 48 56 66

E-mail	:	wepalquasimeme@wur.nl

Internet:	www.wepal.nl

International Harmonised Protocol for Proficiency Testing of (Chemical) Analytical Laboratories.  M Thompson, R Wood, Journal of AOAC International Vol. 76, No. 4, 1993