Jan 11, 2011 (marinova)

Hobo Redeployment – University Valley Dec 2010

- All Hobo stations (shallow, medium, deep ice) were downloaded; all the data looks good.

- All batteries were replaced at the shallow ice site with no problems or difficulties. All sensors were relaunched to record every 1 hr, but not to record their battery voltage; this allowed us to keep to same recording inverval, but to have them last for > 2 years.

- At the medium ice site, the T/RH sensor under the rock stopped functioning after the battery was changed (showed -94C). This is likely due to a disconnect between the sensor and the circuit board, but this was not discovered at the time. The sensor will likely not record any useful data from that point on. The 4-channel external industrial refused to connect to the computer after the battery was changed. The reason for this is completely unknown. All sensors (that could be connected to) were relaunched with a 1 hr measuring interval and no recording of the battery voltage.

- At the deep ice site, the 4-channel external industrial sensors had the same failure mode as that at the medium ice site. All other sensors functioned fine after the battery change. The sensors were relaunched to a 1 hr interval and no battery voltage recording.