Grant Agreement No: 101004761

AIDAinnova

Advancement and Innovation for Detectors at Accelerators

Horizon 2020 Research Infrastructures project AIDAINNOVA

deliverable report

Evaluate Non-Ionizing Energy Loss

(NIEL) of irradiation facilities with

dedicated dosimeter structures

Deliverable: D4.2

| Document identifier: | AIDAinnova._Deliverable_D4.2 |
| --- | --- |
| Due date of deliverable: | End of Month 42 (30.11.2024) |
| Justification for delay: | Production of dosimetry diodes took longer than anticipated. The date for the deliverable report was shifted to Month 52 (30.9.2025) to include a report on first measurements of new dosimetry diodes produced in the framework of this project. |
| Report release date: | 30/07/2025 |
| Work package: | WP4: Upgrade of Irradiation and Characterization Facilities |
| Lead beneficiary: | CERN |
| Document status: | FINAL |

Abstract:

The Non-Ionizing Energy Loss (NIEL) – Hypothesis is assuming that radiation damage to solid-state detectors originating from displacement damage to the crystalline bulk of the sensors scales linear with NIEL. Within this concept every radiation field or monoenergetic particle source can be scaled to a 1 MeV neutron equivalent fluence by a hardness factor. The hardness factor is given by the ratio of NIEL deposited by the particle or particle field in question to the NIEL deposited by a 1 MeV neutron in the same sensor material. The NIEL can be calculated by considering all possible interactions an impinging particle can have with the detector material. Due to the multitude of possible interactions and their complexity, Monte Carlo based simulators for particle passage through matter, like Geant4 or FLUKA, are the most adequate tools for NIEL calculations. They can thus be used to produce hardness factors for any kind of radiation source. On the experimental side, hardness factors are usually obtained by comparing the amount of radiation damage observed after exposure to the radiation source in question as compared to the radiation damage observed after exposure to a radiation source with a well-known hardness factor.

In this work we review the limits of the hardness factor concept based on NIEL for silicon detectors and we propose a revised NIEL concept that considers the spatial distribution of displacement damage in form of single point defects and clustered crystalline defects. First estimations on the generation ratio of point to clustered defects in silicon for various particles and particle energies are given. They are based on a combined Geant4, SRIM and OPTICS simulations approach. At the same time, we propose to consolidate the quantity and quality of the available experimental data on hardness factors for various radiation facilities. The key deliverable in this context is the production of a set of silicon dosimeters for hardness factor evaluations at various radiation facilities. First measurements on the silicon dosimeters, that have been produced in form of p-i-n diodes, are presented.

AIDAinnova Consortium, 2025

For more information on AIDAinnova, its partners and contributors please see http://aidainnova.web.cern.ch/

The Advancement and Innovation for Detectors at Accelerators (AIDAinnova) project has received funding from the European Union’s Horizon 2020 Research and Innovation programme under Grant Agreement no. 101004761. AIDAinnova began in April 2021 and will run for 4 years.

Delivery Slip

|  | Name | Partner | Date |
| --- | --- | --- | --- |
| Authored by | Michael Moll Vendula Subert Ian Dawson Marie Muehlnikel | CERN CERN QMUL CERN | 01/07/2025 |
| Edited by | Michael Moll | CERN | 10/07/2025 |
| Reviewed by | Federico Ravotti [Task coordinator] | CERN | 30/07/2025 |
| Approved by | Giovanni Calderini [Scientific coordinator] Steering Committee |  | 12/08/2025 |

TABLE OF CONTENTS

1.	Introduction and Scope of the Deliverable	6

2.	NIEL: The Non-Ionizing Energy Loss	7

2.1.	The concept of NIEL	7

2.2.	The NIEL scaling Hypothesis	7

2.3.	Limitations of the NIEL scaling hypothesis	9

3.	Simulation Work Towards a NEW NIEL scaling Approach	10

3.1.	Primary damage - Geant4 simulations	11

3.2.	Damage cascades - SRIM simulations	11

3.3.	NIEL calculations – Combining Geant4 and SRIM simulations	12

3.4.	Distribution of Displaced Atoms and Vacancies	13

3.5.	The OPTICS Clustering algorithm	14

3.6.	Outlook	16

4.	Experimental Evaluation of Hardness FacTors	16

4.1.	Sensors for Dosimetry AnD further Studies	16

4.2.	Hardness FActors of Irradiation Facilities	19

5.	Summary, ConcluSions and Outlook	21

Annex: Glossary	22

References	22

Executive summary:

The Non-Ionizing Energy Loss (NIEL) – Hypothesis is assuming that radiation damage to solid-state detectors originating from displacement damage to the crystalline bulk of the sensors scales linear with NIEL. Within this concept every radiation field or monoenergetic particle source can be scaled to a 1 MeV neutron equivalent fluence by a hardness factor. In this work we review the shortcomings of the NIEL-based hardness factor concept, and we propose a revised scaling concept that considers the spatial distribution of the displaced atoms in form of single point defects and clustered defects. A simulation work on the full damage chain from primary interaction towards the final displacement in the crystalline sensor bulk is presented. Furthermore, we report on the design and production of a large set of silicon sensors serving the experimental determination of hardness factors for a wide range of radiation sources and radiation-test facilities.

# Introduction and Scope of the Deliverable

Radiation damage is posing a serious challenge to the long-term operation of silicon detectors in the harsh radiation fields of high-luminosity hadron collider experiments and other applications areas with strong radiation fields such as for example space, fusion or medical applications. Over recent decades the radiation tolerance of silicon detector technologies was strongly enhanced up to the point that it is now ready to face the challenge of the upcoming luminosity upgrade of the LHC []. Still, in perspective of future accelerator upgrades including the very ambitious FCC-hh project, further developments towards even more radiation tolerant technologies are required and are ongoing [, ].

One important aspect of radiation damage to devices operating in radiation fields composed of different particle type and particle energies is the understanding of how the radiation damage depends on the particle type and energy and how this damage can be scaled or modelled to give a consistent measure for the impacted radiation damage. The common, and very successful, way of doing this, is to scale the particle fluence to a 1 MeV neutron equivalent fluence by a factor that is usually called hardness factor (k). This hardness factor represents the scaling factor for the inflicted damage. For example, stating a hardness factor of 2.1 for 24 MeV monoenergetic protons [] means that the damage produced by 24 MeV protons is 2.1 times larger than for 1 MeV neutrons at the same particle fluence level. Experimentally, this value is most often obtained for silicon detectors by measuring the leakage current after irradiation and a well-defined post-irradiation annealing procedure []. Theoretically, the hardness factor is obtained most often by applying the Non-Ionizing Energy Loss (NIEL)-Hypothesis. This hypothesis assumes that the radiation damage as observed by the measurable properties of the detector (like e.g. the leakage current) is scaling directly with the energy that the impinging particle deposits in the detector in any form other than ionization. The energy deposition itself is calculated considering all possible forms of interaction of the particle with the silicon bulk of the detector with energy transfers above the displacement threshold.

The scope of the deliverable described in this report is to provide a significant quantity of silicon devices that allow to perform leakage current measurements for hardness factor evaluation of monoenergetic radiation sources and radiation fields, predominantly for the evaluation of hardness factors of radiation facilities in use by the High Energy Physics (HEP) detector community.

This deliverable is placed in the wider scope of evaluating the soundness of the NIEL hypothesis and correspondingly the concept of hardness factors for radiation sources. It is well known that the NIEL hypothesis has limits in validity (see section 2). A revised NIEL hypothesis extending the reach and therefore the predictive power of radiation damage models based on the NIEL concept, is therefore desirable for the R&D on future detector technologies and for understanding radiation damage to presently operating devices. To achieve this, more advanced NIEL calculations are targeted in this work to give a quality to the produced displacement damage in form of single point defects and clustered defects rather than just summing up on deposited energy (see section 3). Additionally, a wider statistics and better precision in experimental measurements are required and are fostered in this work by the production of silicon sensors for NIEL measurements (see section 4). This latter activity is also serving the above-mentioned high precision measurements of radiation facilities hardness factors (via leakage current) for the HEP community.

This document is organized as follows: Section 2 introduces the concept of NIEL, the use of NIEL as scaling parameter for radiation damage and gives examples of experimental findings for which the NIEL scaling hypothesis is not applicable. In section 3 we describe the theoretical and simulation work performed within AIDAinnova WP4 to advance in the development of a revised NIEL scaling approach. In section 4 we demonstrate that the silicon sensors for hardness factor evaluation have been produced, measured and are partly already distributed to radiation facilities. A summary of the work with conclusions and an outlook is given in section 5.

# NIEL: The Non-Ionizing Energy Loss

## The concept of NIEL

The non-ionizing energy loss (NIEL) gives the portion of energy lost by a traversing particle which does not go into ionization and eventually leads to displacement damage. However, only a fraction of the NIEL leads to displacements as part of the energy is dissipated in phonons. This fraction depends on the energy of the impinging particle. NIEL is deﬁned in units of MeVcm2/g or as NIEL cross-section and called displacement damage function D with units of MeVmb. A reference value of 1 MeV neutron equivalent (neq) has been ﬁxed to 95 MeVmb. Calculated values of the NIEL cross-sections for various particles as presently used for radiation damage predictions in the HEP community are shown in Figure 1 (see [, ] for more details).

Figure 1: Non-ionizing energy loss (NIEL) cross-sections normalized to 95 MeV mb (see text for details and references).

## The NIEL scaling Hypothesis

The NIEL hypothesis assumes that radiation damage effects scale linearly with NIEL, irrespective of the distribution of the primary displacements over energy and space. For the simulated examples shown in Figure 2 the number of vacancies should give a measure of the damage irrespective of their distribution, whether homogeneously scattered over a relatively wide volume (as for the case of low energetic proton or gamma-ray damage) or clustered in high density in small regions (as in the case of neutron damage). Consequently, the damage produced by different particles or particles with different energies should be scalable via their NIEL (i.e., the number of displacements) and the data given in Figure 1 should allow us to normalize the damage from different particles or particles with different energies.

Figure 2: Initial distribution of vacancies produced by 10 MeV protons (left), 23 GeV protons (middle), and 1 MeV neutrons (right). The plots are projected over 1 µm depth (z) and correspond to a ﬂuence of 1014 particles/cm2. Figure taken from Ref. [].

There are several areas were NIEL scaling has been proven to be a powerful method for coping with displacement damage predictions in complex radiation ﬁelds dominated by fast hadrons. The most prominent example is the increase of leakage current and HEP experiments [1, 7]. The increase of leakage current is linear with fast hadron fluence over several orders of magnitude and independent of the used silicon bulk material, as shown in Figure 3(a). Applying the NIEL scaling to the complex radiation fields of the LHC experiments allows to predict the evolution of the detector leakage current over the operational time of the experiment. This requires in addition a good knowledge on the temperature dependence and the annealing behaviour of the current. This is shown in Figure 3(b), where a good agreement between the measured and simulated leakage current of the ATLAS Pixel outer layers are shown []. More examples for other LHC experiments can be found in [7].

| (a) | (b) |
| --- | --- |
| Figure 3: (a) Radiation induced leakage current density increase as a function of particle ﬂuence for various silicon detectors made from silicon materials produced by various process technologies with different resistivities and conduction types (taken from []). (b) The absolute leakage current as a function of the integrated luminosity in Run 2 of the LHC for the outer pixel layers of the ATLAS detector. A comparison between the simulation and the measured data is shown (taken from [9]). The simulation model is called for historical reasons “Hamburg Model”. |  |

## Limitations of the NIEL scaling hypothesis

There are several experimental examples demonstrating the shortcoming of the NIEL scaling hypothesis presented in section 2.2. They are mainly resulting from the fact that point-like and clustered defects contribute differently to some device damage parameters such as for example the effective space charge of semiconductor radiation detectors. The space charge is ruling the detector internal electric field distribution and consequently the detector depletion behaviour and has strong impact on the charge collection efficiency. The distribution of vacancies in clustered and more point like appearances shown in Figure 2 for different particles gives an intuitive understanding of this shortcoming in damage scaling. In Figure 4 we provide two examples for NIEL scaling violations. Figure 4(a) shows that for detectors made from silicon with a high oxygen concentration, show a different radiation induced change in effective space charge (i.e. change in depletion voltage) for charged hadrons (protons and pions) as for neutrons when all particle fluences are normalized to the 1 MeV neutron equivalent particle fluence. Neutrons inflict a higher damage than protons on the oxygen enriched silicon detectors (see [] for more details). Figure 4(b) gives another example showing the radiation induced degradation of the gain layer (GL) of Low Gain Avalanche Detectors (LGADs). The fraction of remaining dopants in the gain layer as extracted from the measured gain layer depletion voltage is given. Like in the previous example, proton and neutron induced damage does not scale with the 1 MeV neutron equivalent fluence. However, in this case the protons are more damaging than the neutrons (see [] for more details).

| (a) | (b) |
| --- | --- |
| Figure 4: (a) Dependence of the effective space charge (Neﬀ) on the accumulated 1 MeV neutron equivalent ﬂuence for standard and oxygen enriched FZ silicon detectors irradiated with reactor neutrons (Ljubljana), 23 GeV protons (CERN PS), and 192 MeV pions (PSI) (taken from [11]). (b) Fraction of the remaining gain layer (GL) doping as function of 1 MeV neutron equivalent fluence for the same LGAD types exposed to either reactor neutrons (blue circles) or 23 GeV protons (red squares) (taken from [12]). |  |

The observed NIEL scaling violation leads to the conclusion that a more profound understanding of the damage mechanism is needed to produce a comprehensive modelling approach. A more subtle differentiation of the primary displacements produced by the different particles is needed. A simple approach of summing up the energy transferred into crystal atom displacements is seemingly not sufficient. The distribution of the primary and secondary displacements within the crystal lattice in form of single displacements up to complex accumulations of many displacements in form of damage clusters must be considered and weighted with their relative impact on the macroscopic damage parameter concerned. A revision of the used damage functions for the 1 MeV neutron equivalent ﬂuence calculation is thus of interest. In the scope of this work, we have taken a first step and target the separation of NIEL leading to individual displacements and NIEL resulting in more complex damage structures.

# Simulation Work Towards a NEW NIEL scaling Approach

In section 2.3 we have described the advantages but also several shortcomings of the presently used NIEL Hypothesis and the normalization of radiation damage induced by various particles of various energies towards a 1 MeV neutron equivalent damage. To formulate a more comprehensive modelling approach that takes also into account the quality of the produced damage rather than only the energy loss of the impinging particles going into displacements, we performed an in-depth simulation study. The aim is to take the spatial distribution of the produced displacements in form of single displacements up to the formation to complex damage clusters into account. A simulation framework was created as part of a PhD thesis [] that is based on (at least) 3 simulation steps that are depicted graphically in Figure 5: It starts with the generation of the Primary Knock-on Atom (PKA) spectra with the Geant4 (Geometry ANd Tracking) simulation framework [] considering Coulomb elastic, nuclear elastic and nuclear inelastic scattering processes. The output spectra of the Geant4 simulations are fed into a SRIM (Stopping Range of Transport in Matter) simulation, which then outputs a 3D map of displacements, or more precisely, a 3D map of vacancies and the energies of the corresponding interstitials. In practice, these two simulation steps don’t have to be done sequentially, as the SRIM calculations can be produced in form of look-up tables that are used after the Geant4 simulations. The combined Geant4/SRIM simulations are then handed over to a customized OPTICS (Ordering Points To Identify the Clustering Structure) simulation that results into the number of created point and clustered defects, respectively the NIEL going into point and clustered defects. All these steps are described in more detail in the following sections (sections 3.1 to 3.5) and are concluded by an outlook towards further steps (section 3.6).

Figure 5: Flowchart illustrating the simulations and algorithmic tools utilized for the damage modelling. Inputs, shown as black parallelograms, include high-energy protons, neutrons, electrons, and γ-rays, which are simulated using Geant4. Ions with atomic numbers ranging from 1 to 15 are simulated using SRIM. By combining the results from Geant4 and SRIM standard “NIEL curves” as e.g. published by the RD48 collaboration [6] can be produced. By applying the OPTICS algorithm on the simulation results provided by SRIM for the low energy ions we develop a model characterizing clustered and isolated displacements for low energy ions. By combination of the outputs of Geant4 and SRIM with the subsequent analysis by OPTICS as indicated by the black arrows in the sketch we develop a model separating NIEL going into clustered and NIEL going into isolated displacements for high energy incident particles (graph taken from [13]).

## Primary damage - Geant4 simulations

The typical simulation setting in Geant4 was a pencil beam of the impinging particles directed towards a 100 mm thick silicon substrate representing a generic silicon detector. Coulomb elastic, nuclear elastic and nuclear inelastic scattering processes were taken into account in the simulations.

The Coulomb scattering process, i.e. the interaction of impinging charged particles and the electromagnetic field of the silicon atom, was initially simulated with Geant4 using the appropriate single scattering modules but quickly found to be extremely CPU-intensive and thus unpractical. An analytical model was therefore used instead [] after verifying that the Geant4 scattering and the analytical model resulted in almost identical results for the cross sections.

The nuclear elastic scattering is mediated by the strong force between the impinging hadronic particle and the silicon nucleus. Due to the nature of the strong force the energy loss of the incoming particle is in average higher than in electromagnetic interactions, but appearing with a lower probability, i.e. a lower total cross section. Nuclear elastic scattering can also be mediated by the electromagnetic force when charged particles scatter at the electromagnetic field of the silicon nucleus. All nuclear elastic scattering processes were entirely simulated with Geant4, mainly using the QGSP_BERT_HP physics list and for protons and neutrons above 3 GeV the FTFP_BERT_HP physics list.

The nuclear inelastic scattering is complex as nuclear fragments from alpha particles up to silicon must be considered. These fragments were produced within the Geant4 simulation, but then further tracked for their subsequent damage cascades in the silicon crystal by the SRIM simulation (see section 3.2).

The typical output obtained from the Geant4 simulations for the differential cross sections of the various silicon recoils and nuclear fragments are shown in Figure 6 for 200 MeV protons and 100 MeV neutrons.

| Figure 6: Differential cross sections ∂σ/∂(lnErecoil) for 200 MeV protons and 100 MeV neutrons incident on a silicon target. |  |
| --- | --- |

## Damage cascades - SRIM simulations

In the previous sections we discussed the primary interactions of the impinging particles with the silicon atoms resulting in spectra of PKAs and nuclear spallation products. These primary displaced particle spectra were further treated with the help of SRIM (Stopping and Range of Ions in Matter) simulations. SRIM is a collection of Monte Carlo simulation tools that is widely used to calculate many features of the transport of ions in matter. Typical applications are ion stopping and range in targets, sputtering processes and ion transmission. The simulation outputs the spatial distribution of vacancies and the final energetic distribution of interstitials for each impinging ion, in this case PKAs or spallation fragments. The 3D distribution of vacancies delivered as output of the SRIM simulations is then subjected in this work to an algorithm that is described in the next section (Sec. 3.5) to attribute the observed displacements into either clustered displacements or individual displacements. Another approach that is not followed in this work due to its complexity is to subject displaced atoms and Nuclear Fragments below a certain kinetic energy threshold to Molecular Dynamics (MD) simulations followed by Kinetic Monte Carlo (KMC) calculations, as described for example in []. In any case, summing up over all energy transfer not going into ionization, allows to calculate the classical NIEL. Comprehensive benchmarking efforts of our SRIM simulation approach against available literature and theoretical data were performed. Two examples are given in Figure 7. The Figure 7 (left) shows the total number of vacancies extracted from the simulation as function of the silicon PKA energy. Two methods were used to extract the numbers, a direct reading from the simulation GUI and a reading using a customized script to run the simulation and extract the data. Both yielded the same results within the rounding errors, as expected. A comparison against the Norgett-Robinson-Torrens (NRT) model [] is as well given showing only small differences. Figure 7 (right) gives the NIEL (in units of keV) extracted from the SRIM simulations for various PKA energies for silicon and other ions. The results for silicon are close to the classic Lindhard approximation [] and thus confirm the correct implementation of the SRIM simulation tool in our scripting.

| Figure 7:(left) Cross-comparison of SRIM script with SRIM GUI and NRT-DPA data. The script used to launch, save, and read out data from SRIM matches the SRIM GUI results within the rounding limits of the GUI (two-digit precision). The NRT-DPA reports a slightly lower number of displaced atoms. (right) NIEL in units of keV as a function of PKA energy. NIEL values for helium, carbon, neon, and silicon are compared against the Lindhard approximation for silicon. |  |
| --- | --- |

## NIEL calculations – Combining Geant4 and SRIM simulations

Classical NIEL calculations were performed combining the cross sections for the production of primaries (PKAs and nuclear fragments) from Geant4 simulations (section 3.1) with the SRIM simulations on the subsequent cascades of the primaries in silicon (see section 3.2). The results of the combined simulations are shown in Figure 8 for protons, neutrons and electrons in a wide energy range. The data are in good agreement with previously obtained data and thus further validate the approach taken.

Figure 8: Comparison between literature data (“The RD48 data standard” [6]) and the Geant4-SRIM simulations performed in this work. The simulation is represented by coloured circles, whereas the RD48 standard is expressed by solid lines. The black circles represent the compilation of measured data for the original RD48 ﬁt of the high energy range. Y-axis shows the damage equivalent (D(E)) as a function of energy for various incident particles: Low energy neutrons (blue) by Griﬃn and mid-range neutrons by Konobeyev (green) are matching the simulated neutron data (cyan) excellently. Electrons by Summers (pink) and low energy protons Summers (red) are also matching simulation data (pink and red respectively) very well. High energy neutrons and protons data show a slight curvature where the ﬁt data from RD48 were linear. See [6] for more details on the literature data and [13] for a more detailed explanation of the NIEL curves obtained in this work.

## Distribution of Displaced Atoms and Vacancies

The displacement damage curves in the previous sections were obtained by running first Geant4 to obtain the energetic distribution of displaced silicon atoms and nuclear fragment produced by the primary particle. Then SRIM was used to obtain the number of vacancies for the displacement damage function calculations. To investigate the structure of clusters in heavily irradiated silicon, we follow the same simulation order, while the position, kinetic energy and direction of each particle from Geant4 is stored and used as input to produce with the help of SRIM a true 3D distribution of vacancies for multiple primary particles passing the silicon slice under investigation. Three typical examples obtained for 23 GeV protons, 1 MeV neutrons and 27 MeV electrons impinging on a silicon slice of 100 mm thickness are shown in Figure 9. The sampling volume is of 1mm x 1mm x 1mm and located several tens of mm in the depth of the silicon slice. The silicon was exposed to 106 particles and the vacancies are shown in a 2D projection of the volume as black dots. The recoil production threshold was set to 21 eV. As is visible from the plots the distribution of the vacancies is very different. For the 23 GeV proton case, individual vacancies are visible originating from the coulomb interaction of the proton with the silicon, while also dense long tracks of clustered vacancies can be seen originating from the products of nuclear fragmentation. For the 1 MeV neutrons the nuclear elastic scattering is dominating and producing larger conglomerates of clustered vacancies. Finally, for the 27 MeV electrons only a few clustered vacancies can be seen while the spread of individual vacancies dominates the picture. This visualization clearly indicates that a pure measure of produced vacancies is not be enough to estimate macroscopic damage to silicon detectors, as defect spectroscopic measurements have already shown that clustered and individual point defects in the silicon lattice have substantial different impact on detector performance [1]. A measure for the clustering of vacancies is therefore needed. A first approach to this problem is given in the next section.

| (a) 23 GeV protons | (b) 1MeV neutrons | (c) 27 MeV electrons |
| --- | --- | --- |
| Figure 9: Comparison of the 2D projected Geant4 simulated vacancy distribution in a volume of 1mm x 1mm x 1mm within a piece of silicon exposed to 106 particles/mm2. The figure shows typical example for (a) 23 GeV protons, (b) 1 MeV neutrons and (c) 27 MeV electrons. |  |  |

## The OPTICS Clustering algorithm

Traditional NIEL models typically sum over total energy deposition, assuming a response in damage parameters that is scaling with the total number of displacements. However, experimental observations, particularly via deep-level transient spectroscopy (DLTS) and thermally stimulated current (TSC) measurements, show that spatial clustering of defects takes place and significantly affects sensor performance. The fraction of clustered defects as compared to the overall number of created defects can have an impact on the leakage current and the changes in depletion voltage and therefore lead to effects that are not proportional to the simple total energy loss, i.e. the classical NIEL. In order to get an estimate of the fraction of clustered displacements we applied a clustering algorithm to the datasets produced by the SRIM simulations. These simulations provided a three-dimensional map of vacancies for each simulated impinging ion to which we applied the OPTICS Algorithm. OPTICS stands for Ordering Points To Identify the Clustering Structure and is a density-based clustering algorithm developed by the Machine Learning (ML) community to identify clustered data in data sets []. The algorithm translates in our case the three-dimensional map of displacements into a linearly ordered list of displacements in such a way that spatially closest displacements become neighbours in the ordering. Clusters are then identified on the basis of this list that can also be visualized in form of a so-called reachability plot (for more details on the algorithm see [13,19]). An example is given in Figure 10 for a single event originating from a SRIM simulation of a 100 keV PKA in silicon. Out of the 142 produced vacancies, 103 are attributed to be part of a cluster of vacancies. The output is strongly impacted by the setting of the OPTICS parameters and the displacement damage threshold energy. This is shown in Figure 11 for a wide range in PKA recoil energies in silicon and three different displacement thresholds of 10, 21 and 30 eV. In addition, the parameter epsilon (ε) of the OPTICS algorithm was modified to match the nearest-neighbour distance (2.35 Å) and multiples of it (4.7 and 7.05 Å). As the parameters cannot be easily constraint, further experimental results are needed to converge on a solid parameter set for more profound simulations (see [13] for more details). The algorithm needs to be tuned on experimental data sets before gaining predictive power. This has so far not been reached in the present work, mostly because it is unclear how to translate the obtained three-dimensional distribution of displacements into a reliable measured units of electrical active defects in the silicon bandgap, respectively into generation centres ruling the silicon sensor leakage current or charged defects ruling the silicon sensors space charge and thus the depletion voltage.

Figure 10: Example of clusters extracted from a 100 keV silicon recoil event simulated by SRIM. The numbers in all four sub-figures indicate the 5 diﬀerent clusters as identified by the OPTICS algorithm. The coloured points represent displacements (i.e. vacancies) within clusters, while the grey areas denote isolated displacements. (Fig. a) Reachability plot of the 142 displacements produced in the event. (Fig. b) Three-dimensional coordinate plot of the displacements and projections in the Z-X (Fig. c) and Z-Y (Fig. d) plane.

Figure 11: Fraction of the NIEL deposited in clusters as evaluated with the OPTICS algorithm with three different parameter settings for the parameter ε of the OPTICS clustering algorithm (Yellow: ε = 2.4, Red: ε = 4.7, Blue: ε = 7.05) and three different displacement thresholds for the SRIM calculations (10,21 and 30 eV). A strong variation of the results with the parameters is observed, while the dependence on PKA energy above about 200 eV is small in all cases up to the MeV range.

## Outlook

We are presently working on an optimized clustering algorithm. A drawback of the approach used in the previous section is that the 3D map of vacancies originating from the SRIM simulations is used without considering that the vacancies and interstitial will move in the crystal lattice after their creation. They migrate through the lattice and might recombine when a vacancy encounters an interstitial, might form higher order intrinsic defects like di- or tri-vacancies or eventually might escape the primary damage region and interact with already existing defects or impurities. Such simulation would require the consideration of Molecular Dynamics (MD) simulations for the initial phase of movements and for the long-range migration Kinetic Monte Carlo (KMC) simulations. While full scale MD and KMC simulations are out of the scope of our present work, a simplified approach of moving and clustering vacancies after their generation in a 3D modelling approach similar to KMC modelling for long range migration is under development (see [13] for more details). However, as for MD and KMC simulations, also this simplified modelling approach will have to be tuned to match the macroscopic radiation damage observations and therefore further experimental data are needed. Work towards an extended experimental data set will be presented in the next section.

# Experimental Evaluation of Hardness FacTors

## Sensors for Dosimetry AnD further Studies

For the precise determination of the radiation induced leakage current density, PIN diodes made from high resistive silicon with a thickness of some hundred microns, some square millimetre area, and a well-defined active volume are ideal devices. The active volume is constrained by using a guard ring and the effective dimension of the volume is determined from Capacitance-Voltage (CV) and/or laser scanning measurements. Several hundred identical diodes are needed to allow for the characterization of several irradiation sources with sufficient statistics on the experimentally derived hardness factors that build on precise leakage current measurements. A cost-effective solution was found with the help of the RD50 collaboration and the sensor producer CNM, Barcelona [], which offered to use an already existing 8-inch mask set for p-type sensors with almost ideal dimensions. The wafer mask drawing is shown in Figure 12. It was originally designed for Low Gain Avalanche Diodes (LGADs) but allowed to produce p-type PIN devices by simply not implementing the gain layer implant.

Figure 12: CNM Mask set 869 for 150 mm diameter silicon wafers used to produce the PIN sensors.

The mask holds a total of 554 PIN diodes. 536 devices on a 5x5 mm2 die with an active area of 3.3x3.3 mm2 and 18 devices on a 3x3 mm2 die with an active area of 1.3x1.3 mm2.

| (a) | (b) |
| --- | --- |
| Figure 13: (a) Metal layer on the frontside (n+-contact) of the PIN sensor with 3.3x3.3 mm2 active area on a 5x5 mm2die. The central opening in the metal is implemented to allow for laser testing of the devices. A single guard ring is surrounding the active area. (b) Metal layer on the backside (p+-contact). The opening in the metal allows for laser injection. |  |

The production with the CNM run number 17933 and employing the mask set 869 was launched in January 2024 and finished in late 2024. During production 2 out of the 10 high resistivity FZ silicon wafers broke. The remaining 8 wafers underwent current-voltage (IV) measurements up to -200 V reverse bias at CNM showing a good yield of more than 90% whereas good devices were characterized as giving a current of less than 1 µA at 200V. Figure 14(a) shows the leakage current for all 443 PIN sensors of wafer 4 that were qualified as good out of the overall 472 sensors that were measured. This corresponds to a yield of 94%.

| (a) | (b) |
| --- | --- |
| Figure 14: (a) IV characteristic of all 443 “good” PIN sensors of wafer 4 with I(-200 V) < 1 µA and (b) distribution of current values at -200 V for all measured 472 devices of the wafer including the “bad” sensors data. Measurements were performed with a wafer prober at CNM at room temperature. |  |

In January 2025 the 8 fully processed and diced wafers were delivered to CERN. A photograph of one diced wafer is shown in Figure 15.

Figure 15: Photograph of the 150 mm diameter silicon wafer as delivered to CERN. Note that the two middle rows of sensors are missing as they had to be removed for the dicing of the smaller sensors (a drawing of the full wafer is given in Figure 12).

A subset of sensors underwent further characterization at CERN. From each wafer three samples classified as good by the foundry and one sample failing the criterion of I(-200V) < 1µA, i.e. a sample classified as bad were re-measured at CERN. The corresponding results are shown in Figure 16.

| (a) | (b) |
| --- | --- |
| Figure 16: (a) Leakage current as measured on the central pad of the diced sensors at CERN with the guard ring connected to ground. The results measured before dicing at the foundry lie within the red shaded area. (b): Total current measurements performed at CERN. Both figures show only measurements of diodes that passed the selection criteria, i.e. “good” sensors. |  |

The measured currents, both from CNM and from CERN, are slightly higher than the typical values expected for pristine silicon sensors of this size. Measurements performed at CERN after dicing (Figure 16(a)) consistently yield lower values. Even diodes that were declared ‘bad’ when first measured at CNM, were found to pass the selection criteria (I(-200V) < 1 µA) in the measurement performed at CERN. As for both measurements, at CNM and at CERN, the guard ring was connected to the same potential as the pad, no obvious reason was identified for this difference in measured currents. In Figure 16(b) the total current, i.e. the sum of currents flowing through the pad and the guard ring are displayed.

The variation among measurements from different PIN diodes is bigger than the one observed for the pad current. Finally, the capacitance vs. voltage behaviour was measured for the diodes and is displayed for three sensors in Figure 17. The depletion voltage is found to be between -30V and -50 V for the set of received diodes.

Figure 17: Capacitance Voltage (CV) measurement for two diodes.

## Hardness FActors of Irradiation Facilities

The diodes described in section 4.1 are currently being prepared for distribution to different irradiation facilities. Based on the NIEL scaling hypothesis, the change in leakage current from pre- to post- irradiation (∆I), can be used to determine the current related damage factor  defined as

α=ΔIV×Φ                                                                              (Eq. 1)

where V is the active volume of the device and the  Φ is the absolute particle fluence. Comparing the damage factor obtained with a monoenergetic 1 MeV neutron irradiation (αeq) results in the hardness factor κ for the facility, respectively the particle type and energy in question:

κ=ααeq                                                                              (Eq. 2)

The same can be obtained for a complete particle spectrum. For the experimental evaluation it is very important to take into account the temperature history of the irradiated sensor after irradiation (i.e. the annealing) and the measurement temperature. More details about the measurement procedure and annealing effects can be found in [10, 5].

An exemplary dataset and fit for the extraction of the hardness factor for the 24 MeV proton irradiation facility (MC40 cyclotron) in Birmingham is depicted in Figure 18 (a). A set of diodes is exposed to different fluences of 24 MeV protons and the extracted increase of the measured leakage current, subsequent to an annealing for 80 minutes at 80 C, is used to obtain the hardness factor [4]. Another set of data is given in Figure 18 for 7-10 MeV protons from experiments performed at the Tandem Van de Graaﬀ accelerator of the Laboratoire des Faisceaux ioniques de l’Université de Montréal, Canada and for 24 GeV/c protons obtained at the IRRAD facility of CERN, Switzerland [].

| (a) | (b) |
| --- | --- |
| Figure 18: (a) Radiation induced increase of the leakage current as a function of 24 MeV proton ﬂuence for diodes irradiated at the MC40 cyclotron facility in Birmingham (taken from [4]); (b) The leakage current as function of proton fluence for a set of sensors exposed to protons of the energy given in the legend (taken from [22]). |  |

Each of those measurement campaigns results in a hardness factor that can be compared against the theoretical expectations from the NIEL hypothesis and visualized as the Displacement Damage Functions given in Figure 1. Several experimental hardness factors have been determined for different irradiation facilities. A selection of those data, as obtained from literature, is given in Figure 19. The number in the legend indicates the reference for the shown data points (references in order of appearance in the legend: 1:[4], 2:[], 3:[], 4:[5], 5: [], 6: [], 7: [], 8: [], 9: []).

The set of dosimetric diodes produced in the scope of the AIDAinnova WP4 project will be used to continue this work. A large set of identical diodes is now available that will ease the cross-comparison together with a homogenised irradiation, annealing and measurement procedure. The use of identical diodes across all facilities ensures consistency and comparability.

A better experimental definition of the hardness factor for various particles and particle energies is expected to give an important input for a revised NIEL scaling concept.

Figure 19: Selection of hardness factors for different irradiation facilities (see text and Figure legend). Uncertainties were not reported for all facilities.

# Summary, ConcluSions and Outlook

In this work we reviewed the shortcomings of the NIEL-based hardness factor concept, and we proposed a revised scaling concept that considers the spatial distribution of the displaced atoms in form of single point defects and clustered defects. A simulation work on the full damage chain from primary interaction towards the final displacement in the crystalline sensor bulk was presented. It is based on a combination of the Geant4 and SRIM simulation frameworks for the generation of displacements in the silicon crystal and on the OPTICS algorithm for estimating the ratio between point defects and clustered defects. This work is still in progress but has already delivered a first preliminary estimate. Furthermore, we reported on the design and production of a large set of silicon sensors serving the experimental determination of hardness factors of a wide range of radiation sources and radiation facilities. A part of the sensors has been distributed to radiation facilities already and further experiments are under way in the framework of a PhD project at CERN.

In conclusion, the production of sensors that form the core of the AIDAinnova D4.2. deliverable serve twofold to the community. On the one hand, we progress on the fundamental understanding of radiation damage with an extended set of experimental hardness factor data. These are serving to benchmark the complex simulations for extending the NIEL scaling concept. On the other hand, we contribute to producing reliable hardness factors for the radiation facilities used in the HEP community.

# Annex: Glossary

| Acronym | Definition |
| --- | --- |
| GL | Gain Layer (of an LGAD detector) |
| HEP | High Energy Physics |
| KMC | Kinetic Monte Carlo |
| LGAD | Low Gain Avalanche Detector |
| MD | Molecular Dynamics |
| NIEL | Non Ionizing Energy Loss |
| OPTICS | Ordering Points To Identify the Clustering Structure (Name of an algorithm) |
| SRIM | Stopping and Range of Ions in Matter (Name of a software package) |
| TRIM | Transport of Ions in Matter (Name of a software package) |

# References

[] 	M. Moll, "Displacement Damage in Silicon Detectors for High Energy Physics," in IEEE Transactions on Nuclear Science, vol. 65, no. 8, pp. 1561-1582, Aug. 2018, https://doi.org/10.1109/TNS.2018.2819506

[] 	ECFA Detector R&D Roadmap Process Group, “The 2021 ECFA detector research and development roadmap”;  http://dx.doi.org/10.17181/CERN.XDPL.W2EX

[] 	The DRD3 collaboration “Solid State Detectors”, https://drd3.web.cern.ch/ (visited 1.7.2025)

[] 	P. Allport et al., Experimental determination of proton hardness factors at several irradiation facilities, Journal of Instrumentation, 2019 JINST 14 P12004, https://doi.org/10.1088/1748-0221/14/12/P12004.

[] 	M. Moll et al., Leakage current of hadron irradiated silicon detectors – material dependence,  NIMA, Volume 426, Issue 1, 21 April 1999, Pages 87-93, https://doi.org/10.1016/S0168-9002(98)01475-2

[] 	A. Vasilescu and G. Lindström, Displacement damage in silicon, http://rd50.web.cern.ch/NIEL/default.html  (visited 1.4.2025)

[] 	I. Dawson (Editor), Radiation effects in the LHC experiments: Impact on detector performance and operation; CERN Yellow Report Vol.1 (2021). https://doi.org/10.23731/CYRM-2021-001

[] 	M. Huhtinen, Nucl. Instrum. Meth. A491 (2002) 194, https://doi.org/10.1016/S0168-9002(02)01227-5

[] 	ATLAS Collaboration, Measurement of radiation damage through leakage current monitoring of the ATLAS pixel detector, ATL-INDET-PUB-2019-001 (CERN, Geneva, 2019); https://cds.cern.ch/record/2699903.

[] 	M. Moll, Radiation damage in silicon particle detectors: Microscopic defects and macroscopic properties, PhD thesis: Hamburg University, 1999; https://doi.org/10.3204/PUBDB-2016-02525

[] 	G. Lindström et al., Nucl. Instrum. Meth. A466 (2001) 308, https://doi.org/10.1016/S0168-9002(01)00560-5

[] 	E. Currás Rivera et al., Gain layer degradation study after neutron and proton irradiations in Low Gain Avalanche Diodes, 2023 JINST 18 P10020; https://doi.org/10.1088/1748-0221/18/10/P10020

[] 	Vendula Maulerova, CERN and University of Hamburg, private communication April 2025 and draft of dissertation work with preliminary title “Characterization and Modelling of Non Ionizing Energy Loss in Silicon Sensors for High Energy Physics”.

[] 	Geant4 – A toolkit for the simulation of the passage of particles through matter, https://geant4.web.cern.ch/ (online, last visited 10.6.2025)

[] 	M.J. Boschini, C. Consolandi, M. Gervasi, S. Giani, D. Grandi, V. Ivanchenko, S. Pensotti, P.G. Rancoita, and M. Tacconi, Nuclear and Non-Ionizing Energy-Loss for Coulomb Scattered Particles from Low Energy up to Relativistic Regime in Space Radiation Environment, in Astroparticle, Particle, Space Physics, Radiation Interaction, Detectors and Medical Physics Applications Cosmic Rays for Particle and Astroparticle Physics, pp. 9-23 (2011) https://doi.org/10.1142/9789814329033_0002

[] 	M. Raine et al., Simulation of Single Particle Displacement Damage in Silicon – Part I: Global Approach and Primary Interaction Simulation, IEEE Transactions on Nuclear Science, vol. 64, no. 1, pp. 133-140, Jan. 2017, https://doi.org/10.1109/TNS.2016.2615133

[] 	M. J. Norgett, M. T. Robinson, and I. M. Torrens. A proposed method of calculating displacement dose rates. Nuclear Engineering and Design, 33(1):50–54, 1975.

[] 	A. Akkerman and J. Barak. New partition factor calculations for evaluating the damage of low energy ions in silicon. IEEE Transactions on Nuclear Science, 53(6):3667–3674, 2006.

[] 	Mihael Ankerst; Markus M. Breunig; Hans-Peter Kriegel; Jörg Sander (1999). OPTICS: Ordering Points To Identify the Clustering Structure. ACM SIGMOD international conference on Management of data. ACM Press. pp. 49–60. CiteSeerX 10.1.1.129.6542.

[] 	Instituto de Microelectrónica de Barcelona (IMB-CNM), Barcelona, Spain.

[] 	D.Bechevet et al., Results of irradiation tests on standard planar silicon detectors with 7–10 MeV protons, Nucl. Instrum. Meth. A479, Volume 479, Issues 2–3, 1 March 2002, Pages 487-497, https://doi.org/10.1016/S0168-9002(01)00925-1

[] 	E. Liu et al., Measurements of proton hardness factors in silicon at energies between 10 MeV and 25 MeV, JINST 20 P01024, 2025

[] 	J.Sorenson et al., Investigation of LGADs exposed to proton fleunces beyond 10e15neq/cm2, VERTEX 2023 poster, https://agenda.infn.it/event/35597/contributions/211782/attachments/111665/159378/vertex_draft_5_joey.pdf

[] 	G. P. Summers et al., Correlation of particle-induced displacement damage in silicon. IEEE Transactions on Nuclear Science, 34(6):1133–1139, 1987

[] 	H.J. Ziock et al., Measurement of proton induced radiation damage to cmos transistors and p-i-n diodes. IEEE Transactions on Nuclear Science, 37(3):1238–1241, 1990

[] 	E Barberis et al., Temperature effects on radiation damage to silicon detectors. NIMA: 0168-9002, 1993

[] 	T. Ohsugi et al., Radiation damage in silicon microstrip detectors. NIMA 265(1):105–111, 1988

[] 	M. Nakamura et al,. Radiation damage test of silicon multistrip detectors, NIMA 270(1):42–55, 1988