## Slide 1: AMS data analysis pipeline: from PIKA output to SOFI input
Julian Weng :: Paul Scherrer Institute
Offline data analysis concepts series: 30.04.2024

## Slide 2: What’s not covered here
Data analysis steps in Squirrel & PIKA (see video tutorials and documentation)
Error propagation details
Ion filtering
Implementations of the steps and analysis platforms

## Slide 3: What do we want to get
For each sample and AMS ion, ambient concentration scaled data [ug/m3] and related uncertainties that can go to further analysis, e.g. PMF

## Slide 4: What do we start with?
4 AMS data files exported from PIKA (frequency units)
AMS HROrg
AMS HRNO3
AMS HRSO4
AMS AirBeam

## Slide 5: What do we start with?
Autosampler file: info about sample and waterblank starting and ending time
samples
waterblanks
Sample 1
Sample 2
WB 1
WB 2

## Slide 6: What do we start with?
Sample info:
linking each sample to the filter sampling date & effective extraction volume

## Slide 7: What do we start with?
WSOC (Water soluble organic carbon) from Shimadzu analyzer:
Retrieved ambient concentrations [ug/m3]

## Slide 8: Remove bad run numbers based on N2 signal
Cutoff defined by visual inspection

## Slide 9: Sample and wb assignment to AMS run numbers
samples
waterblanks
Notes: Combine info of autosampler data and PIKA output files

## Slide 10: Sample and wb assignment to AMS run numbers
Full campaign

## Slide 11: Sample and wb assignment to AMS run numbers
1 Day

## Slide 12: Based on 15NO3 signal: filtering of outliers & RN with NaN values
Before
Notes: Outlier removal
Despite blacklisting samples in Squirrel/PIKA based on AB

## Slide 13: Based on 15NO3 signal: filtering of outliers & RN with NaN values
After

## Slide 14: Based on 15NO3 signal: filtering of outliers & RN with NaN values
Outlier removal: Median absolute deviation
For each RN:
difference to its corresponding sample median
For each RN
difference / median(differences of all RN)
If difference / median(differences of all RN) > threshold
-> remove RN as outlier
Sample i median
RN of sample i
Typical threshold: 3

## Slide 15: Normalization by internal standard
AMS Signal intensity (Hz)
Extract Concentration (ug/mL)
Note: Typically, normalization done before averaging, i.e. normalization on a point by point basis (rather than for integrated period)
Pro:
By normalization, variability introduced by unstable nebulization process compensated; this variability should not be considered a measurement uncertainty
Con:
If most variability (before norm) would rather come from instrument instability (e.g. high AB fluctuations), normalization on integrated period basis (i.e. first averaging, then normalization) could make sense to avoid introduction of additional noise -> Usually not the case

## Slide 16: Normalization by internal standard
For AMS ion i:
(Spike = Internal standard)

## Slide 17: Normalization by internal standard
2 options: 15N-nitrate or 34S-sulfate
Before
After (normalized)

## Slide 18: Normalization by internal standard
15N-nitrate
After (normalized)

## Slide 19: Averaging
For each sample and its corresponding waterblank:
Average raw data and normalized data
Use RN to find values in AMS files
AMS HRNO3
average
raw
normalized

## Slide 20: Effective extraction volume correction
Problem: for samples where filters were extracted, effective spiking volume does not equal the theoretical one, as volume is held back in vials, syringes and filter punches
=> Internal standard concentration effectively changed (higher than expected)
Corrected for by weighing vials before (empty) and after extraction
Theoretical: 12 mL
Experience from JFJ: 10-25 % deviation

## Slide 21: Effective extraction volume correction
>1

## Slide 22: Effective extraction volume correction
Only on normalized data of ‘samples’ where filters were extracted
i.e. not on waterblanks, QC samples like ANAS, Spiked ANAS, …
Before
After
Theoretical: 12 mL

## Slide 23: Waterblank subtraction
For normalized data, for each ion subtract sample average (effective volume corrected) with it’s corresponding wb average
sample – wb = waterblank subtracted sample

## Slide 24: Error estimation
Assumption
Main source of uncertainty:
ion-to-ion signal variability at the detector & ion counting statistics
Fluctuations of aerosol due to nebulizer instability ruled out by normalization with internal standard (should not be considered as measurement uncertainty)

## Slide 25: Error estimation
Sample 1
Sample 2
WB 1
WB 2
Error estimated from fluctuation of signal
within the runs corresponding to
a sample/waterblank
Sample standard deviation
‘(n-1)’
For independent samples

## Slide 26: Error estimation
Sample 1
Sample 2
WB 1
WB 2
BUT: runs are autocorrelated
past values of the variable (signal intensity)
are predictive of future values
=> this has enhancing effect on uncertainty
To account for this, number of runs per
sample n is reduced to effectvie
number of runs neff

## Slide 27: Error estimation
Sample 1
Sample 2
WB 1
WB 2
Calculation of autocorrelation coefficients
r_k for different lag times k (k=1,2,3,...)
between AMS runs within a sample
The higher is k, the less similar are runs
-> r_k decreases with higher k.
A coefficient n_c is set equal to
k corresponding to last positive value
of r_k within a sample.

## Slide 28: Error estimation
From this, get neff
And corrected standard deviation and standard error:
Standard Deviation ~ variability within a set of measurements
Estimate of Standard Error ~ variability within the means from multiple sets of measurements
~ precision of the mean

## Slide 29: Error propagation
We have estimated the error at this stage of the data processing (i.e. at averaging step)
Question: how should it be propagated in the following steps (i.e. mathematical operations)? And which uncertainties should be considered?
-> This depends on what analysis we want to make (or how we want to report the data)

## Slide 30: Error propagation
Just reporting ambient concentations… Propagate errors through all mathematical expressions (i.e. data treatment steps) considering the uncertainties of variables that flow into the pipeline:
Gaussian error propagation
Assume uncertainties of variables normal distributed & independent
For function:
Combined uncertainty derives from partial derivative with respect to each variable:
Examples:

## Slide 31: Filtering highly problematic peak fit ions
These are ions (and their isotopes) that can considerably affect further steps when calculating ion ensample properties (OM/OC rations etc) if kept
Example: CH2N (in shoulder of N2 at m/z 28) & isotopes
[CH4 (in shoulder of O at m/z 16)]
How to identify? -> scatter matrix of all organic ions in repeated QC samples
-> check the outliers (i.e. that are highly fluctuating between identical samples)

## Slide 32: Filtering highly problematic peak fit ions
CO2
CH2N
CH4

## Slide 33: Pieber correction
Perform Pieber correction on CO2 and ions that are scaled by it:
[CO2]pieber = [CO2] - ([14NO3] + [15NO3]) * Pieber ratio
Sum of internal standard and natural sample nitrate

## Slide 34: Pieber correction
Before
After
ANAS (no spiked)

## Slide 35: Ambient quantification
2 options:
Scaling to WSOA (from WSOC of Shimadzu TOC analyzer, using OM/OC rations from AMS)
Internal standard method

## Slide 36: Ambient quantification – Scaling to WSOA
Retrieve WSOA [ug/m3] from WSOC [ug/m3]:
WSOA = WSOC * OM/OC
Then scale normalized intensities of each AMS ion i to WSOA for each sample j:
Ion conc. (i)j = (signal(i)j)/(total signalj) * WSOAj
(after scale, sum of all ion intensities will equal WSOA)
With OM/OC: Organic Matter/Organic Carbon <-> Retrieve from AMS spectra

## Slide 37: Retrieve OM/EC from AMS spectra
Relative mass concentrations Mi of a sample for element i (i = C, O, N, H):
Element mass fraction of element i in AMS ion j
e.g. F(C)CHO+ = M(C) / M(CHO) = 12.0107 / 29.00284
Ion intensity of ion j

## Slide 38: Retrieve OM/EC from AMS spectra
Element mass fraction of element i in AMS ion j, e.g. F(C)CHO+ = M(C) / M(CHO) = 12.0107 / 29.00284

## Slide 39: Retrieve OM/EC from AMS spectra

## Slide 40: Retrieve OM/EC from AMS spectra
Theoretical atomic ratios: O/C, N/C, S/C, H/C
from relative mass concentrations Mi by (example for O/C):
O/C = [MO/16] / [MC/12]
However, several biases...
fragmentation with respect to electronegativity
rearrangements of ions in EI process
=> Need for empirical calibration factor α
Atomic mass of O and C
Default calibration factors from APES maybe not good for offline… (& depending on instrument status)…

## Slide 41: Retrieve OM/EC from AMS spectra

## Slide 42: Retrieve OM/EC from AMS spectra
OM/EC from calibrated atomic ratios: O/C, N/C, S/C, H/C
*also include S/C
(though no theoretical calibration
factor for that)
+ *

## Slide 43: Retrieve OM/EC from AMS spectra
Sanity check with levoglucosan QC samples: C6H10O5 -> Theory: OM/OC = 2.25
Default O/C cal factor (APES)

## Slide 44: Ambient quantification – Scaling to WSOA
Caveats of scaling to WSOA:
Assumption: WSOA derived from Shimadzu = WSOA from AMS
i.e. same analytes covered…
However, covered by AMS but not by Shimadzu
Inorganic carbon
Whatever organic is cleaved and released by acid and purging with air (therefore Shimadzu calls it non-purgeable organic carbon – NPOC)
Calibration factors for atomic ratios and hence OM/OC species-dependent
Detection limits and sensitivities of TOC analyzer

## Slide 45: Ambient quantification – Internal standard method
Transformation of waterblank subtracted, internal standard normalized data, i.e. quantified water extract concentrations [ug/mL] to ambient concentrations [ug/m3]

## Slide 46: Ambient quantification – Internal standard method
Caveats of internal standard method:
Matrix effects that effect internal standard (formation of insoluble or (more) refractory complexes with sample, evaporation)
RIEs (relative ionization efficiencies) species-dependent… we need to apply ensample value for organics
At the moment (and at least since late 2022): RIEs of L-ToF instrument appear to be perturbed in a way, that mass balance to WSOA from Shimadzu impossible
=> Use Levoglucosan and Spiked ANAS QC samples to retrieve effective RIEs (and always check WSOA scaled data as reference/validation method)

## Slide 47: Fieldblank subtraction
Until this step: Fieldblanks are treated just like any other samples
i.e. they are also normalized, waterblank subtracted, WSOA scaled/ambient quantified…
FB subtraction for sample j:
But which fieldblank(s) is/are that?
=> Strategy different for every dataset

## Slide 48: Fieldblank subtraction strategy
Check FB characteristics to decide what strategy for FB subtraction to apply
Loadings of WSOC (Shimadzu)
Loadings of WSOA (Shimadzu + OM/OC from AMS)
Loadings of WSOA from AMS (quantified by internal standard method)
Chemical composition of AMS fieldblanks
Properties like AMS derived OM/OC
Factors that might influence/bias FB
Sampling time/operator
Seasons
PM size (if two samplers run in parallel)
Extraction process on different days (contaminations)
…

## Slide 49: Fieldblank subtraction strategy
Example JFJ
Clear seasonal trend in loadings (summer > winter)
No clear trend in chemical composition (similar contamination source)
WSOA

## Slide 50: Fieldblank subtraction strategy
Example JFJ
Clear seasonal trend in loadings (summer > winter)
No trend in chemical composition (similar contamination source)
FB subtraction strategy:
Take consistent relative chemical composition
Scale FB loading to subtract according to sampling season of each sample

## Slide 51: Error propagation of Fieldblank subtraction
It‘s very likely/unavoidable that you have to calculate means and or medians of fieldblank loadings & relative mass spectra
Often, there‘s no analytical solution for error propagation in such cases…
=> Estimation of uncertainties (e.g. in median) by bootstrap resampling

## Slide 52: What do we get from all this?
For each AMS ion, ambient concentration data [ug/m3] and related uncertainties that can go to further analysis, e.g. PMF…. But also many other info required for understanding….
To further analysis (e.g. PMF)
Check correlation & mass balance with external data
Check raw data when inconsistencies at later stages
Use to obtain effective AMS RIEs
Check reproducibility of repetitive QC samples
Background levels and contaminants
Apply for AMS Ion filtering

## Slide 53: What do we get from all this?
For each AMS ion, ambient concentration data [ug/m3] and related uncertainties that can go to further analysis, e.g. PMF
Apply for AMS Ion filtering