## Belading NPL-VSL 29&30-5-2017
| Datum : | 29 en 30 - 05-2017
| Klant : | NPL
| concentratie generator: |  | 2.915 | ng/min
mfc
| flow meten: | 500.2 | ml/min
|  | 500.4 | ml/min
|  | 500.3 | ml/min
|  | 500.3 | ml/min
|  | 500.2 | ml/min
| gemiddelde: | 500.28
| MFM | 92 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 20.015 | l/min | 10.025 | l/min | 10.065 | l/min | 5.9 | l/min | 5.9 | l/min | 10.005 | l/min
| 19.999 | l/min | 10.023 | l/min | 10.053 | l/min | 5.9013 | l/min | 5.9013 | l/min | 10.003 | l/min
| 20.006 | l/min | 10.027 | l/min | 10.054 | l/min | 5.9005 | l/min | 5.9005 | l/min | 10.004 | l/min
| 20.011 | l/min | 10.025 | l/min | 10.063 | l/min | 5.9003 | l/min | 5.9003 | l/min | 10.003 | l/min
| 20.014 | l/min | 10.026 | l/min | 10.064 | l/min | 5.9005 | l/min | 5.9005 | l/min | 10.003 | l/min
flow | 20.009 |  | 10.0252 |  | 10.0598 |  | 5.9005 |  | 5.9005 |  | 10.0036
| positie 1 | 92 | % | 500.28 | ml/min
| positie 2 | 92 | % | 500.28 | ml/min
| positie 3 | 92 | % | 500.28 | ml/min
| positie 4 | 92 | % | 500.28 | ml/min
| positie 5 | 92 | % | 500.28 | ml/min
| positie 6 | 92 | % | 500.28 | ml/min
| Belading 1: | 28 | min. | positie op beladingsrek | buis |  | Belading 2: | 28 | min. | positie op beladingsrek | buis
| positie 1 | 2.0407 | ng/buis | 1 | X186 NPL |  | positie 1 | 4.0731 | ng/buis | 1 | X188 NPL
| positie 2 | 2.0407 | ng/buis | 2 | X187 NPL |  | positie 2 | 4.0731 | ng/buis | 2 | X189 NPL
| positie 3 | 2.0407 | ng/buis | 3 | 2 VSL |  | positie 3 | 4.0731 | ng/buis | 3 | 4 VSL
| positie 4 | 2.0407 | ng/buis | 4 | 5 VSL |  | positie 4 | 4.0731 | ng/buis | 4 | 6 VSL
| positie 5 | 2.0407 | ng/buis | 5 | 7 VSL |  | positie 5 | 4.0731 | ng/buis | 5 | 8 VSL
| positie 6 | 2.0407 | ng/buis | 6 |  |  | positie 6 | 4.0731 | ng/buis | 6
| Belading 3: | 42.5 | min. | positie op beladingsrek | buis |  | Belading 4: | 33.5 | min. | positie op beladingsrek | buis
| positie 1 | 6.161 | ng/buis | 1 | X190 NPL |  | positie 1 | 8.2796 | ng/buis | 1 | X192 NPL
| positie 2 | 6.161 | ng/buis | 2 | X191 NPL |  | positie 2 | 8.2796 | ng/buis | 2 | X193 NPL
| positie 3 | 6.161 | ng/buis | 3 | 2 VSL |  | positie 3 | 8.2796 | ng/buis | 3 | 4 VSL
| positie 4 | 6.161 | ng/buis | 4 | 5 VSL |  | positie 4 | 8.2796 | ng/buis | 4 | 6 VSL
| positie 5 | 6.161 | ng/buis | 5 | 7 VSL |  | positie 5 | 8.2796 | ng/buis | 5 | 8 VSL
| positie 6 | 6.161 | ng/buis | 6 |  |  | positie 6 | 8.2796 | ng/buis | 6
| Belading 5: | 42 | min. | positie op beladingsrek | buis |  | Belading 6: | 42 | min. | positie op beladingsrek | buis
| positie 1 | 10.3804 | ng/buis | 1 | X194 NPL |  | positie 1 | 6.1228 | ng/buis | 1 | X190 NPL
| positie 2 | 10.3804 | ng/buis | 2 | X195 NPL |  | positie 2 | 6.1228 | ng/buis | 2 | X191 NPL
| positie 3 | 10.3804 | ng/buis | 3 | 2 VSL |  | positie 3 | 6.1228 | ng/buis | 3 | 4 VSL
| positie 4 | 10.3804 | ng/buis | 4 | 5 VSL |  | positie 4 | 6.1228 | ng/buis | 4 | 6 VSL
| positie 5 | 10.3804 | ng/buis | 5 | 7 VSL |  | positie 5 | 6.1228 | ng/buis | 5 | 8 VSL
| positie 6 | 10.3804 | ng/buis | 6 |  |  | positie 6 | 6.1228 | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | X186 NPL | X187 NPL | 2 VSL | 5 VSL | 7 VSL | 0
|  | meting ( ng) | x | x | x | 2.027 | 2.2726 | 2.2877 | x
|  |  | buis dummy | X186 NPL | X187 NPL | 2 VSL | 5 VSL | 7 VSL | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | -0.6713 | 11.3631 | 12.103 | #VALUE!
opmerkingen
2 buizen NPL beladen + 3 buizen van onszelf (VSL)
Belading 2
|  | meet volgorde | buis dummy | X188 NPL | X189 NPL | 4 VSL | 6 VSL | 8 VSL | 0
|  | meting ( ng) | x | x | x | 3.666 | 3.7376 | 3.8171 | x
|  |  | buis dummy | X188 NPL | X189 NPL | 4 VSL | 6 VSL | 8 VSL | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | -9.9943 | -8.2351 | -6.2853 | #VALUE!
opmerkingen
2 buizen NPL beladen + 3 buizen van onszelf (VSL)
Belading 3
|  | meet volgorde | buis dummy | X190 NPL | X191 NPL | 2 VSL | 5 VSL | 7 VSL | 0
|  | meting ( ng) |  | 5.6807 | 5.9538 | 4.8215 | 5.572 | 0.02435
|  |  | buis dummy | X190 NPL | X191 NPL | 2 VSL | 5 VSL | 7 VSL | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | -7.7962 | -3.3637 | -21.7428 | -9.5605 | -99.6048 | -100
opmerkingen
2 buizen NPL beladen + 3 buizen van onszelf (VSL)
Belading 4
|  | meet volgorde | buis dummy | X192 NPL | X193 NPL | 4 VSL | 6 VSL | 8 VSL | 0
|  | meting ( ng) |  |  |  | 7.6158 | 7.5297 | 7.5568
|  |  | buis dummy | X192 NPL | X193 NPL | 4 VSL | 6 VSL | 8 VSL | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -8.0174 | -9.0571 | -8.7293 | -100
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | X194 NPL | X195 NPL | 2 VSL | 5 VSL | 7 VSL | 0
|  | meting ( ng) |  |  |  | 8.963 | 9.3109 | 10.7993
|  |  | buis dummy | X194 NPL | X195 NPL | 2 VSL | 5 VSL | 7 VSL | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -13.6546 | -10.3035 | 4.035 | -100
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | X190 NPL | X191 NPL | 4 VSL | 6 VSL | 8 VSL | 0
|  | meting ( ng) |  |  |  | 5.2468 | 5.6129 | 5.709
|  |  | buis dummy | X190 NPL | X191 NPL | 4 VSL | 6 VSL | 8 VSL | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -14.3071 | -8.3279 | -6.758 | -100
opmerkingen
buizen van NPL nog een keer beladen met 6 ng. De eerste keer met 6 ng werd aan getwijfeld, omdat buis 7 helemaal leeg was. Na meten bleek dat ze wel goed waren.(zie meting bij belading 3)

## Results NPL 5-6-2017
Hello all,
The 12 tubes were analysed today:
Tube number | VSL loaded (ng) | NPL measured (ng) | Result normalised | to ng
X186 | 2.0407 | 1.966 | 1.9268 | 2
X187 | 2.0407 | 2.014 | 1.9738 | 2
X188 | 4.0731 | 3.712 | 3.6454 | 4
X189 | 4.0731 | 3.85 | 3.7809 | 4
X190 | 6.1228 | 5.636 | 5.523 | 6
X191 | 6.1228 | 5.901 | 5.7827 | 6
X192 | 8.2796 | 7.642 | 7.3839 | 8
X193 | 8.2796 | 7.573 | 7.3173 | 8
X194 | 10.3804 | 9.171 | 8.8349 | 10
X195 | 10.3804 | 9.255 | 8.9158 | 10
X102 | Blank | -0.009
X107 | Blank | 0.059
These come with an estimated uncertainty of about 10%
Kind regards
Liz

## Belading NPL-VSL 31-7-2017
| Datum : | 2017-07-31
| Klant : | VSL/NPL
| concentratie generator: |  | 2.915 | ng/min
mfc
| flow meten: | 500.1 | ml/min
|  | 500.2 | ml/min
|  | 500.2 | ml/min
|  | 500.2 | ml/min
|  | 500 | ml/min
| gemiddelde: | 500.14
| MFM | 91.3 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6 |  | belading 7
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 6.0234 | l/min | 6.0234 | l/min | 6.0234 | l/min | 6.0234 | l/min | 10.001 | l/min | 10.001 | l/min | 14.981 | l/min
| 6.0216 | l/min | 6.0216 | l/min | 6.0216 | l/min | 6.0216 | l/min | 10.004 | l/min | 10.004 | l/min | 14.989 | l/min
| 6.0221 | l/min | 6.0221 | l/min | 6.0221 | l/min | 6.0221 | l/min | 9.9989 | l/min | 9.9989 | l/min | 14.973 | l/min
|  | l/min |  | l/min |  | l/min |  | l/min | 10.002 | l/min | 10.002 | l/min | 14.981 | l/min
|  | l/min |  | l/min |  | l/min |  | l/min | 10.006 | l/min | 10.006 | l/min | 14.99 | l/min
flow | 6.0224 |  | 6.0224 |  | 6.0224 |  | 6.0224 |  | 10.0024 |  | 10.0024 |  | 14.9828
| positie 1 | 91.3 | % | 500.14 | ml/min
| positie 2 | 91.3 | % | 500.14 | ml/min
| positie 3 | 91.3 | % | 500.14 | ml/min
| positie 4 | 91.3 | % | 500.14 | ml/min
| positie 5 | 91.3 | % | 500.14 | ml/min
| positie 6 | 91.3 | % | 500.14 | ml/min
| Belading 1: | 33 | min. | positie op beladingsrek | buis |  | Belading 2: | 33 | min. | positie op beladingsrek | buis
| positie 1 | 7.9888 | ng/buis | 1 | 16 VSL |  | positie 1 | 7.9888 | ng/buis | 1 | 10 VSL
| positie 2 | 7.9888 | ng/buis | 2 | 17 VSL |  | positie 2 | 7.9888 | ng/buis | 2 | 11 VSL
| positie 3 | 7.9888 | ng/buis | 3 | 18 VSL |  | positie 3 | 7.9888 | ng/buis | 3 | 12 VSL
| positie 4 | 7.9888 | ng/buis | 4 | 19 VSL |  | positie 4 | 7.9888 | ng/buis | 4 | 13 VSL
| positie 5 | 7.9888 | ng/buis | 5 | 1 VSL |  | positie 5 | 7.9888 | ng/buis | 5 | 14 VSL
| positie 6 | 7.9888 | ng/buis | 6 | 20 VSL |  | positie 6 | 7.9888 | ng/buis | 6 | 15 VSL
| Belading 3: | 42 | min. | positie op beladingsrek | buis |  | Belading 4: | 34 | min. | positie op beladingsrek | buis
| positie 1 | 10.1675 | ng/buis | 1 | X194 NPL |  | positie 1 | 8.2309 | ng/buis | 1 | X192 NPL
| positie 2 | 10.1675 | ng/buis | 2 | X195 NPL |  | positie 2 | 8.2309 | ng/buis | 2 | X193 NPL
| positie 3 | 10.1675 | ng/buis | 3 | 2 VSL |  | positie 3 | 8.2309 | ng/buis | 3 | 8 VSL
| positie 4 | 10.1675 | ng/buis | 4 | 4 VSL |  | positie 4 | 8.2309 | ng/buis | 4 | 9 VSL
| positie 5 | 10.1675 | ng/buis | 5 | 1 VSL |  | positie 5 | 8.2309 | ng/buis | 5 | 3 VSL
| positie 6 | 10.1675 | ng/buis | 6 | 6 VSL |  | positie 6 | 8.2309 | ng/buis | 6 | 10 VSL
| Belading 5: | 42 | min. | positie op beladingsrek | buis |  | Belading 6: | 28 | min. | positie op beladingsrek | buis |  | Belading 7: | 21 | min. | positie op beladingsrek | buis
| positie 1 | 6.1218 | ng/buis | 1 | X190 NPL |  | positie 1 | 4.0812 | ng/buis | 1 | X188 NPL |  | positie 1 | 2.0434 | ng/buis | 1 | X186 NPL
| positie 2 | 6.1218 | ng/buis | 2 | X191 NPL |  | positie 2 | 4.0812 | ng/buis | 2 | X189 NPL |  | positie 2 | 2.0434 | ng/buis | 2 | X187 NPL
| positie 3 | 6.1218 | ng/buis | 3 | 11 VSL |  | positie 3 | 4.0812 | ng/buis | 3 | 14 VSL |  | positie 3 | 2.0434 | ng/buis | 3 | 17 VSL
| positie 4 | 6.1218 | ng/buis | 4 | 12 VSL |  | positie 4 | 4.0812 | ng/buis | 4 | 15 VSL |  | positie 4 | 2.0434 | ng/buis | 4 | 18 VSL
| positie 5 | 6.1218 | ng/buis | 5 | 5 VSL |  | positie 5 | 4.0812 | ng/buis | 5 | 7 VSL |  | positie 5 | 2.0434 | ng/buis | 5 | 19 VSL
| positie 6 | 6.1218 | ng/buis | 6 | 13 VSL |  | positie 6 | 4.0812 | ng/buis | 6 | 16 VSL |  | positie 6 | 2.0434 | ng/buis | 6 | 20 VSL
Belading 1
|  | meet volgorde | buis dummy | 16 VSL | 17 VSL | 18 VSL | 19 VSL | 1 VSL | 20 VSL
|  | meting ( ng) | 0.02867 | 0.97181 | 1.451 | 2.0405 | 2.4731 | 0.09655 | 2.3193
|  |  | buis dummy | 16 VSL | 17 VSL | 18 VSL | 19 VSL | 1 VSL | 20 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | -87.8353 | -81.8368 | -74.4582 | -69.0427 | -98.7914 | -70.9681
opmerkingen
6x VSL buisjes beladen als test van de nieuwe buisjes. Begonnen met buis 1 en daarna 16-20
Belading 2
|  | meet volgorde | buis dummy | 10 VSL | 11 VSL | 12 VSL | 13 VSL | 14 VSL | 15 VSL
|  | meting ( ng) |  | 2.6138 | 7.85 | 9.0017 | 9.0245 | 8.422 | 7.5572
|  |  | buis dummy | 10 VSL | 11 VSL | 12 VSL | 13 VSL | 14 VSL | 15 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | -67.2815 | -1.7369 | 12.6797 | 12.9651 | 5.4223 | -5.402
opmerkingen
6x VSL buisjes beladen als test van de nieuwe buisjes. Na buis 10 is de tubing vervangen, en is het 'stoffilter' ertussenuit gehaald
Belading 3
|  | meet volgorde | buis dummy | X194 NPL | X195 NPL | 2 VSL | 4 VSL | 1 VSL | 6 VSL
|  | meting ( ng) |  | x | x | 9.0724 | 7.8642 | x | 8.051
|  |  | buis dummy | X194 NPL | X195 NPL | 2 VSL | 4 VSL | 1 VSL | 6 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | -10.7709 | -22.6536 | #VALUE! | -20.8166
opmerkingen
Gemeten 2017-09-06
Belading 4
|  | meet volgorde | buis dummy | X192 NPL | X193 NPL | 8 VSL | 9 VSL | 3 VSL | 10 VSL
|  | meting ( ng) |  | x | x | 8.3911 | 8.6371 | x | 9.3642
|  |  | buis dummy | X192 NPL | X193 NPL | 8 VSL | 9 VSL | 3 VSL | 10 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | 1.9464 | 4.9361 | #VALUE! | 13.7696
opmerkingen
Gemeten 2017-09-06
Belading 5
|  | meet volgorde | buis dummy | X190 NPL | X191 NPL | 11 VSL | 12 VSL | 5 VSL | 13 VSL
|  | meting ( ng) |  | x | x | 5.6739 | 4.9662 | x | 4.1234
|  |  | buis dummy | X190 NPL | X191 NPL | 11 VSL | 12 VSL | 5 VSL | 13 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | -7.3167 | -18.8775 | #VALUE! | -32.6445
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | X188 NPL | X189 NPL | 14 VSL | 15 VSL | 7 VSL | 16 VSL
|  | meting ( ng) |  | x | x | kapot | 2.5755 | x | 3.4407
|  |  | buis dummy | X188 NPL | X189 NPL | 14 VSL | 15 VSL | 7 VSL | 16 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | #VALUE! | -36.8933 | #VALUE! | -15.6947
opmerkingen
Belading 7
|  | meet volgorde | buis dummy | X186 NPL | X187 NPL | 17 VSL | 18 VSL | 19 VSL | 20 VSL
|  | meting ( ng) |  | x | x | 2.7433 | 2.2221 | 2.1868 | 2.234
|  |  | buis dummy | X186 NPL | X187 NPL | 17 VSL | 18 VSL | 19 VSL | 20 VSL
afwijking t.o.v theoretisch berekend (%) |  |  | x | #VALUE! | #VALUE! | 34.2493 | 8.7422 | 7.0181 | 9.3236
opmerkingen

## Results NPL 31-7-2017
Morning Ruben and all,
Here are the results from round 2 which were analysed on 4th August:
Tube number | VSL loaded | NPL measured | Result normalised | to ng
| (ng) | (ng)
X186 | 2.0434 | 1.941 | 1.8997 | 2
X187 | 2.0434 | 1.926 | 1.8851 | 2
X188 | 4.0812 | 3.743 | 3.6685 | 4
X189 | 4.0812 | 3.771 | 3.696 | 4
X190 | 6.1218 | 6.064 | 5.9433 | 6
X191 | 6.1218 | 6.131 | 6.009 | 6
X192 | 8.2309 | 7.412 | 7.2041 | 8
X193 | 8.2309 | 7.349 | 7.1429 | 8
X194 | 10.1675 | 9.372 | 9.2176 | 10
X195 | 10.1675 | 9.14 | 8.9894 | 10
X102 | Blank | -0.033
X107 | Blank | 0.09
Shall I start the motions for posting them back to you, or hang on?
Thanks,
Liz

## Belading NPL-VSL 6-9-2017
| Datum : | 2017-09-06
| Klant : | VSL
| concentratie generator: |  | 2.915 | ng/min
mfc
| flow meten: | 499.87 | ml/min
|  | 499.87 | ml/min
|  | 499.79 | ml/min
|  | 499.79 | ml/min
|  | 499.98 | ml/min
| gemiddelde: | 499.86
| MFM | 91.2 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 6.019 | l/min | 6.019 | l/min | 6.019 | l/min | 6.019 | l/min |  | l/min |  | l/min
| 6.0262 | l/min | 6.0262 | l/min | 6.0262 | l/min | 6.0262 | l/min |  | l/min |  | l/min
| 6.0137 | l/min | 6.0137 | l/min | 6.0137 | l/min | 6.0137 | l/min |  | l/min |  | l/min
| 6.0266 | l/min | 6.0266 | l/min | 6.0266 | l/min | 6.0266 | l/min |  | l/min |  | l/min
| 6.0204 | l/min | 6.0204 | l/min | 6.0204 | l/min | 6.0204 | l/min |  | l/min |  | l/min
flow | 6.0212 |  | 6.0212 |  | 6.0212 |  | 6.0212 |  | #DIV/0! |  | #DIV/0!
| positie 1 | 91.2 | % | 499.86 | ml/min
| positie 2 | 91.2 | % | 499.86 | ml/min
| positie 3 | 91.2 | % | 499.86 | ml/min
| positie 4 | 91.2 | % | 499.86 | ml/min
| positie 5 | 91.2 | % | 499.86 | ml/min
| positie 6 | 91.2 | % | 499.86 | ml/min
| Belading 1: | 34 | min. | positie op beladingsrek | buis |  | Belading 2: | 42 | min. | positie op beladingsrek | buis
| positie 1 | 8.2279 | ng/buis | 1 | 3 |  | positie 1 | 10.1638 | ng/buis | 1 | 5
| positie 2 | 8.2279 | ng/buis | 2 | 21 |  | positie 2 | 10.1638 | ng/buis | 2 | 7
| positie 3 | 8.2279 | ng/buis | 3 | 22 |  | positie 3 | 10.1638 | ng/buis | 3 | 26
| positie 4 | 8.2279 | ng/buis | 4 | 23 |  | positie 4 | 10.1638 | ng/buis | 4 | 27
| positie 5 | 8.2279 | ng/buis | 5 | 24 |  | positie 5 | 10.1638 | ng/buis | 5 | 28
| positie 6 | 8.2279 | ng/buis | 6 | 25 |  | positie 6 | 10.1638 | ng/buis | 6 | 30
| Belading 3: | 34 | min. | positie op beladingsrek | buis |  | Belading 4: | 42 | min. | positie op beladingsrek | buis
| positie 1 | 8.2279 | ng/buis | 1 | NPL X192 |  | positie 1 | 10.1638 | ng/buis | 1 | NPL X194
| positie 2 | 8.2279 | ng/buis | 2 | NPL X193 |  | positie 2 | 10.1638 | ng/buis | 2 | NPL X195
| positie 3 | 8.2279 | ng/buis | 3 | VSL 8 |  | positie 3 | 10.1638 | ng/buis | 3 | VSL 2
| positie 4 | 8.2279 | ng/buis | 4 | VSL 9 |  | positie 4 | 10.1638 | ng/buis | 4 | VSL 4
| positie 5 | 8.2279 | ng/buis | 5 | VSL 3 |  | positie 5 | 10.1638 | ng/buis | 5 | VSL 5
| positie 6 | 8.2279 | ng/buis | 6 | VSL 10 |  | positie 6 | 10.1638 | ng/buis | 6 | VSL 6
| Belading 5: | 42 | min. | positie op beladingsrek | buis |  | Belading 6: | 42 | min. | positie op beladingsrek | buis
| positie 1 | #DIV/0! | ng/buis | 1 |  |  | positie 1 | #DIV/0! | ng/buis | 1
| positie 2 | #DIV/0! | ng/buis | 2 |  |  | positie 2 | #DIV/0! | ng/buis | 2
| positie 3 | #DIV/0! | ng/buis | 3 |  |  | positie 3 | #DIV/0! | ng/buis | 3
| positie 4 | #DIV/0! | ng/buis | 4 |  |  | positie 4 | #DIV/0! | ng/buis | 4
| positie 5 | #DIV/0! | ng/buis | 5 |  |  | positie 5 | #DIV/0! | ng/buis | 5
| positie 6 | #DIV/0! | ng/buis | 6 |  |  | positie 6 | #DIV/0! | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | 3 | 21 | 22 | 23 | 24 | 25
|  | meting ( ng) |  | 7.0799 | 7.8328 | 7.9926 | 8.0966 | 7.8678 | 7.867
|  |  | buis dummy | 3 | 21 | 22 | 23 | 24 | 25
afwijking t.o.v theoretisch berekend (%) |  |  | x | -13.9519 | -4.8014 | -2.8598 | -1.5952 | -4.3759 | -4.3864
opmerkingen
Calibratie met 200 400 800 1000 ul @ 18 C
Belading 2
|  | meet volgorde | buis dummy | 5 | 7 | 26 | 27 | 28 | 30
|  | meting ( ng) |  | 9.3571 | 4.357 | 11.3827 | 9.6535 | 8.9207 | 9.2349
|  |  | buis dummy | 5 | 7 | 26 | 27 | 28 | 30
afwijking t.o.v theoretisch berekend (%) |  |  | x | -7.9374 | -57.1324 | 11.9924 | -5.021 | -12.2305 | -9.1399
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | NPL X192 | NPL X193 | VSL 8 | VSL 9 | VSL 3 | VSL 10
|  | meting ( ng) |  |  |  | 6.0371 | 3.6158 | 6.702 | 4.1058
|  |  | buis dummy | NPL X192 | NPL X193 | VSL 8 | VSL 9 | VSL 3 | VSL 10
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -26.6265 | -56.0542 | -18.5451 | -50.0992
opmerkingen
Buis 10 nog 2 x gemeten, om memory effect vast te stellen: 1,00239 ng & 0,19621 ng. Hierna buis 9 ook nog eens gemeten: 0,11185 ng, en buis 8: 0,08001 ng --> NPL buisjes schoongemaakt mbv PSA trap cleaner, en alles opnieuw beladen.
Belading 4
|  | meet volgorde | buis dummy | NPL X194 | NPL X195 | VSL 2 | VSL 4 | VSL 5 | VSL 6
meting ( ng)
|  |  | buis dummy | NPL X194 | NPL X195 | VSL 2 | VSL 4 | VSL 5 | VSL 6
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | -100 | -100
opmerkingen
Buis 5 op positie 5, maar als 1e gemeten. Daarna 21-25
Belading 5
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen

## Belading NPL-VSL 7-9-2017
| Datum : | 2017-09-07
| Klant : | VSL/NPL
| concentratie generator: |  | 2.915 | ng/min
mfc
| flow meten: | 499.87 | ml/min
|  | 500.1 | ml/min
|  | 499.87 | ml/min
|  | 499.79 | ml/min
|  | 499.98 | ml/min
| gemiddelde: | 499.922
| MFM | 91.4 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 5.9931 | l/min | 9.9918 | l/min | 9.9918 | l/min | 15.022 | l/min |  | l/min |  | l/min
| 5.9945 | l/min | 9.9876 | l/min | 9.9876 | l/min | 15.031 | l/min |  | l/min |  | l/min
| 5.9956 | l/min | 9.9884 | l/min | 9.9884 | l/min | 15.043 | l/min |  | l/min |  | l/min
| 5.9975 | l/min | 9.9902 | l/min | 9.9902 | l/min | 15.035 | l/min |  | l/min |  | l/min
| 5.9953 | l/min | 9.9895 | l/min | 9.9895 | l/min | 15.059 | l/min |  | l/min |  | l/min
flow | 5.9952 |  | 9.9895 |  | 9.9895 |  | 15.038 |  | #DIV/0! |  | #DIV/0!
| positie 1 | 91.4 | % | 499.922 | ml/min
| positie 2 | 91.4 | % | 499.922 | ml/min
| positie 3 | 91.4 | % | 499.922 | ml/min
| positie 4 | 91.4 | % | 499.922 | ml/min
| positie 5 | 91.4 | % | 499.922 | ml/min
| positie 6 | 91.4 | % | 499.922 | ml/min
| Belading 1: | 34 | min. | positie op beladingsrek | buis |  | Belading 2: | 42 | min. | positie op beladingsrek | buis
| positie 1 | 8.2646 | ng/buis | 1 | NPL X192 |  | positie 1 | 6.127 | ng/buis | 1 | X190 NPL
| positie 2 | 8.2646 | ng/buis | 2 | NPL X193 |  | positie 2 | 6.127 | ng/buis | 2 | X191 NPL
| positie 3 | 8.2646 | ng/buis | 3 | VSL 8 |  | positie 3 | 6.127 | ng/buis | 3 | VSL 21
| positie 4 | 8.2646 | ng/buis | 4 | VSL 9 |  | positie 4 | 6.127 | ng/buis | 4 | VSL 22
| positie 5 | 8.2646 | ng/buis | 5 | VSL 1 |  | positie 5 | 6.127 | ng/buis | 5 | VSL 3
| positie 6 | 8.2646 | ng/buis | 6 | VSL 10 |  | positie 6 | 6.127 | ng/buis | 6 | VSL 23
| Belading 3: | 28 | min. | positie op beladingsrek | buis |  | Belading 4: | 21 | min. | positie op beladingsrek | buis
| positie 1 | 4.0847 | ng/buis | 1 | X188 NPL |  | positie 1 | 2.035 | ng/buis | 1 | X186 NPL
| positie 2 | 4.0847 | ng/buis | 2 | X189 NPL |  | positie 2 | 2.035 | ng/buis | 2 | X187 NPL
| positie 3 | 4.0847 | ng/buis | 3 | VSL 24 |  | positie 3 | 2.035 | ng/buis | 3 | VSL 27
| positie 4 | 4.0847 | ng/buis | 4 | VSL 25 |  | positie 4 | 2.035 | ng/buis | 4 | VSL 28
| positie 5 | 4.0847 | ng/buis | 5 | VSL 29 |  | positie 5 | 2.035 | ng/buis | 5 | VSL 7
| positie 6 | 4.0847 | ng/buis | 6 | VSL 26 |  | positie 6 | 2.035 | ng/buis | 6 | VSL 30
| Belading 5: | 42 | min. | positie op beladingsrek | buis |  | Belading 6: | 42 | min. | positie op beladingsrek | buis
| positie 1 | #DIV/0! | ng/buis | 1 |  |  | positie 1 | #DIV/0! | ng/buis | 1
| positie 2 | #DIV/0! | ng/buis | 2 |  |  | positie 2 | #DIV/0! | ng/buis | 2
| positie 3 | #DIV/0! | ng/buis | 3 |  |  | positie 3 | #DIV/0! | ng/buis | 3
| positie 4 | #DIV/0! | ng/buis | 4 |  |  | positie 4 | #DIV/0! | ng/buis | 4
| positie 5 | #DIV/0! | ng/buis | 5 |  |  | positie 5 | #DIV/0! | ng/buis | 5
| positie 6 | #DIV/0! | ng/buis | 6 |  |  | positie 6 | #DIV/0! | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | NPL X192 | NPL X193 | VSL 8 | VSL 9 | VSL 1 | VSL 10
|  | meting ( ng) |  |  |  | 8.06 | 6.8385 |  | 6.5049
|  |  | buis dummy | NPL X192 | NPL X193 | VSL 8 | VSL 9 | VSL 1 | VSL 10
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -2.4746 | -17.2555 | -100 | -21.2914
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | X190 NPL | X191 NPL | VSL 21 | VSL 22 | VSL 3 | VSL 23
meting ( ng)
|  |  | buis dummy | X190 NPL | X191 NPL | VSL 21 | VSL 22 | VSL 3 | VSL 23
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | -100 | -100
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | X188 NPL | X189 NPL | VSL 24 | VSL 25 | VSL 29 | VSL 26
meting ( ng)
|  |  | buis dummy | X188 NPL | X189 NPL | VSL 24 | VSL 25 | VSL 29 | VSL 26
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | -100 | -100
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | X186 NPL | X187 NPL | VSL 27 | VSL 28 | VSL 7 | VSL 30
meting ( ng)
|  |  | buis dummy | X186 NPL | X187 NPL | VSL 27 | VSL 28 | VSL 7 | VSL 30
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | -100 | -100
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen

## Results NPL 12-9-2017
Hello all,
Here are the results of the analysis of the traps in round 3, analysed yesterday (12/09/17)
Tube number | VSL loaded (ng) | NPL measured (ng) | Result normalised | to ng
X186 | 2.035 | 1.86 | 1.828 | 2
X187 | 2.035 | 1.886 | 1.8535 | 2
X188 | 4.0847 | 3.538 | 3.4647 | 4
X189 | 4.0847 | 3.615 | 3.5401 | 4
X190 | 6.127 | 5.428 | 5.3155 | 6
X191 | 6.127 | 5.292 | 5.1823 | 6
X192 | 8.2646 | 7.55 | 7.3083 | 8
X193 | 8.2646 | 7.471 | 7.2319 | 8
X194 | 10.1638 | 9.277 | 9.1275 | 10
X195 | 10.1638 | 9.128 | 8.9809 | 10
X102 | Blank | 0.071
X107 | Blank | 0.035
Kind regards
Liz

## Summary
2 ng
Round | Tube | Entry | Reference value | Uncertainty | Verification | Unceratainty | Difference
nr | nr |  | ng | ng | ng | ng
1 | 2 VSL | R1 T1 VSL | 2.0407 | 0.063263 | 2.027 | 0.2027 | -0.6758%
1 | 5 VSL | R1 T2 VSL | 2.0407 | 0.063263 | 2.2726 | 0.22726 | 10.2%
1 | 7 VSL | R1 T3 VSL | 2.0407 | 0.063263 | 2.2877 | 0.22877 | 10.8% | VSL | 2.1958 | 0.21958
1 | X186 NPL | R1 T1 NPL | 2.0407 | 0.063263 | 1.966 | 0.1966 | -3.802%
1 | X187 NPL | R1 T2 NPL | 2.0407 | 0.063263 | 2.014 | 0.2014 | -1.328% | NPL | 1.99 | 0.199
2 | X186 NPL | R2 T1 NPL | 2.0434 | 0.063346 | 1.941 | 0.1941 | -5.277%
2 | X187 NPL | R2 T2 NPL | 2.0434 | 0.063346 | 1.926 | 0.1926 | -6.097%
3 | X186 NPL | R3 T1 NPL | 2.035 | 0.063086 | 1.86 | 0.186 | -9.411%
3 | X187 NPL | R3 T2 NPL | 2.035 | 0.063086 | 1.886 | 0.1886 | -7.903% | NPL | 1.9322 | 0.19322
4 ng
Round | Tube |  | Reference value | Uncertainty | Verification | Unceratainty | Difference
nr | nr |  | ng | ng | ng | ng
1 | 4 VSL | R1 T1 VSL | 4.0731 | 0.12626 | 3.666 | 0.3666 | -11.1%
1 | 6 VSL | R1 T2 VSL | 4.0731 | 0.12626 | 3.7376 | 0.37376 | -8.974%
1 | 8 VSL | R1 T3 VSL | 4.0731 | 0.12626 | 3.8171 | 0.38171 | -6.707% | VSL | 3.7402 | 0.37402
1 | X188 NPL | R1 T1 NPL | 4.0731 | 0.12626 | 3.712 | 0.3712 | -9.727%
1 | X189 NPL | R1 T2 NPL | 4.0731 | 0.12626 | 3.85 | 0.385 | -5.794% | NPL | 3.781 | 0.3781
2 | X188 NPL | R2 T1 NPL | 4.0812 | 0.12652 | 3.743 | 0.3743 | -9.036%
2 | X189 NPL | R2 T2 NPL | 4.0812 | 0.12652 | 3.771 | 0.3771 | -8.226%
3 | X188 NPL | R3 T1 NPL | 4.0847 | 0.12663 | 3.538 | 0.3538 | -15.45%
3 | X189 NPL | R3 T2 NPL | 4.0847 | 0.12663 | 3.615 | 0.3615 | -12.99% | NPL | 3.7048 | 0.37048
6 ng
Round | Tube |  | Reference value | Uncertainty | Verification | Unceratainty | Difference
nr | nr |  | ng | ng | ng | ng
1 | 4 VSL | R1 T1 VSL | 6.1228 | 0.18981 | 5.2468 | 0.52468 | -16.7%
1 | 6 VSL | R1 T2 VSL | 6.1228 | 0.18981 | 5.6129 | 0.56129 | -9.084%
1 | 8 VSL | R1 T3 VSL | 6.1228 | 0.18981 | 5.709 | 0.5709 | -7.248% | VSL | 5.5229 | 0.55229
1 | X190 NPL | R1 T1 NPL | 6.1228 | 0.18981 | 5.636 | 0.5636 | -8.637%
1 | X191 NPL | R1 T2 NPL | 6.1228 | 0.18981 | 5.901 | 0.5901 | -3.758% | NPL | 5.7685 | 0.57685
2 | X190 NPL | R2 T1 NPL | 6.1218 | 0.18978 | 6.064 | 0.6064 | -0.9532%
2 | X191 NPL | R2 T2 NPL | 6.1218 | 0.18978 | 6.131 | 0.6131 | 0.15%
3 | X190 NPL | R3 T1 NPL | 6.127 | 0.18994 | 5.428 | 0.5428 | -12.88%
3 | X191 NPL | R3 T2 NPL | 6.127 | 0.18994 | 5.292 | 0.5292 | -15.78% | NPL | 5.742 | 0.5742
8 ng
Round | Tube |  | Reference value | Uncertainty | Verification | Unceratainty | Difference
nr | nr |  | ng | ng | ng | ng
1 | 4 VSL | VSL 1.1 | 8.2796 | 0.25667 | 7.6158 | 0.76158 | -8.716%
1 | 6 VSL | VSL 1.2 | 8.2796 | 0.25667 | 7.5297 | 0.75297 | -9.959%
1 | 8 VSL | VSL 1.3 | 8.2796 | 0.25667 | 7.5568 | 0.75569 | -9.564% | VSL | 7.5675 | 0.75675
1 | X192 NPL | NPL 1.1 | 8.2796 | 0.25667 | 7.642 | 0.7642 | -8.343%
1 | X193 NPL | NPL 1.2 | 8.2796 | 0.25667 | 7.573 | 0.7573 | -9.331% | NPL | 7.6075 | 0.76075
2 | X192 NPL | NPL 2.1 | 8.2309 | 0.25516 | 7.412 | 0.7412 | -11.05%
2 | X193 NPL | NPL 2.2 | 8.2309 | 0.25516 | 7.349 | 0.7349 | -12%
3 | NPL X192 | NPL 3.1 | 8.2646 | 0.2562 | 7.55 | 0.755 | -9.464%
3 | NPL X193 | NPL 3.2 | 8.2646 | 0.2562 | 7.471 | 0.7471 | -10.62% | NPL | 7.4995 | 0.74995
10 ng
Round | Tube |  | Reference value | Uncertainty | Verification | Unceratainty | Difference
nr | nr |  | ng | ng | ng | ng
1 | 2 VSL | R1 T1 VSL | 10.3804 | 0.32179 | 8.963 | 0.8963 | -15.81%
1 | 5 VSL | R1 T2 VSL | 10.3804 | 0.32179 | 9.3109 | 0.93109 | -11.49%
1 | 7 VSL | R1 T3 VSL | 10.3804 | 0.32179 | 10.7993 | 1.0799 | 3.879% | VSL | 9.691 | 0.9691
1 | X194 NPL | R1 T1 NPL | 10.3804 | 0.32179 | 9.171 | 0.9171 | -13.19%
1 | X195 NPL | R1 T2 NPL | 10.3804 | 0.32179 | 9.255 | 0.9255 | -12.16% | NPL | 9.213 | 0.9213
2 | X194 NPL | R2 T1 NPL | 10.1675 | 0.31519 | 9.372 | 0.9372 | -8.488%
2 | X195 NPL | R2 T2 NPL | 10.1675 | 0.31519 | 9.14 | 0.914 | -11.24%
3 | NPL X194 | R3 T1 NPL | 10.1638 | 0.31508 | 9.277 | 0.9277 | -9.56%
3 | NPL X195 | R3 T2 NPL | 10.1638 | 0.31508 | 9.128 | 0.9128 | -11.35% | NPL | 9.2238 | 0.92238
[5 empty] | Average difference |  | -7.885% |  | 5.91%
[5 empty] | VSL difference |  | -6.077% |  | 8.424%
[5 empty] | NPL difference |  | -8.79% |  | 4.024%
[5 empty] | 2 ng difference |  | -4.928% |  | 3.234%
[5 empty] | 4 ng difference |  | -9.779% |  | 3.025%
[5 empty] | 6 ng difference |  | -8.32% |  | 6.075%
[5 empty] | 8 ng difference |  | -9.894% |  | 1.157%
[5 empty] | 10 ng difference |  | -9.934% |  | 5.582%

## Uncertainty
|  |  |  | x(Hg) = (Q18 + Q19 +Q20) | [10 empty] | Q(Hg) = Q18 + Q19 +Q20
Uncertainty mercury concentration in the generator | [10 empty] | Uncertainty sum diffusion
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y) |  | X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
Q 18 | ng/min |  |  | #DIV/0! | 1 | 0 | 1/d | 0.16609 | 0 | a | Q 18 | ng/min | 0 | 0 | #DIV/0! | 1 | 0 | 1 | 1 | 0 | a
Q 19 | ng/min | 1.4767 | 0.031011 | 2.1% | 1 | 0.031011 | 1/d | 0.16609 | 0.0051505 | b | Q 19 | ng/min | 1.4767 | 0.031011 | 2.1% | 1 | 0.031011 | 1 | 1 | 0.031011 | b
Q 20 | ng/min | 1.4383 | 0.030204 | 2.1% | 1 | 0.030204 | 1/d | 0.16609 | 0.0050165 | c | Q 20 | ng/min | 1.4383 | 0.030204 | 2.1% | 1 | 0.030204 | 1 | 1 | 0.030204 | c
MFC1 | L/min | 0.5 | 0.0015 | 0.3% | 1 | 0.0015 |  |  |  |  | Q | ng/min | 2.915 | 0.086579 | 2.97% | 2 | 0.04329 | 1
MFC2 | L/min | 5.5 | 0.0165 | 0.3% | 1 | 0.0165
Bios | L/min | 6.021 | 0.012042 | 0.2% | 2 | 0.006021
ρ air | L/min | 6.021 | 0.017628 | 0.2928% | 1 | 0.017628 | -((a+b+c)/d^2) | -0.080409 | -0.0014175 | d
x(Hg) | ng/m3 | 484.1424 | 14.6563 | 3.027% | 2 | 7.3282
Uncertainty mercury mass on transfer standards
m(Hg) = (t* ρ sample) * Q/ρ N2
Uncertainty mercury mass on transfer standards
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
t | min | 42 | 0.03 | 0.07143% | 2 | 0.015 | c(d+e+f)/g | 0.24202 | 0.0036303 | b
ρ sample | L/min | 0.4999 | 0.0017497 | 0.5% | 1 | 0.0017497 | b(d+e+f)/g | 20.334 | 0.035577 | c
MFC | L/min | 0.4999 | 0.0014997 | 0.3% | 1 | 0.0014997
MFM | L/min | 0.4999 | 0.0009998 | 0.2% | 2 | 0.0004999
MPM | L/min | 0.4999 | 0.0014997 | 0.3% | 2 | 0.00074985
Q 18 | ng/min | 0 | 0 | #DIV/0! | 1 | 0 | bc/g | 3.4871 | 0 | d
Q 19 | ng/min | 1.4767 | 0.031011 | 2.1% | 1 | 0.031011 | bc/g | 3.4871 | 0.10814 | e
Q 20 | ng/min | 1.4383 | 0.030204 | 2.1% | 1 | 0.030204 | bc/g | 3.4871 | 0.10533 | f
ρ N2 | L/min | 6.021 | 0.017628 | 0.2928% | 1 | 0.017628 | -bc(f+e+d)/g^2 | -1.6883 | -0.029761 | g
m(Hg) | ng | 10.165 | 0.31592 | 3.108% | 2 | 0.15796
m(Hg) = (t* ρ sample) * x(Hg)/1000
Uncertainty mercury mass on transfer standards
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
t | min | 42 | 0.03 | 0.07143% | 2 | 0.015 | bc/1000 | 0.24202 | 0.0036303 | a
ρ sample | ml/min | 0.4999 | 0.0017497 | 0.5% | 1 | 0.0017497 | ac/1000 | 20.334 | 0.035577 | b
MFC | ml/min | 0.4999 | 0.0014997 | 0.3% | 1 | 0.0014997
MFM | ml/min | 0.4999 | 0.0009998 | 0.2% | 2 | 0.0004999
MPM | ml/min | 0.4999 | 0.0014997 | 0.3% | 2 | 0.00074985
x(Hg) | ng/m3 | 484.1424 | 14.6563 | 3.027% | 2 | 7.3282 | ab/1000 | 0.020996 | 0.15386 | c
m(Hg) | ng | 10.165 | 0.31592 | 3.108% | 2 | 0.15796
Note on A19: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the timer used is 0.03 minutes for a time interval of 0.1 to 100 minutes with a coverage factor of k = 2. For a loading of 25 minutes this leads to a standard uncertainty of 0.0012 % with a coverage factor of k = 2, which is negligible compared to other uncertainty contributions.
Note on A21: Krom, Iris (I.M.J.) de, Mrs: Mass-flow controller output variations are stated by the manufacturer to be < 0.1 % temperature sensitivity of full scale per K of the reading. The uncertainty of mass-flow controller output at full scale, at a temperature uncertainty of 1 K, is equal to (0.1*0.3)2 = 0.03 %. The uncertainty in drift of the mass-flow controller is  0.3 %. By keeping the mass-flow controller set point above 20 %, uMFC  0.3 %.
Note on A22: Krom, Iris (I.M.J.) de, Mrs: The reading of the MFM (mass flow meter) is within ± 0.05 (% set point), resulting in a reading uncertainty of 0.029 (% set point). By keeping the set point above 40%, the relative reading uncertainty is kept below 0.07%. Manufacturers’ specifications of the repeatability of the MFM is < 0.2% of reading. The resulting uncertainty contribution UMFM is 0.2% (k = 2).
Note on A23: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the MPM (mercury piston prover) used is 0.3 % for flow rates from 20 up to 1000 mL.min-1 with a coverage factor of k = 2, UMPM = 0.3 %.
Note on A35: Krom, Iris (I.M.J.) de, Mrs: Mass-flow controller output variations are stated by the manufacturer to be < 0.1 % temperature sensitivity of full scale per K of the reading. The uncertainty of mass-flow controller output at full scale, at a temperature uncertainty of 1 K, is equal to (0.1*0.3)2 = 0.03 %. The uncertainty in drift of the mass-flow controller is  0.3 %. By keeping the mass-flow controller set point above 20 %, uMFC  0.3 %.
Note on A36: Krom, Iris (I.M.J.) de, Mrs: The reading of the MFM (mass flow meter) is within ± 0.05 (% set point), resulting in a reading uncertainty of 0.029 (% set point). By keeping the set point above 40%, the relative reading uncertainty is kept below 0.07%. Manufacturers’ specifications of the repeatability of the MFM is < 0.2% of reading. The resulting uncertainty contribution UMFM is 0.2% (k = 2).
Note on A37: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the MPM (mercury piston prover) used is 0.3 % for flow rates from 20 up to 1000 mL.min-1 with a coverage factor of k = 2, UMPM = 0.3 %.

## ANOVA
2 ng
[16 empty] | ISO 5725-2, ongecorrigeerd | [9 empty] | ISO 5725-2, gecorrigeerd voor vlaggen
date | value #1 | value #2 | value #3 | value #4 | value #5 | value#6 | value#7 | value#8 | value#9 | value#10 | mean | s | n | flag | Z |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2 | AD | s^2 |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2
2020-05-29 | 1.9268 | 1.9738 | [8 empty] | 1.9503 | 0.033264 | 2 |  | 0.97741 |  | 2 | 3.9005 | 7.6071 | 4 | 0.0011065 | 0.0062266 | 0.055797 | 0.0011065 |  | 2 | 3.9005 | 7.6071 | 4 | 0.0011065 | 0.0062266
2020-07-31 | 1.8997 | 1.8851 | [8 empty] | 1.8924 | 0.010381 | 2 |  | -0.036253 |  | 2 | 3.7848 | 7.1624 | 4 | 0.00010777 | 8.5663e-06 | 0.0020696 | 0.00010777 |  | 2 | 3.7848 | 7.1624 | 4 | 0.00010777 | 8.5663e-06
2020-09-06 | 1.828 | 1.8535 | [8 empty] | 1.8407 | 0.018068 | 2 |  | -0.94115 |  | 2 | 3.6815 | 6.7767 | 4 | 0.00032646 | 0.0057733 | 0.053728 | 0.00032646 |  | 2 | 3.6815 | 6.7767 | 4 | 0.00032646 | 0.0057733
1.8945
Statistics (raw) |  |  | Statistics (corrected) | [12 empty] | 6 | 11.3669 | 21.5463 | 12 | 0.0015407 | 0.012008 |  | 0.0015407 |  | 6 | 11.3669 | 21.5463 | 12 | 0.0015407 | 0.012008
m | 1.8945 |  | m | 1.8945
s(r) | 0.022662 | 1.196% | s(r) | 0.022662 | 1.196% | [10 empty] | K | 2 | [7 empty] | K | 2
s(A) | 0.052396 | 2.766% | s(A) | 0.052396 | 2.766% | [10 empty] | s2(d) | 0.0060042 | [7 empty] | s2(d) | 0.0060042
s(R) | 0.057087 | 3.013% | s(R) | 0.057087 | 3.013%
p | 3 |  | p | 3
Median | 1.8924
MAD | 0.053728 | 2.839%
AAD | 0.037198 | 1.966%
4 ng
[16 empty] | ISO 5725-2, ongecorrigeerd | [9 empty] | ISO 5725-2, gecorrigeerd voor vlaggen
date | value #1 | value #2 | value #3 | value #4 | value #5 | value#6 | value#7 | value#8 | value#9 | value#10 | mean | s | n | flag | Z |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2 | AD | s^2 |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2
2020-05-29 | 3.6454 | 3.7809 | [8 empty] | 3.7132 | 0.095831 | 2 |  | 0.65738 |  | 2 | 7.4264 | 27.5755 | 4 | 0.0091835 | 0.01299 | 0.08059 | 0.0091835 |  | 2 | 7.4264 | 27.5755 | 4 | 0.0091835 | 0.01299
2020-07-31 | 3.6685 | 3.696 | [8 empty] | 3.6822 | 0.019405 | 2 |  | 0.40502 |  | 2 | 7.3645 | 27.1179 | 4 | 0.00037656 | 0.0049308 | 0.049653 | 0.00037656 |  | 2 | 7.3645 | 27.1179 | 4 | 0.00037656 | 0.0049308
2020-09-06 | 3.4647 | 3.5401 | [8 empty] | 3.5024 | 0.053318 | 2 |  | -1.0624 |  | 2 | 7.0047 | 24.533 | 4 | 0.0028429 | 0.033927 | 0.13024 | 0.0028429 |  | 2 | 7.0047 | 24.533 | 4 | 0.0028429 | 0.033927
3.6326
Statistics (raw) |  |  | Statistics (corrected) | [12 empty] | 6 | 21.7956 | 79.2264 | 12 | 0.012403 | 0.051847 |  | 0.012403 |  | 6 | 21.7956 | 79.2264 | 12 | 0.012403 | 0.051847
m | 3.6326 |  | m | 3.6326
s(r) | 0.064299 | 1.77% | s(r) | 0.064299 | 1.77% | [10 empty] | K | 2 | [7 empty] | K | 2
s(A) | 0.10438 | 2.873% | s(A) | 0.10438 | 2.873% | [10 empty] | s2(d) | 0.025923 | [7 empty] | s2(d) | 0.025923
s(R) | 0.12259 | 3.375% | s(R) | 0.12259 | 3.375%
p | 3 |  | p | 3
Median | 3.6822
MAD | 0.08059 | 2.189%
AAD | 0.086829 | 2.358%
6 ng
[16 empty] | ISO 5725-2, ongecorrigeerd | [9 empty] | ISO 5725-2, gecorrigeerd voor vlaggen
date | value #1 | value #2 | value #3 | value #4 | value #5 | value#6 | value#7 | value#8 | value#9 | value#10 | mean | s | n | flag | Z |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2 | AD | s^2 |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2
2020-05-29 | 5.523 | 5.7827 | [8 empty] | 5.6528 | 0.18363 | 2 |  | 0.071751 |  | 2 | 11.3057 | 63.9091 | 4 | 0.033719 | 0.001444 | 0.02687 | 0.033719 |  | 2 | 11.3057 | 63.9091 | 4 | 0.033719 | 0.001444
2020-07-31 | 5.9433 | 6.009 | [8 empty] | 5.9762 | 0.046434 | 2 |  | 0.93519 |  | 2 | 11.9524 | 71.4295 | 4 | 0.0021561 | 0.2453 | 0.35022 | 0.0021561 |  | 2 | 11.9524 | 71.4295 | 4 | 0.0021561 | 0.2453
2020-09-06 | 5.3155 | 5.1823 | [8 empty] | 5.2489 | 0.094173 | 2 |  | -1.0069 |  | 2 | 10.4978 | 55.1014 | 4 | 0.0088685 | 0.28439 | 0.37709 | 0.0088685 |  | 2 | 10.4978 | 55.1014 | 4 | 0.0088685 | 0.28439
5.626
Statistics (raw) |  |  | Statistics (corrected) | [12 empty] | 6 | 33.7558 | 190.44 | 12 | 0.044743 | 0.53114 |  | 0.044743 |  | 6 | 33.7558 | 190.44 | 12 | 0.044743 | 0.53114
m | 5.626 |  | m | 5.626
s(r) | 0.12212 | 2.171% | s(r) | 0.12212 | 2.171% | [10 empty] | K | 2 | [7 empty] | K | 2
s(A) | 0.35402 | 6.293% | s(A) | 0.35402 | 6.293% | [10 empty] | s2(d) | 0.26557 | [7 empty] | s2(d) | 0.26557
s(R) | 0.37449 | 6.656% | s(R) | 0.37449 | 6.656%
p | 3 |  | p | 3
Median | 5.6528
MAD | 0.35022 | 6.195%
AAD | 0.25139 | 4.447%
8 ng
[16 empty] | ISO 5725-2, ongecorrigeerd | [9 empty] | ISO 5725-2, gecorrigeerd voor vlaggen
date | value #1 | value #2 | value #3 | value #4 | value #5 | value#6 | value#7 | value#8 | value#9 | value#10 | mean | s | n | flag | Z |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2 | AD | s^2 |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2
2020-05-29 | 7.3839 | 7.3173 | [8 empty] | 7.3506 | 0.047143 | 2 |  | 0.9035 |  | 2 | 14.7012 | 108.0625 | 4 | 0.0022224 | 0.014747 | 0.08587 | 0.0022224 |  | 2 | 14.7012 | 108.0625 | 4 | 0.0022224 | 0.014747
2020-07-31 | 7.2041 | 7.1429 | [8 empty] | 7.1735 | 0.043298 | 2 |  | -0.95991 |  | 2 | 14.347 | 102.918 | 4 | 0.0018747 | 0.016646 | 0.091232 | 0.0018747 |  | 2 | 14.347 | 102.918 | 4 | 0.0018747 | 0.016646
2020-09-06 | 7.3083 | 7.2319 | [8 empty] | 7.2701 | 0.054073 | 2 |  | 0.056412 |  | 2 | 14.5402 | 105.7083 | 4 | 0.0029239 | 5.7491e-05 | 0.0053615 | 0.0029239 |  | 2 | 14.5402 | 105.7083 | 4 | 0.0029239 | 5.7491e-05
7.2647
Statistics (raw) |  |  | Statistics (corrected) | [12 empty] | 6 | 43.5883 | 316.6888 | 12 | 0.0070211 | 0.031451 |  | 0.0070211 |  | 6 | 43.5883 | 316.6888 | 12 | 0.0070211 | 0.031451
m | 7.2647 |  | m | 7.2647
s(r) | 0.048377 | 0.6659% | s(r) | 0.048377 | 0.6659% | [10 empty] | K | 2 | [7 empty] | K | 2
s(A) | 0.081809 | 1.126% | s(A) | 0.081809 | 1.126% | [10 empty] | s2(d) | 0.015726 | [7 empty] | s2(d) | 0.015726
s(R) | 0.095042 | 1.308% | s(R) | 0.095042 | 1.308%
p | 3 |  | p | 3
Median | 7.2701
MAD | 0.08587 | 1.181%
AAD | 0.060821 | 0.8366%
10 ng
[16 empty] | ISO 5725-2, ongecorrigeerd | [9 empty] | ISO 5725-2, gecorrigeerd voor vlaggen
date | value #1 | value #2 | value #3 | value #4 | value #5 | value#6 | value#7 | value#8 | value#9 | value#10 | mean | s | n | flag | Z |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2 | AD | s^2 |  | n | ny | nyy | n2 | (n-1)ss | n(y-m)^2
2020-05-29 | 8.8349 | 8.9158 | [8 empty] | 8.8754 | 0.05722 | 2 |  | -0.93414 |  | 2 | 17.7508 | 157.5449 | 4 | 0.0032742 | 0.036789 | 0.13563 | 0.0032742 |  | 2 | 17.7508 | 157.5449 | 4 | 0.0032742 | 0.036789
2020-07-31 | 9.2176 | 8.9894 | [8 empty] | 9.1035 | 0.16135 | 2 |  | 0.63696 |  | 2 | 18.207 | 165.747 | 4 | 0.026032 | 0.017105 | 0.092479 | 0.026032 |  | 2 | 18.207 | 165.747 | 4 | 0.026032 | 0.017105
2020-09-06 | 9.1275 | 8.9809 | [8 empty] | 9.0542 | 0.10366 | 2 |  | 0.29718 |  | 2 | 18.1083 | 163.9555 | 4 | 0.010746 | 0.0037232 | 0.043146 | 0.010746 |  | 2 | 18.1083 | 163.9555 | 4 | 0.010746 | 0.0037232
9.011
Statistics (raw) |  |  | Statistics (corrected) | [12 empty] | 6 | 54.0661 | 487.2474 | 12 | 0.040052 | 0.057617 |  | 0.040052 |  | 6 | 54.0661 | 487.2474 | 12 | 0.040052 | 0.057617
m | 9.011 |  | m | 9.011
s(r) | 0.11555 | 1.282% | s(r) | 0.11555 | 1.282% | [10 empty] | K | 2 | [7 empty] | K | 2
s(A) | 0.087914 | 0.9756% | s(A) | 0.087914 | 0.9756% | [10 empty] | s2(d) | 0.028808 | [7 empty] | s2(d) | 0.028808
s(R) | 0.14519 | 1.611% | s(R) | 0.14519 | 1.611%
p | 3 |  | p | 3
Median | 9.0542
MAD | 0.092479 | 1.021%
AAD | 0.090417 | 0.9986%