## Summary
Entry | Generator/diffusion (ng/min) | Sample flow (ml/min) | System flow (l/min) | Loading time (min) | Average loading (ng Hg/tube) | uncertainty (ng Hg/tube) | uncertainty (%) (k=2) | Loading day/number | Tube type | Tube type/tube nr. | Average recovery (ng Hg/tube) | uncertainty (ng Hg/tube) | repeatability (rsd), k=1 | Average recovery (%) | u(r,c) NIST 3133, k=1 | expanded combined uncertainty U(r,c) k=2 | expanded combined uncertainty U(r,c) k=2
1 | G2/4.21 | 100 | 3 | 75 | 10.5 | 0.336 | 3.2% | 16-11 nr. 2 | Gold | 1,2,5,6 | 10.7309 | 0.10536 | 0.9818% | 102.1988 | 0.6156% | 0.42399 | 3.951%
2 |  |  |  |  | 10.5 | 0.336 | 3.2% | 16-11 nr. 3 | Gold | 1,2,3,4 | 10.8415 | 0.063216 | 0.5831% | 103.2524 | 0.6156% | 0.39263 | 3.622%
3 | G2/4.21 | 200 | 3 | 37.5 | 10.5 | 0.336 | 3.2% | 14-11 nr. 8 | Gold + PSA tubes | 1,2 | 13.4966 | 0.44251 | 3.279% | 128.5387 | 0.6156% | 0.9987 | 7.4%
4 |  |  |  |  | 10.5 | 0.336 | 3.2% | 16-11 nr. 1 | Gold | 1,2,3,4 | 10.6202 | 0.05268 | 0.496% | 101.1452 | 0.6156% | 0.37907 | 3.569%
5 | G2/4.21 | 500 | 3 | 75 | 52 | 1.664 | 3.2% | 14-11 nr. 6 | Gold | 1,2,3 | 56.9904 | 0.42153 | 0.7396% | 109.597 | 0.6055% | 2.1243 | 3.728%
6 | G2/4.21 | 500 | 3 | 150 | 105 | 3.36 | 3.2% | 14-11 nr. 7 | Gold | 4,5,6 | 106.9116 | 0.52702 | 0.493% | 101.8205 | 0.6055% | 3.8068 | 3.561% | 103.6028 | 3.4366
7 | G1/218.5 | 50 | 3 | 14 | 50 | 0.7 | 1.4% | 14-11 nr. 5 | Gold | 4, 5, 6 | 50.5832 | 0.42153 | 0.8333% | 101.1664 | 0.6055% | 1.2599 | 2.491%
8 | G1/218.5 | 100 | 3 | 15 | 107 | 1.498 | 1.4% | 14-11 nr. 3 | Gold (pink)* | 16 + 15, 8 + 7, 10 + 9 | 103.76 | 0.74837 | 0.7213% | 96.9719 | 0.6055% | 2.435 | 2.347% | * These tubes have been used for the comparisons in the past
- | G1/218.5 | 100 | 3 | 75 | 535 | 7.49 | 1.4% | [8 empty] | 0
9 | G1/218.5 | 500 | 3 | 15 | 535 | 7.49 | 1.4% | 13-11 nr. 3 | JSI Carbon | 5 + 9, 3 + 4, 12 +16, 13 + 8 | 491.4472 | 9.9356 | 2.022% | 91.8593 | 0.6055% | 21.8544 | 4.447%
10 |  |  |  |  | 535 | 7.49 | 1.4% | 14-11 nr. 2 | LUMEX | OL502086, OL502097, OL502125 | 451.9054 | 10.1995 | 2.257% | 84.4683 | 0.6055% | 22.0475 | 4.879%
11 |  |  |  |  | 535 | 7.49 | 1.4% | 14-11 nr. 4 | Gold | 1,2,3 | 530.9889 | 2.7135 | 0.511% | 99.2503 | 0.6055% | 11.2275 | 2.114%
12 |  |  |  |  | 537 | 7.518 | 1.4% | 15-11 nr. 1 | Gold | 1,2,3,4 | 538.6438 | 4.2551 | 0.79% | 100.3061 | 0.6055% | 13.1085 | 2.434%
13 |  |  |  |  | 537 | 7.518 | 1.4% | 15-11 nr. 3 | Gold | 1,2,3,4 | 538.6438 | 7.0848 | 1.315% | 100.3061 | 0.6055% | 17.3259 | 3.217%
14 |  |  |  |  | 537 | 7.518 | 1.4% | 15-11 nr. 4 | Gold | 1,2,5,6 | 532.9739 | 3.2415 | 0.6082% | 99.2503 | 0.6055% | 11.8049 | 2.215%
15 |  |  |  |  | 537 | 7.518 | 1.4% | 15-11 nr. 5 | Gold | 1,2,3,4 | 532.9739 | 3.3365 | 0.626% | 99.2503 | 0.6055% | 11.9103 | 2.235%
16 | G1/218.5 | 500 | 3 | 30 | 1070 | 14.98 | 1.4% | 13-11 nr. 2 | JSI Carbon | 14 + 21, 6 + 11, 17 + 1, 19 +2 | 1132.1749 | 21.4373 | 1.893% | 105.8107 | 0.5184% | 47.1939 | 4.168%
17 |  |  |  |  | 1069 | 14.966 | 1.4% | 14-11 nr. 1 | LUMEX | OL502123, OL502067, OL502040 | 961.4492 | 45.1494 | 4.696% | 89.9391 | 0.5184% | 91.8392 | 9.552%
18 |  |  |  |  | 1073 | 15.022 | 1.4% | 15-11 nr. 2 | JSI Carbon* | 25 + 26, 27 + 28, 30 +29 | [7 empty] | * Tubes filed with a different type of carbon, results do not meet expectations
19 | G1/218.5 | 100 | 3 | 150 | 1075 | 15.05 | 1.4% | 13-11 nr.1 | JSI Carbon | 4 x dubbel tubes | 1057.8428 | 15.0251 | 1.42% | 98.404 | 0.5184% | 35.2513 | 3.332% | 100.0796 | 2.4662
[18 empty] | 101.3379 | 3.232

## NIST 3133 uncertainty
NIST 3133 | 9.954 | 0.053 | k=2.18 | 0.024312 | k=1 | u(r) | 0.2442%
uncertainty due to repeatability of V pip |  |  |  | uncertainty of pipette pip due to temperature
V (pip) |  |  |  | V (mL) | ±ΔT (°C) | u | u(r)
V (mL) | Average | u | u(r) | 1 | 4 | 0.00048497 | 0.0485%
1.0002 | 1.0048 | 0.0032457 | 0.323%
0.9994
1.0082
1.0072
1.005
1.0068
1.0078 |  |  |  | combined uncertainty Vp
1.003 |  |  |  | V (mL) |  | u(c) | u(c,r)
1.0054 |  |  |  | 1 |  | 0.0032665 | 0.3266%
uncertainty of V vf |  |  |  | uncertainty of vf due to temperature |  |  |  |  | 1000 | 100 μg/mL | u(r,c) | 0.5184%
V (mL) | ± | u | u(r) | V (mL) | ±ΔT (°C) | u | u(r)
100 | 0.1 | 0.040825 | 0.04082% | 100 | 4 | 0.048497 | 0.0485%
combined uncertainty of vf
|  |  |  | V (mL) |  | u(c) | u(c,r)
|  |  |  | 100 |  | 0.063393 | 0.06339%
uncertainty of V vf |  |  |  | uncertainty of vf due to temperature |  |  |  | used to prepare all std
V (mL) |  |  |  | V (mL) | ±ΔT (°C) | u | u(r)
V (mL) | Average | u | u(r) | 30 | 4 | 0.014549 | 0.0485%
30.0183 | 30.0146 | 0.009955 | 0.03317%
30.0192
30.0143
30.019
30.0186
30.0235
30.0227 |  |  |  | combined uncertainty Vvf
29.9969 |  |  |  | V (mL) |  | u(c) | u(c,r)
29.9984 |  |  |  | 30 |  | 0.017626 | 0.05875%
uncertainty due to repeatability of V pip |  |  |  | uncertainty of pipette pip due to temperature |  |  |  | only used to prepare std:
V (p4) |  |  |  | V (mL) | ±ΔT (°C) | u | u(r)
V (mL) | Average | u | u(r) | 5 | 4 | 0.0024249 | 0.0485%
4.982 | 5.0108 | 0.015199 | 0.3033% | [5 empty] | 500 | 50 μg/mL | u(r,c) | 0.6055%
5.0145 | [8 empty] | 100 | 10 μg/mL | u(r,c) | 0.6055%
5.011 | [8 empty] | 50 | 5 μg/mL | u(r,c) | 0.6055%
5.0325
5.022
4.9985
5.01 |  |  |  | combined uncertainty Vp
5.0025 |  |  |  | V (mL) |  | u(c) | u(c,r)
5.0245 |  |  |  | 5 |  | 0.015358 | 0.3072%
uncertainty due to repeatability of V p |  |  |  | uncertainty of pipette p due to temperature |  |  |  | only used to prepare std:
V (p) |  |  |  | V (mL) | ±ΔT (°C) | u | u(r)
V (mL) | Average | u | u(r) | 0.3 | 4 | 0.00014549 | 0.0485%
1.0002 | 1.0048 | 0.0032457 | 0.323% | [5 empty] | 10 | 1 μg/mL | u(r,c) | 0.6156%
0.9994
1.0082
1.0072
1.005
1.0068
1.0078 |  |  |  | combined uncertainty Vp
1.003 |  |  |  | V (mL) |  | u(c) | u(c,r)
1.0054 |  |  |  | 0.3 |  | 0.00097994 | 0.3266%
uncertainty due to repeatability of V p |  |  |  | uncertainty of pipette p due to temperature |  |  |  | used for spiking of all std
V (p) |  |  |  | V (uL) | ±ΔT (°C) | u | u(r)
V (uL) | Average | u | u(r) | 10 | 4 | 0.0048497 | 0.0485%
99.7 | 100.3122 | 0.31088 | 0.3099%
100.19
100.37
100.52
100.61
100.68
100.2 |  |  |  | combined uncertainty Vp
100.05 |  |  |  | V (uL) |  | u(c) | u(c,r)
100.49 |  |  |  | 10 |  | 0.031368 | 0.3137%

## NIST 3133 exact
used conc (μg/mL) | used amount (ng) | exact conc (μg/mL) | exact amount (ng) | factor
100 | 1000 | 105.8107 | 1058.1074 | 1.0581
50 | 500 | 52.7927 | 527.9269 | 1.0559
10 | 100 | 10.5404 | 105.4043 | 1.054
5 | 50 | 5.2691 | 52.6909 | 1.0538
1 | 10 | 1.0536 | 10.536 | 1.0536