## 13-11
| Datum : | 2018-11-13
| Klant : | JSI
| concentratie generator: |  | 218.46 | ng/min
| belading 1 | mfc |  |  | belading 2 | mfc
| flow meten: | 100.62 | ml/min |  | flow meten: | 499.95 | ml/min
|  | 100.59 | ml/min |  |  | 499.76 | ml/min
|  | 100.55 | ml/min |  |  | 499.95 | ml/min
|  | 100.55 | ml/min |  |  | 499.78 | ml/min
|  | 100.59 | ml/min |  |  | 499.78 | ml/min
| gemiddelde: | 100.58 |  |  | gemiddelde: | 499.844
| MFM | 37.25 | % |  | MFM | 90.75 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 3.0043 | l/min | 3.0043 | l/min | 3.0043 | l/min |  | l/min |  | l/min |  | l/min
| 3.0045 | l/min | 3.0045 | l/min | 3.0045 | l/min |  | l/min |  | l/min |  | l/min
| 3.0057 | l/min | 3.0057 | l/min | 3.0057 | l/min |  | l/min |  | l/min |  | l/min
| 3.0043 | l/min | 3.0043 | l/min | 3.0043 | l/min |  | l/min |  | l/min |  | l/min
| 3.0052 | l/min | 3.0052 | l/min | 3.0052 | l/min |  | l/min |  | l/min |  | l/min
flow | 3.0048 |  | 3.0048 |  | 3.0048 |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| belading 1 | [5 empty] | belading 2 en 3
| positie 1 | 37.2 | % | 100.445 | ml/min |  | positie 1 | 90.8 | % | 500.1194 | ml/min |  | 100.445
| positie 2 | 37.2 | % | 100.445 | ml/min |  | positie 2 | 90.8 | % | 500.1194 | ml/min
| positie 3 | 37.15 | % | 100.31 | ml/min |  | positie 3 | 90.8 | % | 500.1194 | ml/min
| positie 4 | 37.2 | % | 100.445 | ml/min |  | positie 4 | 90.8 | % | 500.1194 | ml/min
| positie 5 | 37.3 | % | 100.715 | ml/min |  | positie 5 | 90.8 | % | 500.1194 | ml/min
| positie 6 | 37.2 | % | 100.445 | ml/min |  | positie 6 | 0 | % | 0 | ml/min
| Belading 1: | 150 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: | 30 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 1095.408 | ng/buis | 1 | 15 en 23 | 91.4 | positie 1 | 1090.8155 | ng/buis | 1 | 14 en 21 | 103 |  | No breaktrough
| positie 2 | 1095.408 | ng/buis | 2 | 10 en 24 | 59.1 | positie 2 | 1090.8155 | ng/buis | 2 | 6 en 11 | 98.8
| positie 3 | 1093.9357 | ng/buis | 3 | 18 en 20 | 94.1 | positie 3 | 1090.8155 | ng/buis | 3 | 17 en 1 | 102.3
| positie 4 | 1095.408 | ng/buis | 4 | 7 en 22 | 92.5 | positie 4 | 1090.8155 | ng/buis | 4 | 19 en 2 | 100.6
| positie 5 |  | ng/buis | 5 | JSI carbon |  | positie 5 |  | ng/buis | 5 | JSI carbon
| positie 6 |  | ng/buis | 6 | trap |  | positie 6 |  | ng/buis | 6 | trap
| Belading 3: | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: |  | min. | positie op beladingsrek | buis
| positie 1 | 545.4078 | ng/buis | 1 | 5 en 9 | 89 | positie 1 | #DIV/0! | ng/buis | 1
| positie 2 | 545.4078 | ng/buis | 2 | 3 en 4 | 87.2 | positie 2 | #DIV/0! | ng/buis | 2
| positie 3 | 545.4078 | ng/buis | 3 | 12 en 16 | 87.8 | positie 3 | #DIV/0! | ng/buis | 3
| positie 4 | 545.4078 | ng/buis | 4 | 13 en 8 | 84.9 | positie 4 | #DIV/0! | ng/buis | 4
| positie 5 |  | ng/buis | 5 | JSI carbon |  | positie 5 | #DIV/0! | ng/buis | 5
| positie 6 |  | ng/buis | 6 | trap |  | positie 6 | #DIV/0! | ng/buis | 6
| Belading 5: |  | min. | positie op beladingsrek | buis |  | Belading 6: |  | min. | positie op beladingsrek | buis
| positie 1 | #DIV/0! | ng/buis | 1 |  |  | positie 1 | #DIV/0! | ng/buis | 1
| positie 2 | #DIV/0! | ng/buis | 2 |  |  | positie 2 | #DIV/0! | ng/buis | 2
| positie 3 | #DIV/0! | ng/buis | 3 |  |  | positie 3 | #DIV/0! | ng/buis | 3
| positie 4 | #DIV/0! | ng/buis | 4 |  |  | positie 4 | #DIV/0! | ng/buis | 4
| positie 5 | #DIV/0! | ng/buis | 5 |  |  | positie 5 | #DIV/0! | ng/buis | 5
| positie 6 | #DIV/0! | ng/buis | 6 |  |  | positie 6 | #DIV/0! | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | 15 en 23 | 10 en 24 | 18 en 20 | 7 en 22 | JSI carbon | trap
meting ( ng)
|  |  | buis dummy | 15 en 23 | 10 en 24 | 18 en 20 | 7 en 22 | JSI carbon | trap
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | #REF! | #REF! | #REF! | #REF! | JSI carbon | trap
meting ( ng)
|  |  | buis dummy | #REF! | #REF! | #REF! | #REF! | JSI carbon | trap
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | 5 en 9 | 3 en 4 | 12 en 16 | 13 en 8 | JSI carbon | trap
meting ( ng)
|  |  | buis dummy | 5 en 9 | 3 en 4 | 12 en 16 | 13 en 8 | JSI carbon | trap
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen

## 14-11 (G2 = 4.21)
| Datum : | 2018-11-14
| Klant : | JSI
| concentratie generator 1: |  | 218.5 | ng/min |  |  | concentratie generator 2: |  | 4.21 | ng/min
| belading 1&2 | mfc |  |  | belading 3 | mfc |  |  | belading 4 | mfc |  |  | belading 5 | mfc |  |  | belading 6 & 7 | mfc |  |  | belading 8 | mfc
| flow meten: | 499.62 | ml/min |  | flow meten: | 100.2 | ml/min |  | flow meten: | 499.73 | ml/min |  | flow meten: | 49.97 | ml/min |  | flow meten: | 499.7 | ml/min |  | flow meten: | 200.09 | ml/min
|  | 499.68 | ml/min |  |  | 100.17 | ml/min |  |  | 499.7 | ml/min |  |  | 50.04 | ml/min |  |  | 499.62 | ml/min |  |  | 200.06 | ml/min
|  | 499.76 | ml/min |  |  | 100.13 | ml/min |  |  | 499.76 | ml/min |  |  | 49.846 | ml/min |  |  | 499.7 | ml/min |  |  | 199.97 | ml/min
|  | 499.87 | ml/min |  |  | 100.16 | ml/min |  |  | 499.68 | ml/min |  |  | 50.1 | ml/min |  |  | 499.79 | ml/min |  |  | 200.04 | ml/min
|  | 499.85 | ml/min |  |  | 100.24 | ml/min |  |  | 499.68 | ml/min |  |  | 49.926 | ml/min |  |  | 499.7 | ml/min |  |  | 200.01 | ml/min
| gemiddelde: | 499.756 |  |  | gemiddelde: | 100.18 |  |  | gemiddelde: | 499.71 |  |  | gemiddelde: | 49.9764 |  |  | gemiddelde: | 499.702 |  |  | gemiddelde: | 200.034
| MFM | 90.75 | % |  | MFM | 37.15 | % |  | MFM | 90.8 | % |  | MFM | 91.75 | % |  | MFM | 90.4 | % |  | MFM | 93.2 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6 |  | belading 7 |  | belading 8
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 3.006 | l/min | 3.006 | l/min | 3.006 | l/min | 3.006 | l/min | 3.006 | l/min | 3.0165 | l/min | 3.0165 | l/min | 3.0165 | l/min
| 3.0056 | l/min | 3.0056 | l/min | 3.0056 | l/min | 3.0056 | l/min | 3.0056 | l/min | 3.0148 | l/min | 3.0148 | l/min | 3.0148 | l/min
| 3.0005 | l/min | 3.0005 | l/min | 3.0005 | l/min | 3.0005 | l/min | 3.0005 | l/min | 3.0164 | l/min | 3.0164 | l/min | 3.0164 | l/min
| 3.0014 | l/min | 3.0014 | l/min | 3.0014 | l/min | 3.0014 | l/min | 3.0014 | l/min | 3.0153 | l/min | 3.0153 | l/min | 3.0153 | l/min
| 3.0055 | l/min | 3.0055 | l/min | 3.0055 | l/min | 3.0055 | l/min | 3.0055 | l/min | 3.0157 | l/min | 3.0157 | l/min | 3.0157 | l/min
flow | 3.0038 |  | 3.0038 |  | 3.0038 |  | 3.0038 |  | 3.0038 |  | 3.0157 |  | 3.0157 |  | 3.0157
| belading 1 & 2 | [5 empty] | belading 3 | [5 empty] | belading 4 | [5 empty] | belading 5
| positie 1 | 90.7 | % | 499.4807 | ml/min |  | positie 1 | 37.15 | % | 100.18 | ml/min |  | positie 1 | 90.8 | % | 499.71 | ml/min |  | positie 1 | 91.45 | % | 49.813 | ml/min
| positie 2 | 90.8 | % | 500.0313 | ml/min |  | positie 2 | 37.2 | % | 100.3148 | ml/min |  | positie 2 | 90.8 | % | 499.71 | ml/min |  | positie 2 | 91.55 | % | 49.8675 | ml/min
| positie 3 | 90.7 | % | 499.4807 | ml/min |  | positie 3 | 37.2 | % | 100.3148 | ml/min |  | positie 3 | 90.8 | % | 499.71 | ml/min |  | positie 3 | 91.6 | % | 49.8947 | ml/min
| positie 4 | 0 | % | 0 | ml/min |  | positie 4 | 0 | % | 0 | ml/min |  | positie 4 | 0 | % | #DIV/0! | ml/min |  | positie 4 | 0 | % | 0 | ml/min
| positie 5 | 0 | % | 0 | ml/min |  | positie 5 | 0 | % | 0 | ml/min |  | positie 5 | 0 | % | #DIV/0! | ml/min |  | positie 5 | 0 | % | 0 | ml/min
| positie 6 | 0 | % | 0 | ml/min |  | positie 6 | 0 | % | 0 | ml/min |  | positie 6 | 0 | % | #DIV/0! | ml/min |  | positie 6 | 0 | % | 0 | ml/min
| Belading 1: G1 | 30 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: G1 | 15 | min. | positie op beladingsrek | buis |  | belading 6 & 7 | [5 empty] | belading 8
| positie 1 | 1089.9846 | ng/buis | 1 | OL502123 | 89.5 | positie 1 | 544.9923 | ng/buis | 1 | OL502086 | 82.8 | positie 1 | 90.4 | % | 499.702 | ml/min |  | positie 1 | 93.25 | % | 200.1413 | ml/min
| positie 2 | 1091.1863 | ng/buis | 2 | OL502067 | 82.3 | positie 2 | 545.5932 | ng/buis | 2 | OL502097 | 79.2 | positie 2 | 90.4 | % | 499.702 | ml/min |  | positie 2 | 93.25 | % | 200.1413 | ml/min
| positie 3 | 1089.9846 | ng/buis | 3 | OL502040 | 83.1 | positie 3 | 544.9923 | ng/buis | 3 | OL502125 | 81.4 | positie 3 | 90.4 | % | 499.702 | ml/min |  | positie 3 | 93.25 | % | 200.1413 | ml/min
| positie 4 |  | ng/buis | 4 |  | 84.9667 | positie 4 |  | ng/buis | 4 |  |  | positie 4 | 0 | % | #DIV/0! | ml/min |  | positie 4 | 93.25 | % | 200.1413 | ml/min
| positie 5 |  | ng/buis | 5 | LUMEX | 3.9463 | positie 5 |  | ng/buis | 5 | LUMEX |  | positie 5 | 0 | % | #DIV/0! | ml/min |  | positie 5 | 93.25 | % | 200.1413 | ml/min
| positie 6 |  | ng/buis | 6 | tubes | 4.645% | positie 6 |  | ng/buis | 6 | tubes |  | positie 6 | 0 | % | #DIV/0! | ml/min |  | positie 6 | 93.25 | % | 200.1413 | ml/min
| Belading 3: G1 | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: G1 | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 109.3082 | ng/buis | 1 | 16 en 15 | 92.31% | positie 1 | 545.2425 | ng/buis | 1 | 1 | 94.4 |  | 12.4682
| positie 2 | 109.4553 | ng/buis | 2 | 8 en 7 |  | positie 2 | 545.2425 | ng/buis | 2 | 2 | 94.2 |  | 13.0968
| positie 3 | 109.4553 | ng/buis | 3 | 10 en 9 | 91.27% | positie 3 | 545.2425 | ng/buis | 3 | 3 | 93.5
| positie 4 |  | ng/buis | 4 |  |  | positie 4 |  | ng/buis | 4 |  | 94.0333
| positie 5 |  | ng/buis | 5 | Gold (pink) |  | positie 5 |  | ng/buis | 5 | Gold (gold) | 0.47258 |  | 12.7825 | 1.22 | 22%
| positie 6 |  | ng/buis | 6 | tubes |  | positie 6 |  | ng/buis | 6 | tubes | 0.5026%
| Belading 5: G1 | 14 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 6: G2 | 75 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 7: G2 | 150 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 8: G2 | 37.5 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 50.7284 | ng/buis | 1 | 4 | 97.2 | positie 1 | 52.3191 | ng/buis | 1 | 1 | 104.4 | positie 1 | 104.6383 | ng/buis | 1 | 4 | 97.1 | positie 1 | 10.4775 | ng/buis | 1 | 1 | 119
| positie 2 | 50.7839 | ng/buis | 2 | 5 | 95.7 | positie 2 | 52.3191 | ng/buis | 2 | 2 | 105.4 | positie 2 | 104.6383 | ng/buis | 2 | 5 | 96.6 | positie 2 | 10.4775 | ng/buis | 2 | 2 | 125
| positie 3 | 50.8116 | ng/buis | 3 | 6 | 95.7 | positie 3 | 52.3191 | ng/buis | 3 | 3 | 103.7 | positie 3 | 104.6383 | ng/buis | 3 | 6 | 96.1 | positie 3 | 10.4775 | ng/buis | 3 | J+H PSA
| positie 4 |  | ng/buis | 4 |  | 96.2 | positie 4 |  | ng/buis | 4 |  | 104.5 | positie 4 |  | ng/buis | 4 |  |  | positie 4 | 10.4775 | ng/buis | 4 | C+I PSA
| positie 5 |  | ng/buis | 5 | Gold (gold) | 0.86603 | positie 5 |  | ng/buis | 5 | Gold (gold) | 0.8544 | positie 5 |  | ng/buis | 5 | Gold (gold) |  | positie 5 | 10.4775 | ng/buis | 5 | G+F PSA
| positie 6 |  | ng/buis | 6 | tubes | 0.9002% | positie 6 |  | ng/buis | 6 | tubes | 0.8176% | positie 6 |  | ng/buis | 6 | tubes |  | positie 6 | 10.4775 | ng/buis | 6 | E+D PSA
Belading 1
|  | meet volgorde | buis dummy | OL502123 | OL502067 | OL502040 | 0 | LUMEX | tubes
meting ( ng)
|  |  | buis dummy | OL502123 | OL502067 | OL502040 | 0 | LUMEX | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | #REF! | #REF! | #REF! | #REF! | LUMEX | tubes
meting ( ng)
|  |  | buis dummy | #REF! | #REF! | #REF! | #REF! | LUMEX | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | 16 en 15 | 8 en 7 | 10 en 9 | 0 | Gold (pink) | tubes
meting ( ng)
|  |  | buis dummy | 16 en 15 | 8 en 7 | 10 en 9 | 0 | Gold (pink) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | #REF! | 1 | 3 | 0 | Gold (gold) | tubes
meting ( ng)
|  |  | buis dummy | #REF! | 1 | 3 | 0 | Gold (gold) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 4 | 5 | 6 | 0 | Gold (gold) | tubes
meting ( ng)
|  |  | buis dummy | 4 | 5 | 6 | 0 | Gold (gold) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 1 | 2 | 3 | 0 | Gold (gold) | tubes
meting ( ng)
|  |  | buis dummy | 1 | 2 | 3 | 0 | Gold (gold) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen

## 15-11
| Datum : | 2018-11-15
| Klant : | JSI
| concentratie generator 1: |  | 218.5 | ng/min
| belading 1 tm 5 | mfc
| flow meten: | 499.9 | ml/min
|  | 499.92 | ml/min
|  | 499.96 | ml/min
|  | 499.84 | ml/min
|  | 499.98 | ml/min
| gemiddelde: | 499.92
| MFM | 90.8 | %
belading 1 tm 5
systeem
| 2.9956 | l/min
| 2.9932 | l/min
| 2.9952 | l/min
| 2.9928 | l/min
| 2.9964 | l/min
flow | 2.9946
belading 1 tm 5
| positie 1 | 90.85 | % | 500.1953 | ml/min
| positie 2 | 90.8 | % | 499.92 | ml/min
| positie 3 | 90.8 | % | 499.92 | ml/min
| positie 4 | 90.85 | % | 500.1953 | ml/min
| positie 5 | 0 | % | 0 | ml/min
| positie 6 | 0 | % | 0 | ml/min
| Belading 1: | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: | 30 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 547.4414 | ng/buis | 1 | 1 | 95.4 | positie 1 | 1094.8829 | ng/buis | 1 | 25 en 26 | :( | 516.7847
| positie 2 | 547.1402 | ng/buis | 2 | 2 | 95.2 | positie 2 | 1094.2803 | ng/buis | 2 | 27 en 28 | :( | 509.9346
| positie 3 | 547.1402 | ng/buis | 3 | 3 | 93.8 | positie 3 | 1094.2803 | ng/buis | 3 | 30 en 29 | :( | 511.576
| positie 4 | 547.4414 | ng/buis | 4 | 4 | 94.6 | positie 4 |  | ng/buis | 4 |  |  | 515.1424
| positie 5 |  | ng/buis | 5 | Gold | 94.75 | positie 5 |  | ng/buis | 5 | Carbon traps |  | 513.3594 | 0.93774 | -6.226%
| positie 6 |  | ng/buis | 6 | traps | 0.7188 | positie 6 |  | ng/buis | 6 | Igor Self-made
0.7586%
| Belading 3: | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 547.4414 | ng/buis | 1 | 1 | 96 | positie 1 | 547.4414 | ng/buis | 1 | 1 | 94.6
| positie 2 | 547.1402 | ng/buis | 2 | 2 | 95.8 | positie 2 | 547.1402 | ng/buis | 2 | 2 | 93.7
| positie 3 | 547.1402 | ng/buis | 3 | 3 | 95.4 | positie 3 | 547.1402 | ng/buis | 3 | 5 | 94.2
| positie 4 | 547.4414 | ng/buis | 4 | 4 | 93.3 | positie 4 | 547.4414 | ng/buis | 4 | 6 | 93.3
| positie 5 |  | ng/buis | 5 | Gold | 95.125 | positie 5 |  | ng/buis | 5 | Gold | 93.95
| positie 6 |  | ng/buis | 6 | traps | 1.242 | positie 6 |  | ng/buis | 6 | traps | 0.56862
[6 empty] | 1.306% | [5 empty] | 0.6052%
| Belading 5: | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 547.4414 | ng/buis | 1 | 1 | 94.4
| positie 2 | 547.1402 | ng/buis | 2 | 2 | 93.2
| positie 3 | 547.1402 | ng/buis | 3 | 3 | 93.5
| positie 4 | 547.4414 | ng/buis | 4 | 4 | 94.1
| positie 5 |  | ng/buis | 5 | Gold | 93.8
| positie 6 |  | ng/buis | 6 | traps | 0.54772
0.5839%

## 16-11 (G2 = 4.21)
| Datum : | 2018-11-16
| Klant : | JSI
| concentratie generator 2: |  | 4.21 | ng/min
| belading 1 | mfc |  |  | belading 2 | mfc
| flow meten: | 99.87 | ml/min |  | flow meten: | 199.85 | ml/min
|  | 99.95 | ml/min |  |  | 199.98 | ml/min
|  | 99.92 | ml/min |  |  | 199.89 | ml/min
|  | 99.91 | ml/min |  |  | 200.06 | ml/min
|  | 99.86 | ml/min |  |  | 199.9 | ml/min
| gemiddelde: | 99.902 |  |  | gemiddelde: | 199.936
| MFM | 36.95 | % |  | MFM | 73.45 | %
belading 1, 2 & 3
systeem
| 3.0183 | l/min
| 3.0169 | l/min
| 3.0177 | l/min
| 3.0189 | l/min
| 3.0173 | l/min
flow | 3.0178
| belading 1 | [5 empty] | belading 2
| positie 1 | 37.05 | % | 100.1724 | ml/min |  | positie 1 | 73.45 | % | 199.936 | ml/min
| positie 2 | 37.05 | % | 100.1724 | ml/min |  | positie 2 | 73.45 | % | 199.936 | ml/min
| positie 3 | 37 | % | 100.0372 | ml/min |  | positie 3 | 73.5 | % | 200.0721 | ml/min
| positie 4 | 37.05 | % | 100.1724 | ml/min |  | positie 4 | 73.4 | % | 199.7999 | ml/min
| positie 5 | 0 | % | 0 | ml/min |  | positie 5 | 0 | % | 0 | ml/min
| positie 6 | 0 | % | 0 | ml/min |  | positie 6 | 0 | % | 0 | ml/min
| Belading 1: | 75 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: | 37.5 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 10.4809 | ng/buis | 1 | 1 | 95.9 | positie 1 | 10.4595 | ng/buis | 1 | 1 | 98.1 | 10.2399
| positie 2 | 10.4809 | ng/buis | 2 | 2 | 95.3 | positie 2 | 10.4595 | ng/buis | 2 | 2 | 96.4 | 10.3444
| positie 3 | 10.4667 | ng/buis | 3 | 3 | 96.2 | positie 3 | 10.4666 | ng/buis | 3 | 5 | 96.1 | 10.3096
| positie 4 | 10.4809 | ng/buis | 4 | 4 | 96.7 | positie 4 | 10.4524 | ng/buis | 4 | 6 | 97.9 | 10.2015
| positie 5 |  | ng/buis | 5 | Gold | 96.025 | positie 5 |  | ng/buis | 5 | Gold | 97.125 | 10.2739 | 0.98292 | -1.708%
| positie 6 |  | ng/buis | 6 | traps | 0.58523 | positie 6 |  | ng/buis | 6 | traps | 1.021
[6 empty] | 0.6095% | [5 empty] | 1.051%
| Belading 3: | 37.5 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 10.4595 | ng/buis | 1 | 1 | 97.9 | positie 1 | #REF! | ng/buis | 1
| positie 2 | 10.4595 | ng/buis | 2 | 2 | 98.9 | positie 2 | #REF! | ng/buis | 2
| positie 3 | 10.4666 | ng/buis | 3 | 3 | 98.5 | positie 3 | #REF! | ng/buis | 3
| positie 4 | 10.4524 | ng/buis | 4 | 4 | 97.6 | positie 4 | #REF! | ng/buis | 4
| positie 5 |  | ng/buis | 5 |  | 98.225 | positie 5 | #REF! | ng/buis | 5
| positie 6 |  | ng/buis | 6 |  | 0.58523 | positie 6 | #REF! | ng/buis | 6
0.5958%
| Belading 5: |  | min. | positie op beladingsrek | buis |  | Belading 6: |  | min. | positie op beladingsrek | buis
| positie 1 | #REF! | ng/buis | 1 |  |  | positie 1 | #REF! | ng/buis | 1
| positie 2 | #REF! | ng/buis | 2 |  |  | positie 2 | #REF! | ng/buis | 2
| positie 3 | #REF! | ng/buis | 3 |  |  | positie 3 | #REF! | ng/buis | 3
| positie 4 | #REF! | ng/buis | 4 |  |  | positie 4 | #REF! | ng/buis | 4
| positie 5 | #REF! | ng/buis | 5 |  |  | positie 5 | #REF! | ng/buis | 5
| positie 6 | #REF! | ng/buis | 6 |  |  | positie 6 | #REF! | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | 1 | 2 | 3 | 4 | Gold | traps
meting ( ng)
|  |  | buis dummy | 1 | 2 | 3 | 4 | Gold | traps
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | #REF! | #REF! | #REF! | #REF! | Gold | traps
meting ( ng)
|  |  | buis dummy | #REF! | #REF! | #REF! | #REF! | Gold | traps
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | 1 | 2 | 3 | 4 | 0 | 0
meting ( ng)
|  |  | buis dummy | 1 | 2 | 3 | 4 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
opmerkingen

## Summary
Entry | Tube type | Tube numbers | Hg loading | Analysed value (ng) | Average (ng) | Stdev | Difference (%) | Entry | Tube type | Tube numbers | Hg loading | Analysed value (ng) | Average (ng) | Stdev | Difference (%)
|  |  | (ng) |  |  | (ng) |  |  |  |  | (ng) |  |  | (ng)
1 | Gold | 1 | 10.4595 | 10.3 | 10.175 | 0.095743 | -2.72% | 1 | Gold | 1 | 10.5092 | 10.3 | 10.175 | 0.095743 | -3.18%
|  | 2 | 10.4595 | 10.1 | [5 empty] | 2 | 10.5092 | 10.1
|  | 5 | 10.4666 | 10.1 | [5 empty] | 5 | 10.5092 | 10.1 |  |  |  | -2.797% | all
|  | 6 | 10.4524 | 10.2 | [5 empty] | 6 | 10.502 | 10.2 |  |  |  | -3.324% | 10 ng
2 | Gold | 1 | 10.4595 | 10.2 | 10.25 | 0.057735 | -2.003% | 2 | Gold | 1 | 10.5092 | 10.2 | 10.25 | 0.057735 | -2.466% | -1.995% | 50 and 100 ng
|  | 2 | 10.4595 | 10.3 | [5 empty] | 2 | 10.5092 | 10.3 |  |  |  | -3.84% | 500 ng
|  | 3 | 10.4666 | 10.3 | [5 empty] | 3 | 10.5163 | 10.3 |  |  |  | -1.001% | 1000 ng
|  | 4 | 10.4524 | 10.2 | [5 empty] | 4 | 10.502 | 10.2 |  |  |  | -1.951% | 3 mm
3 | Gold + PSA tubes | 1 | 10.4775 | 12.5 | 12.8 | 0.42426 | 18.14% | 3 | Gold + PSA tubes | 1 | 10.5272 | 12.5 | 12.8 | 0.42426 | 17.76% | -3.267% | 33 mm
|  | 2 | 10.4775 | 13.1 | [5 empty] | 2 | 10.5272 | 13.1
4 | Gold | 1 | 10.4809 | 10.1 | 10.075 | 0.05 | -3.873% | 4 | Gold | 1 | 10.5307 | 10.1 | 10.075 | 0.05 | -4.327%
|  | 2 | 10.4809 | 10 | [5 empty] | 2 | 10.5307 | 10
|  | 3 | 10.4667 | 10.1 | [5 empty] | 3 | 10.5165 | 10.1
|  | 4 | 10.4809 | 10.1 | [5 empty] | 4 | 10.5307 | 10.1
5 | Gold | 1 | 52.3191 | 54.7 | 54.7 | 0.4 | 4.551% | 5 | Gold | 1 | 52.5677 | 54.7 | 54.7 | 0.4 | 4.056%
|  | 2 | 52.3191 | 55.1 | [5 empty] | 2 | 52.5677 | 55.1
|  | 3 | 52.3191 | 54.3 | [5 empty] | 3 | 52.5677 | 54.3
6 | Gold | 4 | 104.6383 | 101.6 | 101.1 | 0.5 | -3.381% | 6 | Gold | 4 | 105.1354 | 101.6 | 101.1 | 0.5 | -3.838%
|  | 5 | 104.6383 | 101.1 | [5 empty] | 5 | 105.1354 | 101.1
|  | 6 | 104.6383 | 100.6 | [5 empty] | 6 | 105.1354 | 100.6
7 | Gold | 4 | 50.7284 | 49.3 | 48.8333 | 0.40415 | -3.736% | 7 | Gold | 4 | 49.7533 | 49.3 | 48.8333 | 0.40415 | -1.849%
|  | 5 | 50.7839 | 48.6 | [5 empty] | 5 | 49.8077 | 48.6
|  | 6 | 50.8116 | 48.6 | [5 empty] | 6 | 49.8349 | 48.6
8 | Gold old | 16 + 15 | 109.3082 | 100.9 | 100.4 | 0.70711 | -8.15% | 8 | Gold old | 16 + 15 | 107.2071 | 100.9 | 100.4 | 0.70711 | -6.349%
|  | 8 + 7 | 109.4553 | - | [5 empty] | 8 + 7 | 107.3514 | -
|  | 10 + 9 | 109.4553 | 99.9 | [5 empty] | 10 + 9 | 107.3514 | 99.9
9 | JSI Carbon | 5 + 9 | 545.4078 | 485.5 | 475.825 | 9.4097 | -12.76% | 9 | JSI Carbon | 5 + 9 | 535.0219 | 485.5 | 475.825 | 9.4097 | -11.06%
| Type 1 | 3 + 4 | 545.4078 | 475.7 |  |  |  |  | Type 1 | 3 + 4 | 535.0219 | 475.7
|  | 12 + 16 | 545.4078 | 479 | [5 empty] | 12 + 16 | 535.0219 | 479
|  | 13 + 8 | 545.4078 | 463.1 | [5 empty] | 13 + 8 | 535.0219 | 463.1
10 | LUMEX | 86 | 544.9923 | 451.3 | 442.3333 | 9.6625 | -18.84% | 10 | LUMEX | 86 | 534.5165 | 451.3 | 442.3333 | 9.6625 | -17.25%
|  | 97 | 545.5932 | 432.1 | [5 empty] | 97 | 535.1058 | 432.1
|  | 125 | 544.9923 | 443.6 | [5 empty] | 125 | 534.5165 | 443.6
11 | Gold | 1 | 545.2425 | 514.7 | 512.7 | 2.571 | -5.968% | 11 | Gold | 1 | 534.7619 | 514.7 | 512.7 | 2.571 | -4.126%
|  | 2 | 545.2425 | 513.6 | [5 empty] | 2 | 534.7619 | 513.6
|  | 3 | 545.2425 | 509.8 | [5 empty] | 3 | 534.7619 | 509.8
12 | Gold | 1 | 547.4414 | 522.3 | 518.575 | 4.0261 | -5.273% | 12 | Gold | 1 | 536.9185 | 522.3 | 518.575 | 4.0261 | -3.416%
|  | 2 | 547.1402 | 520.9 | [5 empty] | 2 | 536.623 | 520.9
|  | 3 | 547.1402 | 513.2 | [5 empty] | 3 | 536.623 | 513.2
|  | 4 | 547.4414 | 517.9 | [5 empty] | 4 | 536.9185 | 517.9
13 | Gold | 1 | 547.4414 | 525.5 | 520.625 | 6.7074 | -4.899% | 13 | Gold | 1 | 536.9185 | 525.5 | 520.625 | 6.7074 | -3.035%
|  | 2 | 547.1402 | 524.2 | [5 empty] | 2 | 536.623 | 524.2
|  | 3 | 547.1402 | 522 | [5 empty] | 3 | 536.623 | 522
|  | 4 | 547.4414 | 510.8 | [5 empty] | 4 | 536.9185 | 510.8
14 | Gold | 1 | 547.4414 | 517.8 | 514.175 | 3.0663 | -6.077% | 14 | Gold | 1 | 536.9185 | 517.8 | 514.175 | 3.0663 | -4.236%
|  | 2 | 547.1402 | 512.7 | [5 empty] | 2 | 536.623 | 512.7
|  | 5 | 547.1402 | 515.4 | [5 empty] | 5 | 536.623 | 515.4
|  | 6 | 547.4414 | 510.8 | [5 empty] | 6 | 536.9185 | 510.8
15 | Gold | 1 | 547.4414 | 516.8 | 513.35 | 3.1586 | -6.227% | 15 | Gold | 1 | 536.9185 | 516.8 | 513.35 | 3.1586 | -4.39%
|  | 2 | 547.1402 | 509.9 | [5 empty] | 2 | 536.623 | 509.9
|  | 3 | 547.1402 | 511.6 | [5 empty] | 3 | 536.623 | 511.6
|  | 4 | 547.4414 | 515.1 | [5 empty] | 4 | 536.9185 | 515.1
16 | Gold | 14 + 21 | 1090.8155 | 1128.5 | 1108.125 | 20.263 | 1.587% | 16 | Gold | 14 + 21 | 1070.0438 | 1128.5 | 1108.125 | 20.263 | 3.559%
|  | 6 +11 | 1090.8155 | 1082.5 | [5 empty] | 6 +11 | 1070.0438 | 1082.5
|  | 17 + 1 | 1090.8155 | 1119.3 | [5 empty] | 17 + 1 | 1070.0438 | 1119.3
|  | 19 + 2 | 1090.8155 | 1102.2 | [5 empty] | 19 + 2 | 1070.0438 | 1102.2
17 | LUMEX | 123 | 1089.9846 | 975.5 | 926.4333 | 42.6716 | -15% | 17 | LUMEX | 123 | 1069.0329 | 975.5 | 926.4333 | 42.6716 | -13.34%
|  | 67 | 1091.1863 | 898 | [5 empty] | 67 | 1070.2116 | 898
|  | 40 | 1089.9846 | 905.8 | [5 empty] | 40 | 1069.0329 | 905.8
18 | JSI Carbon type 2 | 25 + 26 | 1094.8829 | Tubes filed with a different type of carbon, results do not meet expectations |  |  |  | 18 | JSI Carbon type 2 | 25 + 26 | 1073.8371 | Tubes filed with a different type of carbon, results do not meet expectations
|  | 27 + 28 | 1094.2803 | [6 empty] | 27 + 28 | 1073.2461
|  | 30 + 29 | 1094.2803 | [6 empty] | 30 + 29 | 1073.2461
19 | JSI Carbon type 1 | 7 + 10 | 1095.408 | 1001.4 | 1014.8 | 14.152 | -7.359% | 19 | JSI Carbon type 1 | 7 + 10 | 1074.5488 | 1001.4 | 1014.8 | 14.152 | -5.56%
|  | 15 + 18 | 1095.408 | [6 empty] | 15 + 18 | 1074.5488
|  | 20 + 22 | 1093.9357 | 1029.6 | [5 empty] | 20 + 22 | 1073.1045 | 1029.6
|  | 23 + 24 | 1095.408 | 1013.4 | [5 empty] | 23 + 24 | 1074.5488 | 1013.4

## Uncertainty 3 mm
|  |  |  | x(Hg) = (Q18 + Q19 +Q20) | [10 empty] | Q(Hg) = Q18 + Q19 +Q20
Uncertainty mercury concentration in the generator | [10 empty] | Uncertainty sum diffusion
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y) |  | X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
Q 18 | ng/min | 1.3739 | 0.030151 | 2.195% | 1 | 0.030151 | 1/d | 0.33156 | 0.0099971 | a | Q 18 | ng/min | 1.3739 | 0.030151 | 2.195% | 1 | 0.030151 | 1 | 1 | 0.030151 | a
Q 19 | ng/min | 1.4415 | 0.044124 | 3.061% | 1 | 0.044124 | 1/d | 0.33156 | 0.01463 | b | Q 19 | ng/min | 1.4415 | 0.044124 | 3.061% | 1 | 0.044124 | 1 | 1 | 0.044124 | b
Q 20 | ng/min | 1.4137 | 0.03596 | 2.544% | 1 | 0.03596 | 1/d | 0.33156 | 0.011923 | c | Q 20 | ng/min | 1.4137 | 0.03596 | 2.544% | 1 | 0.03596 | 1 | 1 | 0.03596 | c
MFC1 | L/min | 0.5 | 0.0015 | 0.3% | 1 | 0.0015 |  |  |  |  | Q | ng/min | 4.229 | 0.12883 | 3.046% | 2 | 0.064414 | 1
MFC2 | L/min | 2.5 | 0.0075 | 0.3% | 1 | 0.0075
Bios | L/min | 3.016 | 0.006032 | 0.2% | 2 | 0.003016
ρ air | L/min | 3.016 | 0.0082217 | 0.2726% | 1 | 0.0082217 | -((a+b+c)/d^2) | -0.46492 | -0.0038224 | d
x(Hg) | ng/m3 | 1402.1848 | 43.3937 | 3.095% | 2 | 21.6968
Uncertainty mercury mass on transfer standards
m(Hg) = (t* ρ sample) * Q/ρ air
Uncertainty mercury mass on transfer standards
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
t | min | 150 | 0.03 | 0.02% | 2 | 0.015 | c(d+e+f)/g | 0.70067 | 0.01051 | b
ρ sample | L/min | 0.4997 | 0.0017489 | 0.5% | 1 | 0.0017489 | b(d+e+f)/g | 210.3277 | 0.36785 | c
MFC | L/min | 0.4997 | 0.0014991 | 0.3% | 1 | 0.0014991
MFM | L/min | 0.4997 | 0.0009994 | 0.2% | 2 | 0.0004997
MPM | L/min | 0.4997 | 0.0014991 | 0.3% | 2 | 0.00074955
Q 18 | ng/min | 1.3739 | 0.030151 | 2.195% | 1 | 0.030151 | bc/g | 24.8525 | 0.74933 | d
Q 19 | ng/min | 1.4415 | 0.044124 | 3.061% | 1 | 0.044124 | bc/g | 24.8525 | 1.0966 | e
Q 20 | ng/min | 1.4137 | 0.03596 | 2.544% | 1 | 0.03596 | bc/g | 24.8525 | 0.8937 | f
ρ air | L/min | 3.016 | 0.0082217 | 0.2726% | 1 | 0.0082217 | -bc(f+e+d)/g^2 | -34.8477 | -0.28651 | g
m(Hg) | ng | 105.1008 | 3.3348 | 3.173% | 2 | 1.6674
m(Hg) = (t* ρ sample) * x(Hg)/1000
Uncertainty mercury mass on transfer standards
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
t | min | 37.5 | 0.03 | 0.08% | 2 | 0.015 | bc/1000 | 0.70067 | 0.01051 | a
ρ sample | ml/min | 0.4997 | 0.0017489 | 0.5% | 1 | 0.0017489 | ac/1000 | 52.5819 | 0.091963 | b
MFC | ml/min | 0.4997 | 0.0014991 | 0.3% | 1 | 0.0014991
MFM | ml/min | 0.4997 | 0.0009994 | 0.2% | 2 | 0.0004997
MPM | ml/min | 0.4997 | 0.0014991 | 0.3% | 2 | 0.00074955
x(Hg) | ng/m3 | 1402.1848 | 43.3937 | 3.095% | 2 | 21.6968 | ab/1000 | 0.018739 | 0.40657 | c
m(Hg) | ng | 26.2752 | 0.83395 | 3.174% | 2 | 0.41697
Note on A19: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the timer used is 0.03 minutes for a time interval of 0.1 to 100 minutes with a coverage factor of k = 2. For a loading of 25 minutes this leads to a standard uncertainty of 0.0012 % with a coverage factor of k = 2, which is negligible compared to other uncertainty contributions.
Note on A21: Krom, Iris (I.M.J.) de, Mrs: Mass-flow controller output variations are stated by the manufacturer to be < 0.1 % temperature sensitivity of full scale per K of the reading. The uncertainty of mass-flow controller output at full scale, at a temperature uncertainty of 1 K, is equal to (0.1*0.3)2 = 0.03 %. The uncertainty in drift of the mass-flow controller is  0.3 %. By keeping the mass-flow controller set point above 20 %, uMFC  0.3 %.
Note on A22: Krom, Iris (I.M.J.) de, Mrs: The reading of the MFM (mass flow meter) is within ± 0.05 (% set point), resulting in a reading uncertainty of 0.029 (% set point). By keeping the set point above 40%, the relative reading uncertainty is kept below 0.07%. Manufacturers’ specifications of the repeatability of the MFM is < 0.2% of reading. The resulting uncertainty contribution UMFM is 0.2% (k = 2).
Note on A23: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the MPM (mercury piston prover) used is 0.3 % for flow rates from 20 up to 1000 mL.min-1 with a coverage factor of k = 2, UMPM = 0.3 %.
Note on A35: Krom, Iris (I.M.J.) de, Mrs: Mass-flow controller output variations are stated by the manufacturer to be < 0.1 % temperature sensitivity of full scale per K of the reading. The uncertainty of mass-flow controller output at full scale, at a temperature uncertainty of 1 K, is equal to (0.1*0.3)2 = 0.03 %. The uncertainty in drift of the mass-flow controller is  0.3 %. By keeping the mass-flow controller set point above 20 %, uMFC  0.3 %.
Note on A36: Krom, Iris (I.M.J.) de, Mrs: The reading of the MFM (mass flow meter) is within ± 0.05 (% set point), resulting in a reading uncertainty of 0.029 (% set point). By keeping the set point above 40%, the relative reading uncertainty is kept below 0.07%. Manufacturers’ specifications of the repeatability of the MFM is < 0.2% of reading. The resulting uncertainty contribution UMFM is 0.2% (k = 2).
Note on A37: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the MPM (mercury piston prover) used is 0.3 % for flow rates from 20 up to 1000 mL.min-1 with a coverage factor of k = 2, UMPM = 0.3 %.

## Uncertainty 33 mm
|  |  |  | x(Hg) = (Q34 + Q35 +Q36)/ ρ N2 | [10 empty] | Q(Hg) = Q34 + Q35 +Q36
Uncertainty mercury concentration in the generator | [10 empty] | Uncertainty sum diffusion
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y) |  | X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
Q 34 | ng/min | 69.7761 | 0.7 | 1.003% | 1 | 0.7 | 1/d | -0.33278 | -0.23295 | a | Q 34 | ng/min | 69.7761 | 0.7 | 1.003% | 1 | 0.7 | 1 | 1 | 0.7 | a
Q 35 | ng/min | 73.2535 | 0.68696 | 0.9378% | 1 | 0.68696 | 1/d | -0.33278 | -0.22861 | b | Q 35 | ng/min | 73.2535 | 0.68696 | 0.9378% | 1 | 0.68696 | 1 | 1 | 0.68696 | b
Q 36 | ng/min | 71.2258 | 0.66704 | 0.9365% | 1 | 0.66704 | 1/d | -0.33278 | -0.22198 | c | Q 36 | ng/min | 71.2258 | 0.66704 | 0.9365% | 1 | 0.66704 | 1 | 1 | 0.66704 | c
MFC1 | L/min | 0.5 | 0.0015 | 0.3% | 1 | 0.0015 |  |  |  |  | Q | ng/min | 214.2554 | 2.3722 | 1.107% | 2 | 1.1861 | 1
MFC2 | L/min | 2.5 | 0.0075 | 0.3% | 1 | 0.0075
Bios | L/min | 3.004 | 0.006008 | 0.2% | 2 | 0.003004
ρ N2 | L/min | 3.005 | 0.0082173 | 0.2735% | 1 | 0.0082173 | -((a+b+c)/d^2) | -23.727 | -0.19497 | d
x(Hg) | ng/m3 | 71299.6213 | 880.4816 | 1.235% | 2 | 440.2408
Uncertainty mercury mass on transfer standards
m(Hg) = (t* ρ sample) * Q/ρ N2
Uncertainty mercury mass on transfer standards
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
t | min | 150 | 0.03 | 0.02% | 2 | 0.015 | c(d+e+f)/g | 7.1585 | 0.10738 | b
ρ sample | ml/min | 0.1004 | 0.0003514 | 0.5% | 1 | 0.0003514 | b(d+e+f)/g | 10694.9432 | 3.7582 | c
MFC | ml/min | 0.1004 | 0.0003012 | 0.3% | 1 | 0.0003012
MFM | ml/min | 0.1004 | 0.0002008 | 0.2% | 2 | 0.0001004
MPM | ml/min | 0.1004 | 0.0003012 | 0.3% | 2 | 0.0001506
Q 34 | ng/min | 69.7761 | 0.7 | 1.003% | 1 | 0.7 | bc/g | 5.0116 | 3.5082 | d
Q 35 | ng/min | 73.2535 | 0.68696 | 0.9378% | 1 | 0.68696 | bc/g | 5.0116 | 3.4428 | e
Q 36 | ng/min | 71.2258 | 0.66704 | 0.9365% | 1 | 0.66704 | bc/g | 5.0116 | 3.343 | f
ρ N2 | L/min | 3.005 | 0.0082173 | 0.2735% | 1 | 0.0082173 | -bc(f+e+d)/g^2 | -357.3286 | -2.9363 | g
m(Hg) | ng | 1073.7723 | 15.2437 | 1.42% | 2 | 7.6219
m(Hg) = (t* ρ sample) * x(Hg)/1000
Uncertainty mercury mass on transfer standards
X | [X] | x | U(x) | U(x)/x | k | u(x) | Afgeleide | c | u(y)
t | min | 150 | 0.03 | 0.02% | 2 | 0.015 | bc/1000 | 7.1585 | 0.10738 | a
ρ sample | ml/min | 0.1004 | 0.0003514 | 0.5% | 1 | 0.0003514 | ac/1000 | 10694.9432 | 3.7582 | b
MFC | ml/min | 0.1004 | 0.0003012 | 0.3% | 1 | 0.0003012
MFM | ml/min | 0.1004 | 0.0002008 | 0.2% | 2 | 0.0001004
MPM | ml/min | 0.1004 | 0.0003012 | 0.3% | 2 | 0.0001506
x(Hg) | ng/m3 | 71299.6213 | 880.4816 | 1.235% | 2 | 440.2408 | ab/100 | 0.01506 | 6.63 | c
m(Hg) | ng | 1073.7723 | 15.2437 | 1.42% | 2 | 7.6219
Note on A20: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the timer used is 0.03 minutes for a time interval of 0.1 to 100 minutes with a coverage factor of k = 2. For a loading of 25 minutes this leads to a standard uncertainty of 0.0012 % with a coverage factor of k = 2, which is negligible compared to other uncertainty contributions.
Note on A22: Krom, Iris (I.M.J.) de, Mrs: Mass-flow controller output variations are stated by the manufacturer to be < 0.1 % temperature sensitivity of full scale per K of the reading. The uncertainty of mass-flow controller output at full scale, at a temperature uncertainty of 1 K, is equal to (0.1*0.3)2 = 0.03 %. The uncertainty in drift of the mass-flow controller is  0.3 %. By keeping the mass-flow controller set point above 20 %, uMFC  0.3 %.
Note on A23: Krom, Iris (I.M.J.) de, Mrs: The reading of the MFM (mass flow meter) is within ± 0.05 (% set point), resulting in a reading uncertainty of 0.029 (% set point). By keeping the set point above 40%, the relative reading uncertainty is kept below 0.07%. Manufacturers’ specifications of the repeatability of the MFM is < 0.2% of reading. The resulting uncertainty contribution UMFM is 0.2% (k = 2).
Note on A24: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the MPM (mercury piston prover) used is 0.3 % for flow rates from 20 up to 1000 mL.min-1 with a coverage factor of k = 2, UMPM = 0.3 %.
Note on A36: Krom, Iris (I.M.J.) de, Mrs: Mass-flow controller output variations are stated by the manufacturer to be < 0.1 % temperature sensitivity of full scale per K of the reading. The uncertainty of mass-flow controller output at full scale, at a temperature uncertainty of 1 K, is equal to (0.1*0.3)2 = 0.03 %. The uncertainty in drift of the mass-flow controller is  0.3 %. By keeping the mass-flow controller set point above 20 %, uMFC  0.3 %.
Note on A37: Krom, Iris (I.M.J.) de, Mrs: The reading of the MFM (mass flow meter) is within ± 0.05 (% set point), resulting in a reading uncertainty of 0.029 (% set point). By keeping the set point above 40%, the relative reading uncertainty is kept below 0.07%. Manufacturers’ specifications of the repeatability of the MFM is < 0.2% of reading. The resulting uncertainty contribution UMFM is 0.2% (k = 2).
Note on A38: Krom, Iris (I.M.J.) de, Mrs: The standard uncertainty of the MPM (mercury piston prover) used is 0.3 % for flow rates from 20 up to 1000 mL.min-1 with a coverage factor of k = 2, UMPM = 0.3 %.

## 13-11 (G1 = 214.3)
| Datum : | 2018-11-13
| Klant : | JSI
| concentratie generator: |  | 214.3 | ng/min
| belading 1 | mfc |  |  | belading 2 | mfc
| flow meten: | 100.62 | ml/min |  | flow meten: | 499.95 | ml/min
|  | 100.59 | ml/min |  |  | 499.76 | ml/min
|  | 100.55 | ml/min |  |  | 499.95 | ml/min
|  | 100.55 | ml/min |  |  | 499.78 | ml/min
|  | 100.59 | ml/min |  |  | 499.78 | ml/min
| gemiddelde: | 100.58 |  |  | gemiddelde: | 499.844
| MFM | 37.25 | % |  | MFM | 90.75 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 3.0043 | l/min | 3.0043 | l/min | 3.0043 | l/min |  | l/min |  | l/min |  | l/min
| 3.0045 | l/min | 3.0045 | l/min | 3.0045 | l/min |  | l/min |  | l/min |  | l/min
| 3.0057 | l/min | 3.0057 | l/min | 3.0057 | l/min |  | l/min |  | l/min |  | l/min
| 3.0043 | l/min | 3.0043 | l/min | 3.0043 | l/min |  | l/min |  | l/min |  | l/min
| 3.0052 | l/min | 3.0052 | l/min | 3.0052 | l/min |  | l/min |  | l/min |  | l/min
flow | 3.0048 |  | 3.0048 |  | 3.0048 |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| belading 1 | [5 empty] | belading 2 en 3
| positie 1 | 37.2 | % | 100.445 | ml/min |  | positie 1 | 90.8 | % | 500.1194 | ml/min |  | 100.445
| positie 2 | 37.2 | % | 100.445 | ml/min |  | positie 2 | 90.8 | % | 500.1194 | ml/min
| positie 3 | 37.15 | % | 100.31 | ml/min |  | positie 3 | 90.8 | % | 500.1194 | ml/min
| positie 4 | 37.2 | % | 100.445 | ml/min |  | positie 4 | 90.8 | % | 500.1194 | ml/min
| positie 5 | 37.3 | % | 100.715 | ml/min |  | positie 5 | 90.8 | % | 500.1194 | ml/min
| positie 6 | 37.2 | % | 100.445 | ml/min |  | positie 6 | 0 | % | 0 | ml/min
| Belading 1: | 150 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: | 30 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 1074.5488 | ng/buis | 1 | 15 en 23 | 91.4 | positie 1 | 1070.0438 | ng/buis | 1 | 14 en 21 | 103 |  | No breaktrough
| positie 2 | 1074.5488 | ng/buis | 2 | 10 en 24 | 59.1 | positie 2 | 1070.0438 | ng/buis | 2 | 6 en 11 | 98.8
| positie 3 | 1073.1045 | ng/buis | 3 | 18 en 20 | 94.1 | positie 3 | 1070.0438 | ng/buis | 3 | 17 en 1 | 102.3
| positie 4 | 1074.5488 | ng/buis | 4 | 7 en 22 | 92.5 | positie 4 | 1070.0438 | ng/buis | 4 | 19 en 2 | 100.6
| positie 5 |  | ng/buis | 5 | JSI carbon |  | positie 5 |  | ng/buis | 5 | JSI carbon
| positie 6 |  | ng/buis | 6 | trap |  | positie 6 |  | ng/buis | 6 | trap
| Belading 3: | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: |  | min. | positie op beladingsrek | buis
| positie 1 | 535.0219 | ng/buis | 1 | 5 en 9 | 89 | positie 1 | #DIV/0! | ng/buis | 1
| positie 2 | 535.0219 | ng/buis | 2 | 3 en 4 | 87.2 | positie 2 | #DIV/0! | ng/buis | 2
| positie 3 | 535.0219 | ng/buis | 3 | 12 en 16 | 87.8 | positie 3 | #DIV/0! | ng/buis | 3
| positie 4 | 535.0219 | ng/buis | 4 | 13 en 8 | 84.9 | positie 4 | #DIV/0! | ng/buis | 4
| positie 5 |  | ng/buis | 5 | JSI carbon |  | positie 5 | #DIV/0! | ng/buis | 5
| positie 6 |  | ng/buis | 6 | trap |  | positie 6 | #DIV/0! | ng/buis | 6
| Belading 5: |  | min. | positie op beladingsrek | buis |  | Belading 6: |  | min. | positie op beladingsrek | buis
| positie 1 | #DIV/0! | ng/buis | 1 |  |  | positie 1 | #DIV/0! | ng/buis | 1
| positie 2 | #DIV/0! | ng/buis | 2 |  |  | positie 2 | #DIV/0! | ng/buis | 2
| positie 3 | #DIV/0! | ng/buis | 3 |  |  | positie 3 | #DIV/0! | ng/buis | 3
| positie 4 | #DIV/0! | ng/buis | 4 |  |  | positie 4 | #DIV/0! | ng/buis | 4
| positie 5 | #DIV/0! | ng/buis | 5 |  |  | positie 5 | #DIV/0! | ng/buis | 5
| positie 6 | #DIV/0! | ng/buis | 6 |  |  | positie 6 | #DIV/0! | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | 15 en 23 | 10 en 24 | 18 en 20 | 7 en 22 | JSI carbon | trap
meting ( ng)
|  |  | buis dummy | 15 en 23 | 10 en 24 | 18 en 20 | 7 en 22 | JSI carbon | trap
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | #REF! | #REF! | #REF! | #REF! | JSI carbon | trap
meting ( ng)
|  |  | buis dummy | #REF! | #REF! | #REF! | #REF! | JSI carbon | trap
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | 5 en 9 | 3 en 4 | 12 en 16 | 13 en 8 | JSI carbon | trap
meting ( ng)
|  |  | buis dummy | 5 en 9 | 3 en 4 | 12 en 16 | 13 en 8 | JSI carbon | trap
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen

## 14-11 (G2 = 4.23)
| Datum : | 2018-11-14
| Klant : | JSI
| concentratie generator 1: |  | 214.3 | ng/min |  |  | concentratie generator 2: |  | 4.23 | ng/min
| belading 1&2 | mfc |  |  | belading 3 | mfc |  |  | belading 4 | mfc |  |  | belading 5 | mfc |  |  | belading 6 & 7 | mfc |  |  | belading 8 | mfc
| flow meten: | 499.62 | ml/min |  | flow meten: | 100.2 | ml/min |  | flow meten: | 499.73 | ml/min |  | flow meten: | 49.97 | ml/min |  | flow meten: | 499.7 | ml/min |  | flow meten: | 200.09 | ml/min
|  | 499.68 | ml/min |  |  | 100.17 | ml/min |  |  | 499.7 | ml/min |  |  | 50.04 | ml/min |  |  | 499.62 | ml/min |  |  | 200.06 | ml/min
|  | 499.76 | ml/min |  |  | 100.13 | ml/min |  |  | 499.76 | ml/min |  |  | 49.846 | ml/min |  |  | 499.7 | ml/min |  |  | 199.97 | ml/min
|  | 499.87 | ml/min |  |  | 100.16 | ml/min |  |  | 499.68 | ml/min |  |  | 50.1 | ml/min |  |  | 499.79 | ml/min |  |  | 200.04 | ml/min
|  | 499.85 | ml/min |  |  | 100.24 | ml/min |  |  | 499.68 | ml/min |  |  | 49.926 | ml/min |  |  | 499.7 | ml/min |  |  | 200.01 | ml/min
| gemiddelde: | 499.756 |  |  | gemiddelde: | 100.18 |  |  | gemiddelde: | 499.71 |  |  | gemiddelde: | 49.9764 |  |  | gemiddelde: | 499.702 |  |  | gemiddelde: | 200.034
| MFM | 90.75 | % |  | MFM | 37.15 | % |  | MFM | 90.8 | % |  | MFM | 91.75 | % |  | MFM | 90.4 | % |  | MFM | 93.2 | %
| belading 1 |  | belading 2 |  | belading 3 |  | belading 4 |  | belading 5 |  | belading 6 |  | belading 7 |  | belading 8
| systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem |  | systeem
| 3.006 | l/min | 3.006 | l/min | 3.006 | l/min | 3.006 | l/min | 3.006 | l/min | 3.0165 | l/min | 3.0165 | l/min | 3.0165 | l/min
| 3.0056 | l/min | 3.0056 | l/min | 3.0056 | l/min | 3.0056 | l/min | 3.0056 | l/min | 3.0148 | l/min | 3.0148 | l/min | 3.0148 | l/min
| 3.0005 | l/min | 3.0005 | l/min | 3.0005 | l/min | 3.0005 | l/min | 3.0005 | l/min | 3.0164 | l/min | 3.0164 | l/min | 3.0164 | l/min
| 3.0014 | l/min | 3.0014 | l/min | 3.0014 | l/min | 3.0014 | l/min | 3.0014 | l/min | 3.0153 | l/min | 3.0153 | l/min | 3.0153 | l/min
| 3.0055 | l/min | 3.0055 | l/min | 3.0055 | l/min | 3.0055 | l/min | 3.0055 | l/min | 3.0157 | l/min | 3.0157 | l/min | 3.0157 | l/min
flow | 3.0038 |  | 3.0038 |  | 3.0038 |  | 3.0038 |  | 3.0038 |  | 3.0157 |  | 3.0157 |  | 3.0157
| belading 1 & 2 | [5 empty] | belading 3 | [5 empty] | belading 4 | [5 empty] | belading 5
| positie 1 | 90.7 | % | 499.4807 | ml/min |  | positie 1 | 37.15 | % | 100.18 | ml/min |  | positie 1 | 90.8 | % | 499.71 | ml/min |  | positie 1 | 91.45 | % | 49.813 | ml/min
| positie 2 | 90.8 | % | 500.0313 | ml/min |  | positie 2 | 37.2 | % | 100.3148 | ml/min |  | positie 2 | 90.8 | % | 499.71 | ml/min |  | positie 2 | 91.55 | % | 49.8675 | ml/min
| positie 3 | 90.7 | % | 499.4807 | ml/min |  | positie 3 | 37.2 | % | 100.3148 | ml/min |  | positie 3 | 90.8 | % | 499.71 | ml/min |  | positie 3 | 91.6 | % | 49.8947 | ml/min
| positie 4 | 0 | % | 0 | ml/min |  | positie 4 | 0 | % | 0 | ml/min |  | positie 4 | 0 | % | #DIV/0! | ml/min |  | positie 4 | 0 | % | 0 | ml/min
| positie 5 | 0 | % | 0 | ml/min |  | positie 5 | 0 | % | 0 | ml/min |  | positie 5 | 0 | % | #DIV/0! | ml/min |  | positie 5 | 0 | % | 0 | ml/min
| positie 6 | 0 | % | 0 | ml/min |  | positie 6 | 0 | % | 0 | ml/min |  | positie 6 | 0 | % | #DIV/0! | ml/min |  | positie 6 | 0 | % | 0 | ml/min
| Belading 1: G1 | 30 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: G1 | 15 | min. | positie op beladingsrek | buis |  | belading 6 & 7 | [5 empty] | belading 8
| positie 1 | 1069.0329 | ng/buis | 1 | OL502123 | 89.5 | positie 1 | 534.5165 | ng/buis | 1 | OL502086 | 82.8 | positie 1 | 90.4 | % | 499.702 | ml/min |  | positie 1 | 93.25 | % | 200.1413 | ml/min
| positie 2 | 1070.2116 | ng/buis | 2 | OL502067 | 82.3 | positie 2 | 535.1058 | ng/buis | 2 | OL502097 | 79.2 | positie 2 | 90.4 | % | 499.702 | ml/min |  | positie 2 | 93.25 | % | 200.1413 | ml/min
| positie 3 | 1069.0329 | ng/buis | 3 | OL502040 | 83.1 | positie 3 | 534.5165 | ng/buis | 3 | OL502125 | 81.4 | positie 3 | 90.4 | % | 499.702 | ml/min |  | positie 3 | 93.25 | % | 200.1413 | ml/min
| positie 4 |  | ng/buis | 4 |  | 84.9667 | positie 4 |  | ng/buis | 4 |  | 81.1333 | positie 4 | 0 | % | #DIV/0! | ml/min |  | positie 4 | 93.25 | % | 200.1413 | ml/min
| positie 5 |  | ng/buis | 5 | LUMEX | 3.9463 | positie 5 |  | ng/buis | 5 | LUMEX | 1.8148 | positie 5 | 0 | % | #DIV/0! | ml/min |  | positie 5 | 93.25 | % | 200.1413 | ml/min
| positie 6 |  | ng/buis | 6 | tubes | 4.645% | positie 6 |  | ng/buis | 6 | tubes | 2.237% | positie 6 | 0 | % | #DIV/0! | ml/min |  | positie 6 | 93.25 | % | 200.1413 | ml/min
| Belading 3: G1 | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: G1 | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 107.2071 | ng/buis | 1 | 16 en 15 | 94.12% | positie 1 | 534.7619 | ng/buis | 1 | 1 | 94.4 |  | 12.5274
| positie 2 | 107.3514 | ng/buis | 2 | 8 en 7 |  | positie 2 | 534.7619 | ng/buis | 2 | 2 | 94.2 |  | 13.159
| positie 3 | 107.3514 | ng/buis | 3 | 10 en 9 | 93.06% | positie 3 | 534.7619 | ng/buis | 3 | 3 | 93.5
| positie 4 |  | ng/buis | 4 |  |  | positie 4 |  | ng/buis | 4 |  | 94.0333
| positie 5 |  | ng/buis | 5 | Gold (pink) |  | positie 5 |  | ng/buis | 5 | Gold (gold) | 0.47258 |  | 12.8432 | 1.22 | 22%
| positie 6 |  | ng/buis | 6 | tubes |  | positie 6 |  | ng/buis | 6 | tubes | 0.5026%
| Belading 5: G1 | 14 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 6: G2 | 75 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 7: G2 | 150 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 8: G2 | 37.5 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 49.7533 | ng/buis | 1 | 4 | 97.2 | positie 1 | 52.5677 | ng/buis | 1 | 1 | 104.4 | positie 1 | 105.1354 | ng/buis | 1 | 4 | 97.1 | positie 1 | 10.5272 | ng/buis | 1 | 1 | 119
| positie 2 | 49.8077 | ng/buis | 2 | 5 | 95.7 | positie 2 | 52.5677 | ng/buis | 2 | 2 | 105.4 | positie 2 | 105.1354 | ng/buis | 2 | 5 | 96.6 | positie 2 | 10.5272 | ng/buis | 2 | 2 | 125
| positie 3 | 49.8349 | ng/buis | 3 | 6 | 95.7 | positie 3 | 52.5677 | ng/buis | 3 | 3 | 103.7 | positie 3 | 105.1354 | ng/buis | 3 | 6 | 96.1 | positie 3 | 10.5272 | ng/buis | 3 | J+H PSA
| positie 4 |  | ng/buis | 4 |  | 96.2 | positie 4 |  | ng/buis | 4 |  | 104.5 | positie 4 |  | ng/buis | 4 |  | 96.6 | positie 4 | 10.5272 | ng/buis | 4 | C+I PSA
| positie 5 |  | ng/buis | 5 | Gold (gold) | 0.86603 | positie 5 |  | ng/buis | 5 | Gold (gold) | 0.8544 | positie 5 |  | ng/buis | 5 | Gold (gold) | 0.5 | positie 5 | 10.5272 | ng/buis | 5 | G+F PSA
| positie 6 |  | ng/buis | 6 | tubes | 0.9002% | positie 6 |  | ng/buis | 6 | tubes | 0.8176% | positie 6 |  | ng/buis | 6 | tubes | 0.5176% | positie 6 | 10.5272 | ng/buis | 6 | E+D PSA
Belading 1
|  | meet volgorde | buis dummy | OL502123 | OL502067 | OL502040 | 0 | LUMEX | tubes
meting ( ng)
|  |  | buis dummy | OL502123 | OL502067 | OL502040 | 0 | LUMEX | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | #REF! | #REF! | #REF! | #REF! | LUMEX | tubes
meting ( ng)
|  |  | buis dummy | #REF! | #REF! | #REF! | #REF! | LUMEX | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | 16 en 15 | 8 en 7 | 10 en 9 | 0 | Gold (pink) | tubes
meting ( ng)
|  |  | buis dummy | 16 en 15 | 8 en 7 | 10 en 9 | 0 | Gold (pink) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | #REF! | 1 | 3 | 0 | Gold (gold) | tubes
meting ( ng)
|  |  | buis dummy | #REF! | 1 | 3 | 0 | Gold (gold) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 4 | 5 | 6 | 0 | Gold (gold) | tubes
meting ( ng)
|  |  | buis dummy | 4 | 5 | 6 | 0 | Gold (gold) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 1 | 2 | 3 | 0 | Gold (gold) | tubes
meting ( ng)
|  |  | buis dummy | 1 | 2 | 3 | 0 | Gold (gold) | tubes
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | #DIV/0! | #DIV/0! | #DIV/0!
opmerkingen

## 15-11 (G1 = 214.3)
| Datum : | 2018-11-15
| Klant : | JSI
| concentratie generator 1: |  | 214.3 | ng/min
| belading 1 tm 5 | mfc
| flow meten: | 499.9 | ml/min
|  | 499.92 | ml/min
|  | 499.96 | ml/min
|  | 499.84 | ml/min
|  | 499.98 | ml/min
| gemiddelde: | 499.92
| MFM | 90.8 | %
belading 1 tm 5
systeem
| 2.9956 | l/min
| 2.9932 | l/min
| 2.9952 | l/min
| 2.9928 | l/min
| 2.9964 | l/min
flow | 2.9946
belading 1 tm 5
| positie 1 | 90.85 | % | 500.1953 | ml/min
| positie 2 | 90.8 | % | 499.92 | ml/min
| positie 3 | 90.8 | % | 499.92 | ml/min
| positie 4 | 90.85 | % | 500.1953 | ml/min
| positie 5 | 0 | % | 0 | ml/min
| positie 6 | 0 | % | 0 | ml/min
| Belading 1: | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: | 30 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 536.9185 | ng/buis | 1 | 1 | 95.4 | positie 1 | 1073.8371 | ng/buis | 1 | 25 en 26 | :( | 506.8511
| positie 2 | 536.623 | ng/buis | 2 | 2 | 95.2 | positie 2 | 1073.2461 | ng/buis | 2 | 27 en 28 | :( | 500.1327
| positie 3 | 536.623 | ng/buis | 3 | 3 | 93.8 | positie 3 | 1073.2461 | ng/buis | 3 | 30 en 29 | :( | 501.7425
| positie 4 | 536.9185 | ng/buis | 4 | 4 | 94.6 | positie 4 |  | ng/buis | 4 |  |  | 505.2403
| positie 5 |  | ng/buis | 5 | Gold | 94.75 | positie 5 |  | ng/buis | 5 | Carbon traps |  | 503.4917 | 0.93774 | -6.226%
| positie 6 |  | ng/buis | 6 | traps | 0.7188 | positie 6 |  | ng/buis | 6 | Igor Self-made
0.7586%
| Belading 3: | 15 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 536.9185 | ng/buis | 1 | 1 | 96 | positie 1 | 536.9185 | ng/buis | 1 | 1 | 94.6
| positie 2 | 536.623 | ng/buis | 2 | 2 | 95.8 | positie 2 | 536.623 | ng/buis | 2 | 2 | 93.7
| positie 3 | 536.623 | ng/buis | 3 | 3 | 95.4 | positie 3 | 536.623 | ng/buis | 3 | 5 | 94.2
| positie 4 | 536.9185 | ng/buis | 4 | 4 | 93.3 | positie 4 | 536.9185 | ng/buis | 4 | 6 | 93.3
| positie 5 |  | ng/buis | 5 | Gold | 95.125 | positie 5 |  | ng/buis | 5 | Gold | 93.95
| positie 6 |  | ng/buis | 6 | traps | 1.242 | positie 6 |  | ng/buis | 6 | traps | 0.56862
[6 empty] | 1.306% | [5 empty] | 0.6052%
| Belading 5: | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 536.9185 | ng/buis | 1 | 1 | 94.4
| positie 2 | 536.623 | ng/buis | 2 | 2 | 93.2
| positie 3 | 536.623 | ng/buis | 3 | 3 | 93.5
| positie 4 | 536.9185 | ng/buis | 4 | 4 | 94.1
| positie 5 |  | ng/buis | 5 | Gold | 93.8
| positie 6 |  | ng/buis | 6 | traps | 0.54772
0.5839%

## 16-11 (G2 = 4.23)
| Datum : | 2018-11-16
| Klant : | JSI
| concentratie generator 2: |  | 4.23 | ng/min
| belading 1 | mfc |  |  | belading 2 | mfc
| flow meten: | 99.87 | ml/min |  | flow meten: | 199.85 | ml/min
|  | 99.95 | ml/min |  |  | 199.98 | ml/min
|  | 99.92 | ml/min |  |  | 199.89 | ml/min
|  | 99.91 | ml/min |  |  | 200.06 | ml/min
|  | 99.86 | ml/min |  |  | 199.9 | ml/min
| gemiddelde: | 99.902 |  |  | gemiddelde: | 199.936
| MFM | 36.95 | % |  | MFM | 73.45 | %
belading 1, 2 & 3
systeem
| 3.0183 | l/min
| 3.0169 | l/min
| 3.0177 | l/min
| 3.0189 | l/min
| 3.0173 | l/min
flow | 3.0178
| belading 1 | [5 empty] | belading 2
| positie 1 | 37.05 | % | 100.1724 | ml/min |  | positie 1 | 73.45 | % | 199.936 | ml/min
| positie 2 | 37.05 | % | 100.1724 | ml/min |  | positie 2 | 73.45 | % | 199.936 | ml/min
| positie 3 | 37 | % | 100.0372 | ml/min |  | positie 3 | 73.5 | % | 200.0721 | ml/min
| positie 4 | 37.05 | % | 100.1724 | ml/min |  | positie 4 | 73.4 | % | 199.7999 | ml/min
| positie 5 | 0 | % | 0 | ml/min |  | positie 5 | 0 | % | 0 | ml/min
| positie 6 | 0 | % | 0 | ml/min |  | positie 6 | 0 | % | 0 | ml/min
| Belading 1: | 75 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 2: | 37.5 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 10.5307 | ng/buis | 1 | 1 | 95.9 | positie 1 | 10.5092 | ng/buis | 1 | 1 | 98.1 | 10.2885
| positie 2 | 10.5307 | ng/buis | 2 | 2 | 95.3 | positie 2 | 10.5092 | ng/buis | 2 | 2 | 96.4 | 10.3936
| positie 3 | 10.5165 | ng/buis | 3 | 3 | 96.2 | positie 3 | 10.5163 | ng/buis | 3 | 5 | 96.1 | 10.3586
| positie 4 | 10.5307 | ng/buis | 4 | 4 | 96.7 | positie 4 | 10.502 | ng/buis | 4 | 6 | 97.9 | 10.25
| positie 5 |  | ng/buis | 5 | Gold | 96.025 | positie 5 |  | ng/buis | 5 | Gold | 97.125 | 10.3227 | 0.98292 | -1.708%
| positie 6 |  | ng/buis | 6 | traps | 0.58523 | positie 6 |  | ng/buis | 6 | traps | 1.021
[6 empty] | 0.6095% | [5 empty] | 1.051%
| Belading 3: | 37.5 | min. | positie op beladingsrek | buis | Recovery (%) | Belading 4: | 15 | min. | positie op beladingsrek | buis | Recovery (%)
| positie 1 | 10.5092 | ng/buis | 1 | 1 | 97.9 | positie 1 | #REF! | ng/buis | 1
| positie 2 | 10.5092 | ng/buis | 2 | 2 | 98.9 | positie 2 | #REF! | ng/buis | 2
| positie 3 | 10.5163 | ng/buis | 3 | 3 | 98.5 | positie 3 | #REF! | ng/buis | 3
| positie 4 | 10.502 | ng/buis | 4 | 4 | 97.6 | positie 4 | #REF! | ng/buis | 4
| positie 5 |  | ng/buis | 5 |  | 98.225 | positie 5 | #REF! | ng/buis | 5
| positie 6 |  | ng/buis | 6 |  | 0.58523 | positie 6 | #REF! | ng/buis | 6
0.5958%
| Belading 5: |  | min. | positie op beladingsrek | buis |  | Belading 6: |  | min. | positie op beladingsrek | buis
| positie 1 | #REF! | ng/buis | 1 |  |  | positie 1 | #REF! | ng/buis | 1
| positie 2 | #REF! | ng/buis | 2 |  |  | positie 2 | #REF! | ng/buis | 2
| positie 3 | #REF! | ng/buis | 3 |  |  | positie 3 | #REF! | ng/buis | 3
| positie 4 | #REF! | ng/buis | 4 |  |  | positie 4 | #REF! | ng/buis | 4
| positie 5 | #REF! | ng/buis | 5 |  |  | positie 5 | #REF! | ng/buis | 5
| positie 6 | #REF! | ng/buis | 6 |  |  | positie 6 | #REF! | ng/buis | 6
Belading 1
|  | meet volgorde | buis dummy | 1 | 2 | 3 | 4 | Gold | traps
meting ( ng)
|  |  | buis dummy | 1 | 2 | 3 | 4 | Gold | traps
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 2
|  | meet volgorde | buis dummy | #REF! | #REF! | #REF! | #REF! | Gold | traps
meting ( ng)
|  |  | buis dummy | #REF! | #REF! | #REF! | #REF! | Gold | traps
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 3
|  | meet volgorde | buis dummy | 1 | 2 | 3 | 4 | 0 | 0
meting ( ng)
|  |  | buis dummy | 1 | 2 | 3 | 4 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | -100 | -100 | -100 | -100 | #DIV/0! | #DIV/0!
opmerkingen
Belading 4
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
opmerkingen
Belading 5
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
opmerkingen
Belading 6
|  | meet volgorde | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
meting ( ng)
|  |  | buis dummy | 0 | 0 | 0 | 0 | 0 | 0
afwijking t.o.v theoretisch berekend (%) |  |  | x | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
opmerkingen