## Physical Description
STANDARD DATA - 1 MINUTE (CW) | [10 empty] | Does the 904nm laser need time to reach a plate? Experiment 904 vs 830 - 3 MINUTE (CW)
Equip. | Company | Model | Wavelenght (nm) | Power Manual (mW) | Power Evaluated (mW) |  |  | Spot Beam Size (cm²) | SBS Evaluated (cm²) | Uncertainty SBS | Equip. | Company | Model | Wavelenght (nm) | Power Manual (mW) | Power (mW) - 1 min / control |  |  | Power (mW) - 3 min / test
[5 empty] | Maximum | Mean | SD | [8 empty] | Maximum | Mean | SD | Maximum | Mean | SD
1 | Ibramed | Antares | 660 | 40 | 44.68 | 42 | 0.009 | 0.07 | 0.098757 | 0.006442 | 5 | Ibramed | Laserpulse | 830 | 30 | 34.07 | 33.06 | 2.27 | 33.91 | 32.69 | 2.18
2 | Ibramed | Laserpulse | 904 | 70 | 43.14 | 41.93 | 4.76 | 0.1309 | 0.084187 | 0.04612 | 6 | Ibramed | Laserpulse | 830 | 30 | 0.8858 | 0.6141 | 0.185 | 5.262 | 4.513 | 0.53
3 | Ibramed | Laserpulse | 450 | 70 | 65.78 | 62.37 | 6.63 | 0.1 | 0.15122 | 0.006528 | 7 | Ibramed | Laserpulse | 830 | 30 | 41.83 | 40.28 | 4.09 | 42.17 | 40.56 | 2.28
4 | Ibramed | Laserpulse | 660 | 30 | 34.07 | 33.06 | 2.27 | 0.0631 | 0.07724 | 0.006564 | 8 | Ibramed | Laserpulse | 830 | 30 | 34.07 | 33.63 | 0.826 | 35.34 | 34.94 | 0.736
5 | Ibramed | Laserpulse | 830 | 30 | 34.07 | 33.06 | 2.27 | 0.116 | 0.048305 | 0.00714 | 9 | Ibramed | Laserpulse | 830 | 30 | 34.53 | 33.77 | 2.83 | 33.91 | 33.53 | 1.04
6 | Ibramed | Laserpulse | 830 | 30 | 0.8858 | 0.6141 | 0.185 | 0.116 |  |  | 10 | Ibramed | Laserpulse | 830 | 30 | 30.9 | 30.4 | 1.77 | 32.97 | 32.43 | 0.977
7 | Ibramed | Laserpulse | 830 | 30 | 41.83 | 40.28 | 4.09 | 0.116 | 0.17261 | 0.011856 | 19 | Ibramed | Laserpulse | 830 | 30 | 32.36 | 31.25 | 2.04 | 34.34 | 32.73 | 1.64
8 | Ibramed | Laserpulse | 830 | 30 | 34.07 | 33.63 | 0.826 | 0.116 | 0.043595 | 0.002504
9 | Ibramed | Laserpulse | 830 | 30 | 34.53 | 33.77 | 2.83 | 0.116 |  |  | 2 | Ibramed | Laserpulse | 904 | 70 | 43.14 | 41.93 | 4.76 | 45.73 | 44.46 | 2.38
10 | Ibramed | Laserpulse | 830 | 30 | 30.9 | 30.4 | 1.77 | 0.116 | 0.086987 | 0.011844 | 13 | Ibramed | Laserpulse | 904 | 70000 | 53.16 | 52.08 | 3.57 | 53.74 | 52.92 | 2.38
11 | Ibramed | Laserpulse | 660 | 30 | 19.14 | 18.08 | 0.835 | 0.0631 | 0.07383 | 0.012238 | 16 | Ibramed | Laserpulse | 904 | 70000 | 51.07 | 50.04 | 4.44 | 51.33 | 50.81 | 2.66
12 | Ibramed | Laserpulse | 660 | 30 | 17.72 | 16.93 | 0.484 | 0.0631 | 0.061049 | 0.007708 | 17 | Ibramed | Laserpulse | 904 | 70000 | 42.63 | 41.86 | 3.96 | 43.38 | 42.59 | 2.43
13 | Ibramed | Laserpulse | 904 | 70 | 53.16 | 52.08 | 3.57 | 0.1309 | 0.14644 | 0.04686 | 18 | Ibramed | Laserpulse | 904 | 70000 | 44.35 | 43.08 | 4.57 | 44.64 | 43.93 | 2.71
14 | Ibramed | Laserpulse | 660 | 30 | 35.12 | 34.28 | 0.416 | 0.0631 | 0.050671 | 0.004396 | 20 | HTM | Fluence | 904 | 13 | 8.014 | 7.639 | 0.306 | 8.978 | 8.177 | 0.472
15 | Ibramed | Laserpulse | 660 | 30 | 28.54 | 27.13 | 3.56 | 0.0631 | 0.17735 | 0.009712 | 21 | Ibramed | Antares | 904 | 70000 | 74.65 | 72.08 | 11.2 | 75.76 | 74.13 | 6.78
16 | Ibramed | Laserpulse | 904 | 70 | 51.07 | 50.04 | 4.44 | 0.1309 | 0.075574 | 0.01499 | 23 | HTM | Fluence | 904 | 13 | 11.61 | 11.43 | 0.593 | 11.8 | 11.68 | 0.0826
17 | Ibramed | Laserpulse | 904 | 70 | 42.63 | 41.86 | 3.96 | 0.1309 | 0.13605 | 0.018542
18 | Ibramed | Laserpulse | 904 | 70 | 44.35 | 43.08 | 4.57 | 0.1309 | 0.094787 | 0.012014
19 | Ibramed | Laserpulse | 830 | 30 | 32.36 | 31.25 | 2.04 | 0.116 | 0.07383 | 0.005376
20 | HTM | Fluence | 904 | 13 | 8.014 | 7.639 | 0.306 | 0.069
21 | Ibramed | Antares | 904 | 70 | 74.65 | 72.08 | 11.2 | 0.07 | 0.26786 | 0.04068
23 | HTM | Fluence | 904 | 13 | 11.61 | 11.43 | 0.593 | 0.069 | 0.37501 | 0.05324
24 | DMC | Therapy XT | 808 | 100 | 99.83 | 95.52 | 13.7 | 0.1884 | 0.13973 | 0.02008
25 | DMC | Therapy XT | 660 | 100 | 100.6 | 103.4 | 2.52 | 0.1884 | 0.10463 | 0.011252
NORMALIZED
DATA
Equip. | Power Evaluated (mW) |  |  | SBS Evaluated (cm²)
| Maximum | Mean | SD
1 | 111.7% | 105% | 0.02143%
2 | 61.63% | 59.9% | 11.35%
3 | 93.97% | 89.1% | 10.63%
4 | 113.6% | 110.2% | 6.866%
5 | 113.6% | 110.2% | 6.866%
6 | 2.953% | 2.047% | 30.13%
7 | 139.4% | 134.3% | 10.15%
8 | 113.6% | 112.1% | 2.456%
9 | 115.1% | 112.6% | 8.38%
10 | 103% | 101.3% | 5.822%
11 | 63.8% | 60.27% | 4.618%
12 | 59.07% | 56.43% | 2.859%
13 | 75.94% | 74.4% | 6.855%
14 | 117.1% | 114.3% | 1.214%
15 | 95.13% | 90.43% | 13.12%
16 | 72.96% | 71.49% | 8.873%
17 | 60.9% | 59.8% | 9.46%
18 | 63.36% | 61.54% | 10.61%
19 | 107.9% | 104.2% | 6.528%
20 | 61.65% | 58.76% | 4.006%
21 | 106.6% | 103% | 15.54%
23 | 89.31% | 87.92% | 5.188%
24 | 99.83% | 95.52% | 14.34%
25 | 100.6% | 103.4% | 2.437%

## General Description
GENERAL CONDITION
Equip. | Company | Model | Wavelenght (nm) | Owner | Last maintenance (month) | Average time of use/week | Acquisition time | Months/Acquisition | last maintenance
1 | Ibramed | Antares | 660 | LARF | 12 | <30 | <12 | 3 | 3
2 | Ibramed | Laserpulse | 904 | LARF | 12 to 24 | <30 | >120 | 168 | 24
3 | Ibramed | Laserpulse | 450 | LARF | 12 to 24 | <30 | >120 | 168 | 24
4 | Ibramed | Laserpulse | 660 | LARF | 12 to 24 | <30 | >120 | 168 | 24
5 | Ibramed | Laserpulse | 830 | LARF | 12 to 24 | <30 | >120 | 168 | 24
6 | Ibramed | Laserpulse | 830 | LARF | 12 to 24 | <30 | >120 | 168 | 24
7 | Ibramed | Laserpulse | 830 | LARF | 12 to 24 | <30 | >120 | 168 | 24
8 | Ibramed | Laserpulse | 830 | CER | over 36 | <30 | >120 | 168 | 36
9 | Ibramed | Laserpulse | 830 | CER | over 36 | <30 | >120 | 168 | 36
10 | Ibramed | Laserpulse | 830 | CER | over 36 | <30 | >120 | 168 | 36
11 | Ibramed | Laserpulse | 660 | CER | over 36 | <30 | >120 | 168 | 36
12 | Ibramed | Laserpulse | 660 | CER | over 36 | <30 | >120 | 168 | 36
13 | Ibramed | Laserpulse | 904 | CER | over 36 | <30 | >120 | 168 | 36
14 | Ibramed | Laserpulse | 660 | Multi | - | <30 | >120 | 168 | 168
15 | Ibramed | Laserpulse | 660 | Multi | - | <30 | >120 | 168 | 168
16 | Ibramed | Laserpulse | 904 | Multi | - | <30 | >120 | 168 | 168
17 | Ibramed | Laserpulse | 904 | Multi | - | <30 | >120 | 168 | 168
18 | Ibramed | Laserpulse | 904 | Multi | - | <30 | >120 | 168 | 168
19 | Ibramed | Laserpulse | 830 | Multi | - | <30 | >120 | 168 | 168
20 | HTM | Fluence | 904 | CSE | 12 to 24 | <30 | 5 to 10 | 96 | 24
21 | Ibramed | Antares | 904 | LARF | 12 | <30 | <12 | 3 | 3
23 | HTM | Fluence | 904 | Supere | 12 | 30 to 90 | 5 to 10 | 72 | 12
24 | DMC | Therapy XT | 808 | First | 12 | 30 to 90 | <12 | 12 | 12
25 | DMC | Therapy XT | 660 | First | 12 | 30 to 90 | <12 | 12 | 12
60

## General (Time)
1 Min de Coleta (mW) - CW |  |  |  | 1 Min de Coleta (mW) - PW 150Hz |  |  |  | 3 Min de Coleta (mW) - CW |  |  |  | 3 Min de Coleta (mW) - PW 150Hz |  |  |  | 450nm
Equipment | Maximum | Mean | Std. Dev. | Equipment | Maximum | Mean | Std. Dev. | Equipment | Maximum | Mean | Std. Dev. | Equipment | Maximum | Mean | Std. Dev. | 660nm
3 | 65.78 | 62.37 | 6.63 | 3 | - | - | - | 3 | - | - | - | 3 | - | - | - | 808nm
830nm
1 | 44.68 | 42 | 0.009 | 1 | - | - | - | 1 | - | - | - | 1 | - | - | - | 904nm
4 | 34.07 | 33.06 | 2.27 | 4 | - | - | - | 4 | - | - | - | 4 | - | - | -
11 | 19.14 | 18.08 | 0.835 | 11 | - | - | - | 11 | - | - | - | 11 | - | - | -
12 | 17.72 | 16.93 | 0.484 | 12 | - | - | - | 12 | - | - | - | 12 | - | - | -
14 | 35.12 | 34.28 | 0.416 | 14 | 19.39 | 18.47 | 1.74 | 14 | - | - | - | 14 | - | - | -
15 | 28.54 | 27.13 | 3.56 | 15 | 17.46 | 16.92 | 1.46 | 15 | - | - | - | 15 | - | - | -
25 | 100.6 | 103.4 | 2.52 | 25 | - | - | - | 25 | - | - | - | 25 | - | - | -
24 | 99.83 | 95.52 | 13.7 | 24 | - | - | - | 24 | - | - | - | 24 | - | - | -
5 | 32.6 | 31.45 | 3.36 | 5 | 15.77 | 15 | 1.82 | 5 | 33.91 | 32.69 | 2.18 | 5 | 15.94 | 15.21 | 1.05
6 | 0.8858 | 0.6141 | 0.185 | 6 | - | - | - | 6 | 5.262 | 4.513 | 0.53 | 6 | - | - | -
7 | 41.83 | 40.28 | 4.09 | 7 | 20.33 | 19.53 | 1.93 | 7 | 42.17 | 40.56 | 2.28 | 7 | 21.53 | 20.5 | 1.35
8 | 34.07 | 33.63 | 0.826 | 8 | 19 | 18.47 | 0.485 | 8 | 35.34 | 34.94 | 0.736 | 8 | 15.25 | 15.07 | 0.38
9 | 34.53 | 33.77 | 2.83 | 9 | 16.21 | 15.95 | 1.03 | 9 | 33.91 | 33.53 | 1.04 | 9 | 16.58 | 16.36 | 0.802
10 | 30.9 | 30.4 | 1.77 | 10 | 12.94 | 12.73 | 0.678 | 10 | 32.97 | 32.43 | 0.977 | 10 | 16.13 | 15.96 | 0.63
19 | 32.36 | 31.25 | 2.04 | 19 | 14.93 | 14.64 | 1.29 | 19 | 34.34 | 32.73 | 1.64 | 19 | 14.93 | 14.77 | 0.759
Mean | 29.5965 | 28.7706 |  |  | 16.53 | 16.0533 |  |  | 31.1289 | 30.199 |  |  | 16.7267 | 16.3117
SD | 13.1445 | 12.8442 |  |  | 2.7081 | 2.5334 |  |  | 11.8174 | 11.6782 |  |  | 2.428 | 2.1354
2 | 43.14 | 41.93 | 4.76 | 2 | - | - | - | 2 | 45.73 | 44.46 | 2.38 | 13 | - | - | -
13 | 53.16 | 52.08 | 3.57 | 13 | - | - | - | 13 | 53.74 | 52.92 | 2.38 | 2 | - | - | -
16 | 51.07 | 50.04 | 4.44 | 16 | - | - | - | 16 | 51.33 | 50.81 | 2.66 | 16 | - | - | -
17 | 42.63 | 41.86 | 3.96 | 17 | - | - | - | 17 | 43.38 | 42.59 | 2.43 | 17 | - | - | -
18 | 44.35 | 43.08 | 4.57 | 18 | - | - | - | 18 | 44.64 | 43.93 | 2.71 | 18 | - | - | -
20 | 8.014 | 7.639 | 0.306 | 20 | - | - | - | 20 | 8.978 | 8.177 | 0.472 | 20 | - | - | -
21 | 74.65 | 72.08 | 11.2 | 21 | - | - | - | 21 | 75.76 | 74.13 | 6.78 | 21 | - | - | -
23 | 11.61 | 11.43 | 0.593 | 23 | - | - | - | 23 | 11.8 | 11.68 | 0.0826 | 23 | - | - | -
Mean | 41.078 | 40.0174 | [6 empty] | 41.9198 | 41.0871
SD | 21.8823 | 21.2334 | [6 empty] | 22.006 | 21.6991

## General (660nm)
Device / 0mm | EQ1 | EQ4 | EQ11 | EQ12 | EQ14 | EQ15 | EQ25 | Device / 3mm | EQ1 | EQ4 | EQ11 | EQ12 | EQ14 | EQ15 | EQ25 | Device / 6mm | EQ1 | EQ4 | EQ11 | EQ12 | EQ14 | EQ15 | EQ25 | Device / 9mm | EQ1 | EQ4 | EQ11 | EQ12 | EQ14 | EQ15 | EQ25 | Device / 12mm | EQ1 | EQ4 | EQ11 | EQ12 | EQ14 | EQ15 | EQ25 | Device / 15mm | EQ1 | EQ4 | EQ11 | EQ12 | EQ14 | EQ15 | EQ25
0 | 0.28 | 0.19 | 0.1 | 0.07 | 0.23 | 0.12 | 0.25 | 0 | 0.16 | 0.15 | 0.028 | 0.03 | 0.09 | 0.06 | 0.159 | 0 | 0.1 | 0.04 | 0.02 | 0.02 | 0.04 | 0.03 | 0.13 | 0 | 0.06 | 0.028 | 0.006 | 0.008 | 0.028 | 0.03 | 0.16 | 0 | 0.038 | 0.028 | 0.009 | 0.008 | 0.023 | 0.02 | 0.06 | 0 | 0.021 | 0.023 | 0.004 | 0.008 | 0.021 | 0.023 | 0.069
0.5 | 0.43 | 0.34 | 0.15 | 0.1 | 0.4 | 0.19 | 0.35 | 0.5 | 0.19 | 0.28 | 0.04 | 0.06 | 0.14 | 0.08 |  | 0.5 | 0.13 | 0.07 |  |  | 0.05 | 0.048 |  | 0.5 | 0.07 |  |  |  | 0.036 |  |  | 0.5 | [7 empty] | 0.5
1 | 0.64 | 0.74 | 0.2 | 0.2 | 0.73 | 0.26 | 1.34 | 1 | 0.26 | 0.44 | 0.057 | 0.08 | 0.21 | 0.1 | 0.42 | 1 | 0.16 | 0.11 | 0.036 |  | 0.07 | 0.069 |  | 1 | 0.08 | 0.05 |  |  | 0.046 | 0.047 |  | 1 | 0.051 |  |  |  | 0.037 |  |  | 1 | 0.025 |  |  |  | 0.03
1.5 | 0.92 | 1.23 | 0.28 | 0.34 | 1.19 | 0.28 | 1.8 | 1.5 | 0.35 | 0.59 | 0.075 | 0.11 | 0.24 | 0.13 |  | 1.5 | 0.19 | 0.16 |  |  | 0.09 | 0.11 |  | 1.5 | 0.089 |  |  |  | 0.056 |  |  | 1.5 |  | 0.068 | [5 empty] | 1.5
2 | 1.31 | 1.86 | 0.38 | 0.4 | 1.73 | 0.41 | 2.5 | 2 | 0.45 | 0.74 | 0.095 | 0.14 | 0.33 | 0.16 | 0.74 | 2 | 0.22 | 0.23 | 0.046 | 0.044 | 0.13 | 0.13 | 0.28 | 2 | 0.094 | 0.1 | 0.011 |  | 0.07 | 0.063 | 0.28 | 2 | 0.068 |  | 0.013 |  | 0.053 | 0.039 | 0.09 | 2 | 0.032 | 0.06 |  |  | 0.04 | 0.038 | 0.08
2.5 | 1.65 | 2.26 | 0.5 | 0.48 | 2.04 | 0.56 | 2.94 | 2.5 | 0.57 | 0.89 | 0.11 | 0.18 | 0.47 | 0.2 |  | 2.5 | 0.26 | 0.28 |  |  | 0.16 | 0.15 |  | 2.5 | 0.1 |  |  |  | 0.087 |  |  | 2.5 | [7 empty] | 2.5
3 | 2 | 2.49 | 0.66 | 0.48 | 2.16 | 0.74 | 4.5 | 3 | 0.7 | 1 | 0.14 | 0.22 | 0.59 | 0.25 | 1.77 | 3 | 0.31 | 0.33 | 0.055 |  | 0.21 | 0.18 |  | 3 | 0.12 | 0.16 |  | 0.022 | 0.11 | 0.087 |  | 3 | 0.081 | 0.11 |  | 0.016 | 0.071 |  |  | 3 | 0.04 |  | 0.01 | 0.012 | 0.051
3.5 | 2.26 | 2.42 | 0.81 | 0.42 | 1.95 | 0.91 | 5.94 | 3.5 | 0.86 | 1.04 | 0.17 | 0.25 | 0.7 | 0.29 |  | 3.5 | 0.35 | 0.38 |  |  | 0.26 | 0.22 |  | 3.5 | 0.15 |  |  |  | 0.13 |  |  | 3.5 | [7 empty] | 3.5
4 | 2.43 | 2.19 | 0.94 | 0.3 | 1.57 | 1.09 | 6.83 | 4 | 1.02 | 1.08 | 0.21 | 0.34 | 0.79 | 0.29 | 3.38 | 4 | 0.39 | 0.47 | 0.075 | 0.083 | 0.29 | 0.24 | 0.62 | 4 | 0.18 | 0.22 | 0.022 |  | 0.15 | 0.11 | 0.4 | 4 | 0.086 |  | 0.019 |  | 0.099 | 0.065 | 0.14 | 4 | 0.055 | 0.1 |  |  | 0.066 | 0.055 | 0.1
4.5 | 2.38 | 1.83 | 0.92 | 0.22 | 0.98 | 1.29 | 7.25 | 4.5 | 1.16 | 1.07 | 0.22 | 0.38 | 0.84 | 0.36 |  | 4.5 | 0.44 | 0.52 |  |  | 0.29 | 0.29 |  | 4.5 | 0.21 |  |  |  | 0.18 |  |  | 4.5 |  | 0.18 | [5 empty] | 4.5
5 | 2.15 | 1.48 | 0.74 | 0.14 | 0.44 | 1.46 | 7.46 | 5 | 1.3 | 1 | 0.24 | 0.37 | 0.86 | 0.4 | 5.2 | 5 | 0.49 | 0.56 | 0.086 |  | 0.32 | 0.25 |  | 5 | 0.24 | 0.29 |  |  | 0.2 | 0.13 |  | 5 | 0.1 |  |  |  | 0.13 |  |  | 5 | 0.06 |  |  |  | 0.085
5.5 | 1.73 | 1.02 | 0.53 | 0.1 | 0.23 | 1.57 | 7.1 | 5.5 | 1.45 | 0.91 | 0.26 | 0.34 | 0.8 | 0.48 |  | 5.5 | 0.54 | 0.56 |  |  | 0.37 | 0.27 |  | 5.5 | 0.27 |  |  |  | 0.22 |  |  | 5.5 | [7 empty] | 5.5
6 | 1.31 | 0.63 | 0.33 | 0.06 |  | 1.62 | 6.8 | 6 | 1.49 | 0.82 | 0.29 | 0.31 | 0.71 | 0.54 | 6.3 | 6 | 0.59 | 0.57 | 0.1 | 0.11 | 0.39 | 0.3 | 1.44 | 6 | 0.31 | 0.33 | 0.036 | 0.043 | 0.24 | 0.17 | 0.9 | 6 | 0.11 | 0.22 | 0.026 | 0.025 | 0.15 | 0.094 | 0.16 | 6 | 0.062 | 0.15 | 0.016 | 0.018 | 0.1 | 0.074 | 0.12
6.5 | 0.86 | 0.4 | 0.22 |  |  | 1.58 | 5.89 | 6.5 | 1.53 | 0.69 | 0.32 | 0.27 | 0.59 | 0.57 |  | 6.5 | 0.65 | 0.58 |  |  | 0.42 | 0.32 |  | 6.5 | 0.34 |  |  |  | 0.26 |  |  | 6.5 | [7 empty] | 6.5
7 | 0.47 | 0.22 | 0.11 |  |  | 1.59 | 4.45 | 7 | 1.53 | 0.54 | 0.28 | 0.23 | 0.46 | 0.63 | 6.7 | 7 | 0.69 | 0.56 | 0.12 |  | 0.43 | 0.33 |  | 7 | 0.36 | 0.36 |  |  | 0.28 | 0.19 |  | 7 | 0.14 |  |  |  | 0.17 |  |  | 7 | 0.082 |  |  |  | 0.11
7.5 | 0.22 | 0.1 |  |  |  | 1.42 | 2.54 | 7.5 | 1.48 | 0.41 | 0.25 | 0.2 | 0.34 | 0.67 |  | 7.5 | 0.74 | 0.54 |  |  | 0.41 | 0.33 |  | 7.5 | 0.39 |  |  |  | 0.29 |  |  | 7.5 |  | 0.24 | [5 empty] | 7.5
8 | [5 empty] | 1.31 | 1.3 | 8 | 1.33 | 0.32 | 0.19 | 0.17 | 0.29 | 0.71 | 6.66 | 8 | 0.76 | 0.52 | 0.13 | 0.16 | 0.4 | 0.35 | 2.26 | 8 | 0.43 | 0.36 | 0.041 |  | 0.24 | 0.22 | 1.14 | 8 | 0.17 |  | 0.028 |  | 0.19 | 0.12 | 0.27 | 8 | 0.093 | 0.17 |  |  | 0.13 | 0.092 | 0.18
8.5 | [5 empty] | 1.05 | 0.23 | 8.5 | 1.16 | 0.23 | 0.16 | 0.14 | 0.18 | 0.7 |  | 8.5 | 0.76 | 0.48 |  |  | 0.37 | 0.36 |  | 8.5 | 0.45 |  |  |  | 0.24 |  |  | 8.5 | [7 empty] | 8.5
9 | [5 empty] | 0.77 | 0.144 | 9 | 0.93 | 0.17 | 0.15 | 0.12 | 0.1 | 0.69 | 4.99 | 9 | 0.78 | 0.42 | 0.15 |  | 0.33 | 0.38 |  | 9 | 0.47 | 0.35 |  | 0.064 | 0.29 | 0.25 |  | 9 | 0.21 | 0.24 |  | 0.035 | 0.21 |  |  | 9 | 0.11 |  | 0.017 | 0.019 | 0.14
9.5 | [5 empty] | 0.51 | 0.02 | 9.5 | 0.76 |  | 0.13 | 0.1 | 0.05 | 0.68 |  | 9.5 | 0.76 | 0.37 |  |  | 0.28 | 0.39 |  | 9.5 | 0.48 |  |  |  | 0.29 |  |  | 9.5 | [7 empty] | 9.5
10 | [5 empty] | 0.3 |  | 10 | 0.56 |  | 0.12 | 0.09 |  | 0.65 | 3.7 | 10 | 0.74 | 0.3 | 0.11 | 0.1 | 0.24 | 0.38 | 2.41 | 10 | 0.5 | 0.34 | 0.071 |  | 0.29 | 0.24 | 1.24 | 10 | 0.24 |  | 0.05 |  | 0.21 | 0.15 | 0.5 | 10 | 0.13 | 0.18 |  |  | 0.15 | 0.11 | 0.31
10.5 | [5 empty] | 0.2 |  | 10.5 | 0.42 |  | 0.11 | 0.09 |  | 0.65 |  | 10.5 | 0.72 | 0.25 |  |  | 0.26 | 0.36 |  | 10.5 | 0.5 |  |  |  | 0.28 |  |  | 10.5 |  | 0.24 | [5 empty] | 10.5
11 | [5 empty] | 0.11 |  | 11 | 0.29 |  | 0.11 | 0.08 |  | 0.61 | 2.58 | 11 | 0.65 | 0.21 | 0.087 |  | 0.2 | 0.34 |  | 11 | 0.49 | 0.28 |  |  | 0.27 | 0.25 |  | 11 | 0.27 |  |  |  | 0.21 |  |  | 11 | 0.15 |  |  |  | 0.15
[8 empty] | 11.5 | 0.19 |  | 0.097 | 0.07 |  | 0.57 |  | 11.5 | 0.61 | 0.17 |  |  | 0.16 | 0.31 |  | 11.5 | 0.49 |  |  |  | 0.24 |  |  | 11.5 | [7 empty] | 11.5
[8 empty] | 12 |  |  | 0.087 | 0.05 |  | 0.5 | 1.32 | 12 | 0.54 | 0.13 | 0.075 | 0.06 | 0.12 | 0.3 | 2.19 | 12 | 0.47 | 0.22 | 0.074 | 0.053 | 0.22 | 0.27 | 1.21 | 12 | 0.3 | 0.21 | 0.051 | 0.028 | 0.2 | 0.18 | 0.68 | 12 | 0.17 | 0.17 | 0.037 | 0.024 | 0.16 | 0.13 | 0.43
[8 empty] | 12.5 |  |  | 0.079 | 0.05 |  | 0.42 |  | 12.5 | 0.45 | 0.1 |  |  | 0.09 | 0.28 |  | 12.5 | 0.43 |  |  |  | 0.19 |  |  | 12.5 | [7 empty] | 12.5
[8 empty] | 13 |  |  | 0.071 | 0.03 |  | 0.34 | 0.45 | 13 | 0.38 | 0.08 | 0.069 |  | 0.06 | 0.25 |  | 13 | 0.38 | 0.17 |  |  | 0.17 | 0.29 |  | 13 | 0.31 |  |  |  | 0.18 |  |  | 13 | 0.18 |  |  |  | 0.15
[8 empty] | 13.5 |  |  | 0.056 |  |  | 0.27 |  | 13.5 | 0.32 | 0.06 |  |  | 0.03 | 0.24 |  | 13.5 | 0.34 |  |  |  | 0.14 |  |  | 13.5 |  | 0.16 | [5 empty] | 13.5
[8 empty] | 14 |  |  | 0.045 |  |  | 0.27 | 0.13 | 14 | 0.25 |  | 0.073 | 0.06 |  | 0.27 | 1.19 | 14 | 0.29 | 0.13 | 0.079 |  | 0.12 | 0.25 | 1.14 | 14 | 0.32 |  | 0.048 |  | 0.15 | 0.18 | 0.68 | 14 | 0.19 | 0.15 |  |  | 0.14 | 0.13 | 0.42
[8 empty] | 14.5 |  |  | 0.035 |  |  | 0.2 |  | 14.5 | 0.2 |  |  |  |  | 0.25 |  | 14.5 | 0.25 |  |  |  | 0.1 |  |  | 14.5 | [7 empty] | 14.5
[8 empty] | 15 | [5 empty] | 0.15 |  | 15 | 0.15 |  | 0.065 |  |  | 0.22 |  | 15 | 0.23 | 0.09 |  | 0.04 | 0.079 | 0.23 |  | 15 | 0.33 | 0.12 |  | 0.028 | 0.12 |  |  | 15 | 0.22 |  | 0.034 | 0.021 | 0.12
[8 empty] | 15.5 | [5 empty] | 0.1 |  | 15.5 | [5 empty] | 0.2 |  | 15.5 | 0.22 |  |  |  | 0.063 |  |  | 15.5 | [7 empty] | 15.5
[8 empty] | 16 | [5 empty] | 0.6 |  | 16 |  |  | 0.062 | 0.05 |  | 0.18 | 0.27 | 16 | 0.19 | 0.06 | 0.071 |  | 0.049 | 0.21 | 0.82 | 16 | 0.31 |  | 0.055 |  | 0.095 | 0.19 | 0.73 | 16 | 0.21 | 0.11 |  |  | 0.1 | 0.14 | 0.45
[16 empty] | 16.5 | [5 empty] | 0.16 |  | 16.5 | 0.18 |  |  |  | 0.034 |  |  | 16.5 |  | 0.07 | [5 empty] | 16.5
[16 empty] | 17 |  |  | 0.048 |  |  | 0.13 |  | 17 | 0.16 | 0.04 |  |  | 0.023 | 0.19 |  | 17 | 0.28 |  |  |  | 0.071 |  |  | 17 | 0.23 |  |  |  | 0.085
[16 empty] | 17.5 | [5 empty] | 0.11 |  | 17.5 | 0.13 | [6 empty] | 17.5 | [7 empty] | 17.5
[16 empty] | 18 |  |  | 0.033 | 0.003 |  | 0.099 |  | 18 | 0.11 | 0.024 | 0.043 | 0.037 |  | 0.17 | 0.37 | 18 | 0.19 | 0.05 | 0.032 | 0.033 | 0.047 | 0.16 | 0.65 | 18 | 0.24 | 0.07 | 0.035 | 0.017 | 0.073 | 0.13 | 0.46
[16 empty] | 18.5 | [5 empty] | 0.084 |  | 18.5 | 0.09 | [6 empty] | 18.5 | [7 empty] | 18.5
[16 empty] | 19 |  |  | 0.025 |  |  | 0.072 |  | 19 | 0.07 |  |  |  |  | 0.14 |  | 19 | 0.15 |  |  |  | 0.031 |  |  | 19 | 0.23 |  |  |  | 0.055
[16 empty] | 19.5 | [5 empty] | 0.06 |  | 19.5 | [7 empty] | 19.5 |  | 0.03 | [5 empty] | 19.5
[16 empty] | 20 |  |  |  | 0.002 |  | 0.053 |  | 20 |  |  | 0.04 |  |  | 0.12 | 0.11 | 20 | 0.13 |  | 0.029 |  | 0.016 | 0.14 | 0.46 | 20 | 0.22 | 0.04 |  |  | 0.045 | 0.11 | 0.46
[16 empty] | 20.5 | [5 empty] | 0.043 |  | 20.5 | [7 empty] | 20.5 | [7 empty] | 20.5
[16 empty] | 21 | [5 empty] | 0.03 |  | 21 |  |  |  | 0.035 |  | 0.1 |  | 21 | 0.13 |  |  | 0.028 |  |  |  | 21 | 0.15 |  | 0.029 | 0.017 | 0.029
[24 empty] | 21.5 | [7 empty] | 21.5 | [7 empty] | 21.5
[24 empty] | 22 |  |  | 0.041 |  |  | 0.083 |  | 22 | 0.11 |  | 0.025 |  |  | 0.11 | 0.26 | 22 | 0.12 | 0.02 |  |  | 0.017 | 0.1 | 0.42
[24 empty] | 22.5 | [7 empty] | 22.5 | [7 empty] | 22.5
[24 empty] | 23 | [5 empty] | 0.061 |  | 23 | 0.08 | [6 empty] | 23 | 0.08
[24 empty] | 23.5 | [7 empty] | 23.5 | [7 empty] | 23.5
[24 empty] | 24 |  |  | 0.039 | 0.02 |  | 0.048 |  | 24 |  |  | 0.025 | 0.025 |  | 0.092 | 0.08 | 24 | 0.091 |  | 0.018 | 0.015 |  | 0.089 | 0.41
[24 empty] | 24.5 | [7 empty] | 24.5 | [7 empty] | 24.5
[24 empty] | 25 | [5 empty] | 0.037 |  | 25 | [7 empty] | 25 | 0.09
[24 empty] | 25.5 | [7 empty] | 25.5 | [7 empty] | 25.5
[24 empty] | 26 |  |  | 0.021 |  |  | 0.028 |  | 26 |  |  | 0.025 |  |  | 0.069 |  | 26 | 0.09 |  |  |  |  | 0.073 | 0.36
[24 empty] | 26.5 | [7 empty] | 26.5 | [7 empty] | 26.5
[24 empty] | 27 |  |  |  | 0.012 |  |  |  | 27 |  |  |  | 0.013 |  |  |  | 27 | 0.07 |  | 0.015 | 0.009
[24 empty] | 27.5 | [7 empty] | 27.5 | [7 empty] | 27.5
[24 empty] | 28 |  |  | 0.01 |  |  |  |  | 28 |  |  | 0.012 |  |  | 0.04 |  | 28 | 0.05 |  |  |  |  | 0.054 | 0.25
[24 empty] | 28.5 | [7 empty] | 28.5 | [7 empty] | 28.5
[24 empty] | 29 | [7 empty] | 29 | [7 empty] | 29
[24 empty] | 29.5 | [7 empty] | 29.5 | [7 empty] | 29.5
[24 empty] | 30 |  |  | 0.01 | 0.006 |  |  |  | 30 |  |  | 0.01 | 0.01 |  | 0.027 |  | 30 |  |  | 0.017 |  |  | 0.038 | 0.13
[32 empty] | 30.5 | [7 empty] | 30.5
[32 empty] | 31 | [7 empty] | 31
[32 empty] | 31.5 | [7 empty] | 31.5
[32 empty] | 32 |  |  | 0.005 |  |  |  |  | 32 | [5 empty] | 0.027 | 0.47
[32 empty] | 32.5 | [7 empty] | 32.5
[32 empty] | 33 |  |  |  | 0.008 |  |  |  | 33 |  |  | 0.014
33.5
34
34.5
35
35.5
[40 empty] | 36 |  |  | 0.008

## General (808nm)
Device / 0mm | EQ5 | Device / 3mm | EQ5 | Device /6mm | EQ5 | Device / 9mm | EQ5 | Device / 12mm | EQ5 | Device / 15mm | EQ5
0 | 0.258 | 0 | 0.114 | 0 | 0.133 | 0 | 0.07 | 0 | 0.046 | 0 | 0.024
0.5 | 0.56 | 0.5 | 0.281 | 0.5 | 0.158 | 0.5 | 0.1 | 0.5 | 0.08 | 0.5
1 | 1.13 | 1 | 0.33 | 1 | 0.191 | 1 | 0.15 | 1 | 0.1 | 1
1.5 | 0.65 | 1.5 | 0.5 | 1.5 | 0.227 | 1.5 | 0.2 | 1.5 | 0.15 | 1.5
2 | 2.09 | 2 | 0.66 | 2 | 0.46 | 2 | 0.22 | 2 | 0.18 | 2 | 0.12
2.5 | 2.33 | 2.5 | 0.88 | 2.5 | 0.56 | 2.5 | 0.24 | 2.5 |  | 2.5
3 | 2.69 | 3 | 1.1 | 3 | 0.69 | 3 | 0.24 | 3 | 0.166 | 3
3.5 | 3.37 | 3.5 | 1.6 | 3.5 | 0.87 | 3.5 | 0.21 | 3.5 |  | 3.5
4 | 3.95 | 4 | 2.22 | 4 | 1.02 | 4 | 0.23 | 4 | 0.1 | 4 | 0.077
4.5 | 4.63 | 4.5 | 2.54 | 4.5 | 1.16 | 4.5 | 0.33 | 4.5 |  | 4.5
5 | 5.11 | 5 | 2.77 | 5 | 1.17 | 5 | 0.39 | 5 | 0.088 | 5
5.5 | 5.22 | 5.5 | 2.88 | 5.5 | 1.32 | 5.5 | 0.45 | 5.5 |  | 5.5
6 | 5.14 | 6 | 2.82 | 6 | 1.38 | 6 | 0.53 | 6 | 0.1 | 6 | 0.057
6.5 | 4.9 | 6.5 | 2.62 | 6.5 | 1.44 | 6.5 | 0.58 | 6.5 |  | 6.5
7 | 4.49 | 7 | 2.32 | 7 | 1.48 | 7 | 0.66 | 7 | 0.16 | 7
7.5 | 4.15 | 7.5 | 1.94 | 7.5 | 1.45 | 7.5 | 0.69 | 7.5 |  | 7.5
8 | 2.8 | 8 | 1.5 | 8 | 1.38 | 8 | 0.72 | 8 | 0.25 | 8 | 0.08
8.5 | 1.18 | 8.5 | 1.12 | 8.5 | 1.17 | 8.5 | 0.73 | 8.5 |  | 8.5
9 | 0.4 | 9 | 0.8 | 9 | 1.01 | 9 | 0.75 | 9 | 0.27 | 9
9.5 | 0.22 | 9.5 | 0.53 | 9.5 | 0.85 | 9.5 | 0.77 | 9.5 |  | 9.5
|  | 10 | 0.37 | 10 | 0.7 | 10 | 0.78 | 10 | 0.36 | 10 | 0.18
|  | 10.5 | 0.24 | 10.5 | 0.63 | 10.5 | 0.77 | 10.5 |  | 10.5
|  | 11 | 0.177 | 11 | 0.56 | 11 | 0.7 | 11 | 0.38 | 11
|  |  |  | 11.5 | 0.52 | 11.5 | 0.68 | 11.5 |  | 11.5
|  |  |  | 12 | 0.45 | 12 | 0.64 | 12 | 0.4 | 12 | 0.24
|  |  |  | 12.5 | 0.41 | 12.5 | 0.6 | 12.5 |  | 12.5
|  |  |  | 13 | 0.36 | 13 | 0.53 | 13 | 0.4 | 13
|  |  |  | 13.5 | 0.26 | 13.5 | 0.49 | 13.5 |  | 13.5
|  |  |  | 14 | 0.156 | 14 | 0.43 | 14 | 0.39 | 14 | 0.27
[6 empty] | 14.5 | 0.38 | 14.5 |  | 14.5
[6 empty] | 15 | 0.3 | 15 | 0.38 | 15
[6 empty] | 15.5 | 0.25 | 15.5 |  | 15.5
[6 empty] | 16 | 0.194 | 16 | 0.41 | 16 | 0.25
[6 empty] | 16.5 | 0.15 | 16.5 |  | 16.5
[6 empty] | 17 | 0.11 | 17 | 0.44 | 17
[6 empty] | 17.5 | 0.1 | 17.5 |  | 17.5
[6 empty] | 18 | 0.099 | 18 | 0.42 | 18 | 0.24
[6 empty] | 18.5 | 0.09 | 18.5 |  | 18.5
[8 empty] | 19 | 0.37 | 19
[8 empty] | 19.5 |  | 19.5
[8 empty] | 20 | 0.31 | 20 | 0.26
[8 empty] | 20.5 |  | 20.5
[8 empty] | 21 | 0.26 | 21
[8 empty] | 21.5 |  | 21.5
[8 empty] | 22 | 0.24 | 22 | 0.2
[8 empty] | 22.5 |  | 22.5
[8 empty] | 23 | 0.18 | 23
[8 empty] | 23.5 |  | 23.5
[8 empty] | 24 | 0.11 | 24 | 0.18
[8 empty] | 24.5 |  | 24.5
[8 empty] | 25 | 0.08 | 25
25.5
[10 empty] | 26 | 0.15
26.5
27
27.5
[10 empty] | 28 | 0.066
28.5
29
29.5
[10 empty] | 30 | 0.037

## General (830nm)
Device / 0mm | EQ5 | EQ6 | EQ7 | EQ8 | EQ9 | EQ10 | EQ19 | Device / 3mm | EQ5 | EQ6 | EQ7 | EQ8 | EQ9 | EQ10 | EQ19 | Device / 6mm | EQ5 | EQ6 | EQ7 | EQ8 | EQ9 | EQ10 | EQ19 | Device / 9mm | EQ5 | EQ6 | EQ7 | EQ8 | EQ9 | EQ10 | EQ19 | Device / 12mm | EQ5 | EQ6 | EQ7 | EQ8 | EQ9 | EQ10 | EQ19 | Device / 15mm | EQ5 | EQ6 | EQ7 | EQ8 | EQ9 | EQ10 | EQ19
0 | 0.25 |  | 0.16 | 0.215 |  | 0.13 | 0.046 | 0 | 0.1 |  | 0.1 | 0.15 |  | 0.08 | 0.035 | 0 | 0.043 |  | 0.07 | 0.048 |  | 0.03 | 0.023 | 0 | 0.03 |  | 0.029 | 0.03 |  | 0.02 | 0.016 | 0 | 0.019 |  | 0.02 | 0.032 |  | 0.017 | 0.015 | 0 | 0.015 |  | 0.01 | 0.018 |  | 0.012 | 0.012
0.5 | 0.55 |  | 0.32 | 0.47 |  | 0.3 | 0.086 | 0.5 | 0.16 |  | 0.15 | 0.31 |  | 0.15 | 0.05 | 0.5 |  |  |  | 0.093 |  |  |  | 0.5 | [7 empty] | 0.5 | [7 empty] | 0.5
1 | 0.9 |  | 0.57 | 0.85 |  | 0.56 | 0.16 | 1 | 0.26 |  | 0.2 | 0.51 |  | 0.22 | 0.068 | 1 | 0.82 |  | 0.1 | 0.18 |  | 0.068 | 0.044 | 1 | 0.06 |  |  | 0.07 |  | 0.04 | 0.02 | 1 | [6 empty] | 0.023 | 1
1.5 | 1.3 |  | 0.83 | 1.32 |  | 0.79 | 0.24 | 1.5 | 0.36 |  | 0.25 | 0.69 |  | 0.25 | 0.09 | 1.5 |  |  |  | 0.28 |  |  |  | 1.5 | [7 empty] | 1.5 |  |  |  | 0.1 |  |  |  | 1.5 |  |  |  | 0.067
2 | 1.73 |  | 1.05 | 1.81 |  | 0.93 | 0.34 | 2 | 0.51 |  | 0.28 | 0.85 |  | 0.27 | 0.1 | 2 | 0.15 |  | 0.14 | 0.39 |  | 0.1 | 0.068 | 2 | 0.09 |  | 0.06 | 0.18 |  | 0.06 | 0.033 | 2 | 0.04 |  | 0.04 |  |  | 0.039 | 0.039 | 2 | 0.032 |  | 0.018 |  |  | 0.047 | 0.022
2.5 | 2.24 |  | 1.24 | 2.09 |  | 0.98 | 0.4 | 2.5 | 0.64 |  | 0.32 | 0.97 |  | 0.3 | 0.13 | 2.5 |  |  |  | 0.43 |  |  |  | 2.5 | [7 empty] | 2.5 | [7 empty] | 2.5
3 | 2.42 |  | 1.39 | 1.95 |  | 1.12 | 0.44 | 3 | 0.8 |  | 0.34 | 1.01 |  | 0.34 | 0.17 | 3 | 0.23 |  | 0.13 | 0.48 |  | 0.12 | 0.096 | 3 | 0.13 |  |  | 3 |  | 0.079 | 0.053 | 3 |  |  |  | 0.21 |  |  | 0.05 | 3 |  |  |  | 0.136
3.5 | 2.03 |  | 1.54 | 1.56 |  | 1.26 | 0.41 | 3.5 | 0.92 |  | 0.43 | 0.98 |  | 0.45 | 0.19 | 3.5 |  |  |  | 0.51 |  |  |  | 3.5 | [7 empty] | 3.5 | [7 empty] | 3.5
4 | 1.42 |  | 1.65 | 1 |  | 1.18 | 0.35 | 4 | 0.98 |  | 0.55 | 0.86 |  | 0.55 | 0.23 | 4 | 0.33 |  | 0.26 | 0.54 |  | 0.19 | 0.12 | 4 | 0.17 |  | 0.08 | 0.34 |  | 0.13 | 0.076 | 4 | 0.085 |  | 0.058 |  |  | 0.12 | 0.058 | 4 | 0.057 |  | 0.041 |  |  | 0.11 | 0.041
4.5 | 0.85 |  | 1.51 | 0.55 |  | 0.93 | 0.29 | 4.5 | 0.98 |  | 0.66 | 0.65 |  | 0.6 | 0.25 | 4.5 |  |  |  | 0.55 |  |  |  | 4.5 | [7 empty] | 4.5 |  |  |  | 0.2 |  |  |  | 4.5 |  |  |  | 0.146
5 | 0.42 |  | 1.41 | 0.207 |  | 0.7 | 0.2 | 5 | 0.96 |  | 0.79 | 0.45 |  | 0.58 | 0.28 | 5 | 0.41 |  | 0.38 | 0.55 |  | 0.3 | 0.16 | 5 | 0.21 |  |  | 0.32 |  | 0.2 | 0.08 | 5 | [6 empty] | 0.069 | 5
5.5 | EQ5: 0.19; EQ7: 1.25; EQ10: 0.43; EQ19: 0.13 | 5.5 | 0.86 |  | 0.85 | 0.28 |  | 0.5 | 0.3 | 5.5 |  |  |  | 0.52 |  |  |  | 5.5 | [7 empty] | 5.5 | [7 empty] | 5.5
6 | EQ7: 1.13; EQ10: 0.28; EQ19: 0.083 | 6 | 0.7 |  | 0.89 | 0.153 |  | 0.4 | 0.29 | 6 | 0.46 |  | 0.47 | 0.45 |  | 0.34 | 0.19 | 6 | 0.24 |  | 0.082 | 0.3 |  | 0.22 | 0.1 | 6 | 0.13 |  | 0.049 | 0.19 |  | 0.15 | 0.069 | 6 | 0.082 |  | 0.033 | 0.139 |  | 0.099 | 0.043
6.5 | EQ7: 0.94; EQ10: 0.15; EQ19: 0.061 | 6.5 | 0.61 |  | 0.87 |  |  | 0.31 | 0.3 | 6.5 |  |  |  | 0.35 |  |  |  | 6.5 | [7 empty] | 6.5 | [7 empty] | 6.5
7 | EQ7: 0.82; EQ19: 0.034 | 7 | 0.49 |  | 0.86 |  |  | 0.2 | 0.3 | 7 | 0.47 |  | 0.46 | 0.22 |  | 0.29 | 0.2 | 7 | 0.26 |  |  | 0.28 |  | 0.19 | 0.13 | 7 | [6 empty] | 0.07 | 7
7.5 | EQ7: 0.66 | 7.5 | 0.37 |  | 0.85 |  |  | 0.13 | 0.29 | 7.5 |  |  |  | 0.13 |  |  |  | 7.5 | [7 empty] | 7.5 |  |  |  | 0.19 |  |  |  | 7.5 |  |  |  | 0.135
8 | EQ7: 0.48 | 8 | 0.26 |  | 0.81 |  |  | 0.09 | 0.26 | 8 | 0.43 |  | 0.48 | 0.072 |  | 0.2 | 0.19 | 8 | 0.27 |  | 0.2 | 0.26 |  | 0.17 | 0.15 | 8 | 0.16 |  | 0.088 |  |  | 0.11 | 0.07 | 8 | 0.1 |  | 0.045 |  |  | 0.082 | 0.038
8.5 | EQ7: 0.35 | 8.5 |  |  | 0.78 |  |  |  | 0.23 | 8.5 |  |  |  | 0.035 |  |  |  | 8.5 | [7 empty] | 8.5 | [7 empty] | 8.5
9 | EQ7: 0.25 | 9 |  |  | 0.73 |  |  |  | 0.2 | 9 | 0.34 |  | 0.47 |  |  | 0.11 | 0.16 | 9 | 0.27 |  |  | 0.16 |  | 0.13 | 0.15 | 9 |  |  |  | 0.17 |  |  | 0.075 | 9 |  |  |  | 0.14
9.5 | EQ7: 0.17 | 9.5 |  |  | 0.68 |  |  |  | 0.16 | 9.5 | [7 empty] | 9.5 | [7 empty] | 9.5 | [7 empty] | 9.5
[8 empty] | 10 |  |  | 0.61 |  |  |  | 0.11 | 10 | 0.26 |  | 0.45 |  |  | 0.05 | 0.19 | 10 | 0.24 |  | 0.28 | 0.056 |  | 0.08 | 0.15 | 10 | 0.18 |  | 0.17 |  |  | 0.08 | 0.075 | 10 | 0.12 |  | 0.064 |  |  | 0.069 | 0.031
[8 empty] | 10.5 |  |  | 0.57 |  |  |  | 0.081 | 10.5 | [7 empty] | 10.5 | [7 empty] | 10.5 |  |  |  | 0.068 |  |  |  | 10.5 |  |  |  | 0.145
[8 empty] | 11 |  |  | 0.51 |  |  |  | 0.057 | 11 | 0.18 |  | 0.42 |  |  | 0.018 | 0.14 | 11 | 0.21 |  |  | 0.021 |  | 0.047 | 0.14 | 11 | [6 empty] | 0.081 | 11
[8 empty] | 11.5 |  |  | 0.44 |  |  |  | 0.033 | 11.5 | [7 empty] | 11.5 | [7 empty] | 11.5 | [7 empty] | 11.5
[8 empty] | 12 |  |  | 0.38 |  |  |  |  | 12 | 0.11 |  | 0.39 |  |  |  | 0.11 | 12 | 0.16 |  | 0.29 |  |  | 0.02 | 0.13 | 12 | 0.17 |  | 0.189 | 0.017 |  | 0.03 | 0.079 | 12 | 0.13 |  | 0.12 | 0.077 |  | 0.038 | 0.045
[8 empty] | 12.5 |  |  | 0.34 |  |  |  |  | 12.5 | [7 empty] | 12.5 | [7 empty] | 12.5 | [7 empty] | 12.5
[8 empty] | 13 |  |  | 0.29 |  |  |  |  | 13 | 0.06 |  | 0.35 |  |  |  | 0.082 | 13 | 0.12 | [5 empty] | 0.11 | 13 | [6 empty] | 0.1 | 13
[8 empty] | 13.5 |  |  | 0.26 |  |  |  |  | 13.5 | [7 empty] | 13.5 | [7 empty] | 13.5 | [7 empty] | 13.5 |  |  |  | 0.019
[8 empty] | 14 |  |  | 0.21 |  |  |  |  | 14 |  |  | 0.31 |  |  |  | 0.051 | 14 | 0.09 |  | 0.3 |  |  |  | 0.1 | 14 | 0.14 |  | 0.2 |  |  | 0.007 | 0.1 | 14 | 0.12 |  | 0.13 |  |  | 0.012 | 0.044
[8 empty] | 14.5 |  |  | 0.18 |  |  |  |  | 14.5 | [7 empty] | 14.5 | [7 empty] | 14.5 | [7 empty] | 14.5
[8 empty] | 15 |  |  | 0.15 |  |  |  |  | 15 |  |  | 0.26 |  |  |  | 0.034 | 15 | 0.06 | [5 empty] | 0.08 | 15 | [6 empty] | 0.096 | 15
[16 empty] | 15.5 | [7 empty] | 15.5 | [7 empty] | 15.5 | [7 empty] | 15.5
[16 empty] | 16 |  |  | 0.21 |  |  |  | 0.017 | 16 | 0.04 |  | 0.27 |  |  |  | 0.055 | 16 | 0.09 |  | 0.196 |  |  |  | 0.07 | 16 | 0.1 |  | 0.14 |  |  | 0.003 | 0.054
[16 empty] | 16.5 | [7 empty] | 16.5 | [7 empty] | 16.5 | [7 empty] | 16.5
[16 empty] | 17 |  |  | 0.18 |  |  |  |  | 17 | [6 empty] | 0.047 | 17 | [6 empty] | 0.044 | 17
[16 empty] | 17.5 | [7 empty] | 17.5 | [7 empty] | 17.5 | [7 empty] | 17.5
[16 empty] | 18 |  |  | 0.14 |  |  |  |  | 18 |  |  | 0.24 |  |  |  | 0.035 | 18 | 0.05 |  | 0.189 |  |  |  | 0.026 | 18 | 0.078 |  | 0.14 |  |  |  | 0.056
[16 empty] | 18.5 | [7 empty] | 18.5 | [7 empty] | 18.5 | [7 empty] | 18.5
[16 empty] | 19 |  |  | 0.11 |  |  |  |  | 19 | [6 empty] | 0.017 | 19 | [6 empty] | 0.027 | 19
[16 empty] | 19.5 | [7 empty] | 19.5 | [7 empty] | 19.5 | [7 empty] | 19.5
[16 empty] | 20 |  |  | 0.09 |  |  |  |  | 20 |  |  | 0.2 |  |  |  |  | 20 | 0.03 |  | 0.177 |  |  |  | 0.018 | 20 | 0.05 |  | 0.138 |  |  |  | 0.036
[16 empty] | 20.5 | [7 empty] | 20.5 | [7 empty] | 20.5 | [7 empty] | 20.5
[16 empty] | 21 |  |  | 0.07 |  |  |  |  | 21 | [7 empty] | 21 | [7 empty] | 21
[24 empty] | 21.5 | [7 empty] | 21.5 | [7 empty] | 21.5
[24 empty] | 22 |  |  | 0.15 |  |  |  |  | 22 | 0.014 |  | 0.157 |  |  |  |  | 22 | 0.035 |  | 0.13 |  |  |  | 0.024
[24 empty] | 22.5 | [7 empty] | 22.5 | [7 empty] | 22.5
[24 empty] | 23 | [7 empty] | 23 | [7 empty] | 23
[24 empty] | 23.5 | [7 empty] | 23.5 | [7 empty] | 23.5
[24 empty] | 24 |  |  | 0.12 |  |  |  |  | 24 |  |  | 0.14 |  |  |  |  | 24 | 0.019 |  | 0.12 |  |  |  | 0.013
[24 empty] | 24.5 | [7 empty] | 24.5 | [7 empty] | 24.5
[24 empty] | 25 | [7 empty] | 25 | [7 empty] | 25
[24 empty] | 25.5 | [7 empty] | 25.5 | [7 empty] | 25.5
[24 empty] | 26 |  |  | 0.08 |  |  |  |  | 26 |  |  | 0.114 |  |  |  |  | 26 |  |  | 0.115
[24 empty] | 26.5 | [7 empty] | 26.5 | [7 empty] | 26.5
[24 empty] | 27 | [7 empty] | 27 | [7 empty] | 27
[24 empty] | 27.5 | [7 empty] | 27.5 | [7 empty] | 27.5
[24 empty] | 28 |  |  | 0.06 |  |  |  |  | 28 |  |  | 0.091 |  |  |  |  | 28 |  |  | 0.105
[24 empty] | 28.5 | [7 empty] | 28.5 | [7 empty] | 28.5
[24 empty] | 29 | [7 empty] | 29 | [7 empty] | 29
[24 empty] | 29.5 | [7 empty] | 29.5 | [7 empty] | 29.5
[24 empty] | 30 |  |  | 0.03 |  |  |  |  | 30 |  |  | 0.074 |  |  |  |  | 30 |  |  | 0.09
[32 empty] | 30.5 | [7 empty] | 30.5
[32 empty] | 31 | [7 empty] | 31
[32 empty] | 31.5 | [7 empty] | 31.5
[32 empty] | 32 |  |  | 0.055 |  |  |  |  | 32 |  |  | 0.07
[32 empty] | 32.5 | [7 empty] | 32.5
[32 empty] | 33 | [7 empty] | 33
[32 empty] | 33.5 | [7 empty] | 33.5
[32 empty] | 34 |  |  | 0.041 |  |  |  |  | 34 |  |  | 0.06
[32 empty] | 34.5 | [7 empty] | 34.5
[32 empty] | 35 | [7 empty] | 35
[32 empty] | 35.5 | [7 empty] | 35.5
[32 empty] | 36 |  |  | 0.029 |  |  |  |  | 36 |  |  | 0.05
36.5
37
37.5
[40 empty] | 38 |  |  | 0.04
38.5
39
39.5
[40 empty] | 40 |  |  | 0.03
40.5
41
41.5
[40 empty] | 42 |  |  | 0.02
42.5
43
43.5
[40 empty] | 44 |  |  | 0.018

## General (904nm)
Device / 0mm | EQ2 | EQ13 | EQ16 | EQ17 | EQ18 | EQ20 | EQ21 | EQ23 | Device / 3mm | EQ2 | EQ13 | EQ16 | EQ17 | EQ18 | EQ20 | EQ21 | EQ23 | Device / 6mm | EQ2 | EQ13 | EQ16 | EQ17 | EQ18 | EQ20 | EQ21 | EQ23 | Device / 9mm | EQ2 | EQ13 | EQ16 | EQ17 | EQ18 | EQ20 | EQ21 | EQ23 | Device / 12mm | EQ2 | EQ13 | EQ16 | EQ17 | EQ18 | EQ20 | EQ21 | EQ23 | Device / 15mm | EQ2 | EQ13 | EQ16 | EQ17 | EQ18 | EQ20 | EQ21 | EQ23
0 | 0.1 | 0.85 | 0.063 | 0.038 | 0.06 |  | 0.079 | 0.005 | 0 | 0.094 | 0.06 | 0.085 | 0.05 | 0.072 |  | 0.072 | 0.004 | 0 | 0.043 | 0.045 | 0.048 | 0.056 | 0.03 |  | 0.046 | 0.004 | 0 | 0.04 | 0.015 | 0.03 | 0.02 | 0.036 |  | 0.035 | 0.0023 | 0 | 0.018 | 0.01 | 0.032 | 0.017 | 0.026 |  | 0.021 | 0.0025 | 0 | 0.01 | 0.01 | 0.024 | 0.013 | 0.017 |  | 0.024 | 0.0015
0.5 | 0.12 | 0.36 | 0.14 | 0.11 | 0.19 |  | 0.1 | 0.01 | 0.5 | 0.13 | 0.09 | 0.13 | 0.095 | 0.12 |  | 0.098 |  | 0.5 | 0.055 |  | 0.079 | [5 empty] | 0.5 |  |  | 0.047 | [5 empty] | 0.5 |  |  | 0.049 | [5 empty] | 0.5
1 | 0.14 | 0.57 | 0.26 | 0.26 | 0.38 |  | 0.15 | 0.017 | 1 | 0.19 | 0.18 | 0.21 | 0.16 | 0.2 |  | 0.15 | 0.009 | 1 | 0.073 | 0.063 | 0.12 | 0.11 | 0.08 |  | 0.079 |  | 1 |  |  | 0.071 |  | 0.076 |  |  |  | 1 |  |  | 0.07 | 0.035 | 0.044 |  |  |  | 1 |  |  | 0.051
1.5 | 0.17 | 0.72 | 0.42 | 0.41 | 0.54 |  | 0.21 | 0.026 | 1.5 | 0.25 | 0.35 | 0.28 | 0.23 | 0.28 |  | 0.22 |  | 1.5 | 0.093 |  | 0.17 | [5 empty] | 1.5 | 0.077 |  | 0.1 | [5 empty] | 1.5 |  |  | 0.099 | [5 empty] | 1.5
2 | 0.24 | 0.93 | 0.51 | 0.49 | 0.69 |  | 0.28 | 0.037 | 2 | 0.35 | 0.41 | 0.37 | 0.34 | 0.38 |  | 0.28 | 0.019 | 2 | 0.11 | 0.087 | 0.23 | 0.2 | 0.17 |  | 0.13 | 0.016 | 2 |  | 0.03 | 0.14 | 0.088 | 0.13 |  | 0.078 | 0.007 | 2 | 0.033 | 0.018 | 0.12 | 0.07 | 0.074 |  | 0.044 | 0.008 | 2 | 0.017 | 0.015 | 0.088 | 0.052 | 0.043 |  | 0.038 | 0.004
2.5 | 0.004 | 1.1 | 0.61 | 0.55 | 0.83 |  | 0.38 | 0.048 | 2.5 | 0.41 | 0.46 | 0.44 | 0.41 | 0.45 |  | 0.36 |  | 2.5 | 0.14 |  | 0.27 | [5 empty] | 2.5 |  |  | 0.18 | [5 empty] | 2.5 |  |  | 0.15 | [5 empty] | 2.5
3 | 0.01 | 1.15 | 0.71 | 0.57 | 0.93 |  | 0.44 | 0.059 | 3 | 0.48 | 0.49 | 0.51 | 0.46 | 0.5 |  | 0.39 | 0.033 | 3 | 0.17 | 0.1 | 0.34 | 0.27 | 0.27 |  | 0.19 |  | 3 | 0.13 |  | 0.2 |  | 0.2 |  |  |  | 3 |  |  | 0.16 |  | 0.11 |  |  |  | 3 |  |  | 0.12
3.5 | 0.06 | 1.09 | 0.79 | 0.59 | 0.95 |  | 0.53 | 0.064 | 3.5 | 0.54 | 0.5 | 0.56 | 0.49 | 0.54 |  | 0.43 |  | 3.5 | 0.2 |  | 0.35 | [5 empty] | 3.5 |  |  | 0.22 | [5 empty] | 3.5 |  |  | 0.16 | [5 empty] | 3.5
4 | 0.3 | 0.99 | 0.83 | 0.59 | 0.79 |  | 0.67 | 0.059 | 4 | 0.63 | 0.51 | 0.6 | 0.51 | 0.58 |  | 0.46 | 0.043 | 4 | 0.25 | 0.12 | 0.37 | 0.29 | 0.38 |  | 0.21 | 0.032 | 4 |  | 0.049 | 0.22 | 0.16 | 0.26 |  | 0.15 | 0.021 | 4 | 0.069 | 0.029 | 0.16 | 0.15 | 0.16 |  | 0.081 | 0.017 | 4 | 0.033 | 0.019 | 0.13 | 0.11 | 0.093 |  | 0.064 | 0.011
4.5 | 0.74 | 0.83 | 0.73 | 0.57 | 0.64 |  | 0.83 | 0.053 | 4.5 | 0.68 | 0.52 | 0.63 | 0.52 | 0.61 |  | 0.5 |  | 4.5 | 0.29 |  | 0.38 | [5 empty] | 4.5 | 0.2 |  | 0.22 | [5 empty] | 4.5 |  |  | 0.15 | [5 empty] | 4.5
5 | 1.05 | 0.63 | 0.57 | 0.53 | 0.53 |  | 0.97 | 0.049 | 5 | 0.73 | 0.51 | 0.64 | 0.53 | 0.64 |  | 0.52 | 0.038 | 5 | 0.36 | 0.16 | 0.4 | 0.32 | 0.42 |  | 0.23 |  | 5 |  |  | 0.22 |  | 0.29 |  |  |  | 5 |  |  | 0.15 |  | 0.2 |  |  |  | 5 |  |  | 0.11
5.5 | 0.87 | 0.43 | 0.38 | 0.5 | 0.44 |  | 1.06 | 0.047 | 5.5 | 0.76 | 0.49 | 0.65 | 0.55 | 0.66 |  | 0.55 |  | 5.5 | 0.39 |  | 0.42 | [5 empty] | 5.5 |  |  | 0.24 | [5 empty] | 5.5 |  |  | 0.15 | [5 empty] | 5.5
6 | 0.68 | 0.3 | 0.12 | 0.42 | 0.36 |  | 1.09 | 0.044 | 6 | 0.78 | 0.44 | 0.65 | 0.57 | 0.68 |  | 0.57 | 0.033 | 6 | 0.41 | 0.21 | 0.44 | 0.33 | 0.44 |  | 0.29 | 0.025 | 6 | 0.25 | 0.065 | 0.25 | 0.17 | 0.33 |  | 0.21 | 0.022 | 6 | 0.11 | 0.037 | 0.16 | 0.17 | 0.22 |  | 0.11 | 0.018 | 6 | 0.059 | 0.027 | 0.11 | 0.12 | 0.15 |  | 0.084 | 0.016
6.5 | 0.79 | 0.21 | 0.039 | 0.31 | 0.21 |  | 1.07 | 0.041 | 6.5 | 0.78 | 0.39 | 0.65 | 0.58 | 0.69 |  | 0.59 |  | 6.5 | 0.43 |  | 0.45 | [5 empty] | 6.5 |  |  | 0.26 | [5 empty] | 6.5 |  |  | 0.16 | [5 empty] | 6.5
7 | 1.17 | 0.11 |  | 0.17 | 0.13 |  | 1.06 | 0.039 | 7 | 0.76 | 0.34 | 0.63 | 0.58 | 0.69 |  | 0.6 | 0.03 | 7 | 0.44 | 0.27 | 0.47 | 0.36 | 0.46 |  | 0.36 |  | 7 |  |  | 0.27 |  | 0.33 |  |  |  | 7 |  |  | 0.17 |  | 0.22 |  |  |  | 7 |  |  | 0.11
7.5 | 1.01 | 0.46 |  | 0.1 | 0.06 |  | 0.96 | 0.036 | 7.5 | 0.72 | 0.29 | 0.58 | 0.56 | 0.68 |  | 0.61 |  | 7.5 | 0.45 |  | 0.48 | [5 empty] | 7.5 | 0.28 |  | 0.28 | [5 empty] | 7.5 |  |  | 0.17 | [5 empty] | 7.5
8 | 0.59 |  |  | 0.04 |  |  | 0.88 | 0.034 | 8 | 0.66 | 0.25 | 0.49 | 0.55 | 0.65 |  | 0.62 | 0.028 | 8 | 0.46 | 0.34 | 0.48 | 0.38 | 0.47 |  | 0.39 | 0.021 | 8 |  | 0.089 | 0.33 | 0.17 | 0.34 |  | 0.26 | 0.018 | 8 | 0.15 | 0.046 | 0.18 | 0.15 | 0.22 |  | 0.14 | 0.015 | 8 | 0.089 | 0.03 | 0.12 | 0.11 | 0.16 |  | 0.1 | 0.013
8.5 | 0.15 | [5 empty] | 0.78 | 0.033 | 8.5 | 0.61 | 0.21 | 0.37 | 0.51 | 0.6 |  | 0.63 |  | 8.5 | 0.48 |  | 0.48 | [5 empty] | 8.5 |  |  | 0.34 | [5 empty] | 8.5 |  |  | 0.19 | [5 empty] | 8.5
9 | [6 empty] | 0.67 | 0.029 | 9 | 0.55 | 0.18 | 0.22 | 0.46 | 0.52 |  | 0.63 | 0.025 | 9 | 0.49 | 0.33 | 0.48 | 0.39 | 0.47 |  | 0.41 |  | 9 | 0.34 |  | 0.35 |  | 0.34 |  |  |  | 9 |  |  | 0.2 |  | 0.22 |  |  |  | 9 |  |  | 0.12
9.5 | [6 empty] | 0.58 | 0.029 | 9.5 | 0.49 | 0.15 | 0.1 | 0.41 | 0.43 |  | 0.63 |  | 9.5 | 0.49 |  | 0.48 | [5 empty] | 9.5 |  |  | 0.35 | [5 empty] | 9.5 |  |  | 0.22 | [5 empty] | 9.5
10 | [6 empty] | 0.46 | 0.027 | 10 | 0.43 | 0.12 | 0.044 | 0.37 | 0.35 |  | 0.62 | 0.023 | 10 | 0.5 | 0.31 | 0.47 | 0.39 | 0.46 |  | 0.43 | 0.018 | 10 |  | 0.14 | 0.35 | 0.21 | 0.35 |  | 0.29 | 0.016 | 10 | 0.19 | 0.061 | 0.23 | 0.14 | 0.22 |  | 0.18 | 0.013 | 10 | 0.11 | 0.036 | 0.13 | 0.11 | 0.15 |  | 0.12 | 0.011
10.5 | [6 empty] | 0.36 | 0.023 | 10.5 | 0.36 | 0.09 |  | 0.28 | 0.27 |  | 0.6 |  | 10.5 | 0.5 |  | 0.45 | [5 empty] | 10.5 | 0.34 |  | 0.35 | [5 empty] | 10.5 |  |  | 0.23 | [5 empty] | 10.5
11 | [6 empty] | 0.18 | 0.018 | 11 | 0.28 | 0.06 |  | 0.18 | 0.2 |  | 0.58 | 0.02 | 11 | 0.49 | 0.29 | 0.4 | 0.39 | 0.44 |  | 0.44 |  | 11 |  |  | 0.35 |  | 0.34 |  |  |  | 11 |  |  | 0.23 |  | 0.23 |  |  |  | 11 |  |  | 0.14
11.5 | [6 empty] | 0.87 | 0.013 | 11.5 | 0.16 |  |  | 0.1 | 0.17 |  | 0.54 |  | 11.5 | 0.48 |  | 0.33 | [5 empty] | 11.5 |  |  | 0.35 | [5 empty] | 11.5 |  |  | 0.23 | [5 empty] | 11.5
12 | [7 empty] | 0.008 | 12 | 0.1 |  |  | 0.049 | 0.14 |  | 0.5 | 0.02 | 12 | 0.47 | 0.24 | 0.22 | 0.37 | 0.4 |  | 0.44 | 0.015 | 12 | 0.34 | 0.16 | 0.36 | 0.23 | 0.33 |  | 0.34 | 0.013 | 12 | 0.21 | 0.089 | 0.23 | 0.12 | 0.23 |  | 0.2 | 0.011 | 12 | 0.13 | 0.049 | 0.15 | 0.1 | 0.16 |  | 0.14 | 0.01
[9 empty] | 12.5 |  |  |  |  | 0.11 |  | 0.45 |  | 12.5 | 0.45 |  | 0.14 | [5 empty] | 12.5 |  |  | 0.36 | [5 empty] | 12.5 |  |  | 0.23 | [5 empty] | 12.5
[9 empty] | 13 |  |  |  |  | 0.09 |  | 0.4 | 0.019 | 13 | 0.42 | 0.15 | 0.085 | 0.28 | 0.32 |  | 0.45 |  | 13 |  |  | 0.35 |  | 0.31 |  |  |  | 13 |  |  | 0.24 |  | 0.23 |  |  |  | 13 |  |  | 0.16
[9 empty] | 13.5 |  |  |  |  | 0.07 |  | 0.35 |  | 13.5 | 0.39 |  | 0.048 | [5 empty] | 13.5 | 0.34 |  | 0.31 | [5 empty] | 13.5 |  |  | 0.25 | [5 empty] | 13.5
[9 empty] | 14 | [6 empty] | 0.3 | 0.016 | 14 | 0.35 | 0.06 |  | 0.13 | 0.19 |  | 0.45 | 0.015 | 14 |  | 0.17 | 0.26 | 0.25 | 0.29 |  | 0.34 | 0.011 | 14 | 0.23 | 0.1 | 0.26 | 0.11 | 0.22 |  | 0.22 | 0.009 | 14 | 0.15 | 0.061 | 0.17 | 0.09 | 0.17 |  | 0.15 | 0.009
[9 empty] | 14.5 | [6 empty] | 0.24 |  | 14.5 | 0.31 | [7 empty] | 14.5 |  |  | 0.18 | [5 empty] | 14.5 |  |  | 0.25 | [5 empty] | 14.5
[9 empty] | 15 | [6 empty] | 0.19 | 0.01 | 15 | 0.26 | 0.03 |  | 0.046 | 0.12 |  | 0.44 |  | 15 | 0.34 |  | 0.12 |  | 0.24 |  |  |  | 15 |  |  | 0.23 |  | 0.21 |  |  |  | 15 |  |  | 0.17
[9 empty] | 15.5 | [6 empty] | 0.16 |  | 15.5 | 0.2 | [7 empty] | 15.5 |  |  | 0.088 | [5 empty] | 15.5 |  |  | 0.19 | [5 empty] | 15.5
[9 empty] | 16 | [6 empty] | 0.13 | 0.004 | 16 | 0.16 |  |  |  | 0.093 |  | 0.42 | 0.01 | 16 |  | 0.156 | 0.06 | 0.22 | 0.17 |  | 0.35 | 0.011 | 16 | 0.22 | 0.1 | 0.16 | 0.11 | 0.19 |  | 0.23 | 0.009 | 16 | 0.16 | 0.66 | 0.16 | 0.082 | 0.16 |  | 0.17 | 0.008
[9 empty] | 16.5 | [6 empty] | 0.1 |  | 16.5 | 0.14 | [7 empty] | 16.5 | 0.33 |  | 0.041 | [5 empty] | 16.5 |  |  | 0.11 | [5 empty] | 16.5
[9 empty] | 17 | [6 empty] | 0.084 |  | 17 | 0.11 |  |  |  | 0.073 |  | 0.39 |  | 17 |  |  | 0.027 | 0.16 | 0.11 |  |  |  | 17 |  |  | 0.08 |  | 0.18 |  |  |  | 17 |  |  | 0.18
[9 empty] | 17.5 | [6 empty] | 0.068 |  | 17.5 | 0.08 | [7 empty] | 17.5 | [8 empty] | 17.5 |  |  | 0.058 | [5 empty] | 17.5
[18 empty] | 18 | 0.069 |  |  |  | 0.055 |  | 0.34 | 0.003 | 18 | 0.29 | 0.1 |  | 0.089 | 0.08 |  | 0.34 | 0.009 | 18 | 0.21 | 0.1 | 0.043 | 0.11 | 0.16 |  | 0.23 | 0.007 | 18 | 0.16 | 0.065 | 0.2 | 0.08 | 0.15 |  | 0.18 | 0.007
[18 empty] | 18.5 | 0.054 | [7 empty] | 18.5 | [8 empty] | 18.5 |  |  | 0.032 | [5 empty] | 18.5
[18 empty] | 19 |  |  |  |  | 0.038 |  | 0.28 |  | 19 |  | 0.06 |  | 0.037 | 0.06 |  |  |  | 19 |  |  |  |  | 0.13 |  |  |  | 19 |  |  | 0.2
[18 empty] | 19.5 | [8 empty] | 19.5 | 0.2 | [7 empty] | 19.5 | [8 empty] | 19.5
[18 empty] | 20 | [6 empty] | 0.19 |  | 20 |  | 0.03 |  |  | 0.051 |  | 0.3 | 0.006 | 20 | 0.2 | 0.09 |  | 0.085 | 0.1 |  | 0.24 | 0.004 | 20 | 0.16 | 0.064 | 0.16 | 0.08 | 0.13 |  | 0.17 | 0.007
[18 empty] | 20.5 | [8 empty] | 20.5 | [8 empty] | 20.5 | [8 empty] | 20.5
[18 empty] | 21 | [6 empty] | 0.14 |  | 21 | 0.13 |  |  |  | 0.046 |  |  |  | 21 |  |  |  | 0.056 | 0.077 |  |  |  | 21 |  |  | 0.1
[18 empty] | 21.5 | [8 empty] | 21.5 | [8 empty] | 21.5 | [8 empty] | 21.5
[18 empty] | 22 | [6 empty] | 0.1 |  | 22 |  |  |  |  | 0.039 |  | 0.24 | 0.0012 | 22 | 0.19 | 0.08 |  | 0.03 | 0.056 |  | 0.23 | 0.001 | 22 | 0.15 | 0.058 | 0.05 | 0.08 | 0.11 |  | 0.18 | 0.006
[18 empty] | 22.5 | [8 empty] | 22.5 | 0.084 | [7 empty] | 22.5 | [8 empty] | 22.5
[18 empty] | 23 | [6 empty] | 0.078 |  | 23 | [8 empty] | 23 |  |  |  |  | 0.044 |  |  |  | 23 |  |  | 0.033
[18 empty] | 23.5 | [8 empty] | 23.5 | [8 empty] | 23.5 | [8 empty] | 23.5
[18 empty] | 24 | [6 empty] | 0.057 |  | 24 | 0.047 | [5 empty] | 0.13 |  | 24 | 0.17 | 0.03 |  |  | 0.037 |  | 0.21 |  | 24 | 0.14 | 0.059 | 0.021 | 0.047 | 0.079 |  | 0.17
[18 empty] | 24.5 | [8 empty] | 24.5 | [8 empty] | 24.5 | [8 empty] | 24.5 | [7 empty] | 0.004
[18 empty] | 25 | [6 empty] | 0.043 |  | 25 | [8 empty] | 25 | [8 empty] | 25
[27 empty] | 25.5 | [8 empty] | 25.5 | [8 empty] | 25.5
[27 empty] | 26 | [6 empty] | 0.081 |  | 26 | 0.12 | 0.013 |  |  |  |  | 0.17 |  | 26 | 0.13 | 0.041 |  | 0.015 | 0.048 |  | 0.17
[27 empty] | 26.5 | [8 empty] | 26.5 | [8 empty] | 26.5
[27 empty] | 27 | [8 empty] | 27 | [8 empty] | 27 | [7 empty] | 0.002
[27 empty] | 27.5 | [8 empty] | 27.5 | [8 empty] | 27.5
[27 empty] | 28 | [6 empty] | 0.051 |  | 28 | 0.082 | [5 empty] | 0.13 |  | 28 | 0.12 | 0.018 |  |  | 0.031 |  | 0.15
[27 empty] | 28.5 | [8 empty] | 28.5 | [8 empty] | 28.5
[27 empty] | 29 | [8 empty] | 29 | [8 empty] | 29
[27 empty] | 29.5 | [8 empty] | 29.5 | [8 empty] | 29.5
[27 empty] | 30 | [6 empty] | 0.03 |  | 30 | 0.049 | [5 empty] | 0.086 |  | 30 | 0.1 | 0.008 |  |  |  |  | 0.12
[36 empty] | 30.5 | [8 empty] | 30.5
[36 empty] | 31 | [8 empty] | 31
[36 empty] | 31.5 | [8 empty] | 31.5
[36 empty] | 32 | [6 empty] | 0.057 |  | 32 | 0.081 | [5 empty] | 0.1
[36 empty] | 32.5 | [8 empty] | 32.5
[36 empty] | 33 | [8 empty] | 33
[36 empty] | 33.5 | [8 empty] | 33.5
[36 empty] | 34 | [6 empty] | 0.037 |  | 34 | 0.056 | [5 empty] | 0.075
34.5
35
35.5
[45 empty] | 36 | [6 empty] | 0.049
36.5
37
37.5
[45 empty] | 38 | [6 empty] | 0.036
38.5
39
39.5
[45 empty] | 40 | [6 empty] | 0.025

## Antares LARF
Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | Antares 660nm (CW) / EQ1 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Antares 904nm (CW) / EQ21 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Antares 808nm (CW) / EQ22 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | LARF |  |  |  | Peak (mW) |  | Power | 40 | OWNER | LARF |  |  |  | Peak (mW) |  | Power | 70mW | OWNER | LARF |  |  |  | Peak (mW) |  | Power
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660nm | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 808
0 | 0.28 | 0.16 | 0.1 | 0.06 | 0.038 | 0.021 |  |  | 0 | 0.079 | 0.072 | 0.046 | 0.035 | 0.021 | 0.024 |  |  | 0
0.5 | 0.43 | 0.19 | 0.13 | 0.07 |  |  | 1min |  | 0.5 | 0.1 | 0.098 |  |  |  |  | 1min |  | 0.5 | [6 empty] | 1min
1 | 0.64 | 0.26 | 0.16 | 0.08 | 0.051 | 0.025 | Maximum | 44.68 | 1 | 0.15 | 0.15 | 0.079 |  |  |  | Maximum | 74.65 | 1 | [6 empty] | Maximum
1.5 | 0.92 | 0.35 | 0.19 | 0.089 |  |  | Mean | 42 | 1.5 | 0.21 | 0.22 |  |  |  |  | Mean | 72.08 | 1.5 | [6 empty] | Mean
2 | 1.31 | 0.45 | 0.22 | 0.094 | 0.068 | 0.032 | Std Dev | 9.34 | 2 | 0.28 | 0.28 | 0.13 | 0.078 | 0.044 | 0.038 | Std Dev | 11.2 | 2 | [6 empty] | Std Dev
2.5 | 1.65 | 0.57 | 0.26 | 0.1 |  |  |  |  | 2.5 | 0.38 | 0.36 | [6 empty] | 2.5
3 | 2 | 0.7 | 0.31 | 0.12 | 0.081 | 0.04 |  |  | 3 | 0.44 | 0.39 | 0.19 |  |  |  | 1min |  | 3 | [6 empty] | 1min
3.5 | 2.26 | 0.86 | 0.35 | 0.15 |  |  | pi |  | 3.5 | 0.53 | 0.43 |  |  |  |  | Maximum | 75.76 | 3.5 | [6 empty] | Maximum
4 | 2.43 | 1.02 | 0.39 | 0.18 | 0.086 | 0.055 | 3.1416 |  | 4 | 0.67 | 0.46 | 0.21 | 0.15 | 0.081 | 0.064 | Mean | 74.13 | 4 | [6 empty] | Mean
4.5 | 2.38 | 1.16 | 0.44 | 0.21 |  |  | r |  | 4.5 | 0.83 | 0.5 |  |  |  |  | Std Dev | 6.78 | 4.5 | [6 empty] | Std Dev
5 | 2.15 | 1.3 | 0.49 | 0.24 | 0.1 | 0.06 | 1.775 |  | 5 | 0.97 | 0.52 | 0.23 | [5 empty] | 5
5.5 | 1.73 | 1.45 | 0.54 | 0.27 |  |  |  |  | 5.5 | 1.06 | 0.55 | [6 empty] | 5.5
6 | 1.31 | 1.49 | 0.59 | 0.31 | 0.11 | 0.062 | circle area |  | 6 | 1.09 | 0.57 | 0.29 | 0.21 | 0.11 | 0.084 |  |  | 6
6.5 | 0.86 | 1.53 | 0.65 | 0.34 |  |  | 9.898 |  | 6.5 | 1.07 | 0.59 | [6 empty] | 6.5
7 | 0.47 | 1.53 | 0.69 | 0.36 | 0.14 | 0.082 |  |  | 7 | 1.06 | 0.6 | 0.36 | [5 empty] | 7
7.5 | 0.22 | 1.48 | 0.74 | 0.39 |  |  |  |  | 7.5 | 0.96 | 0.61 | [6 empty] | 7.5
8 |  | 1.33 | 0.76 | 0.43 | 0.17 | 0.093 |  |  | 8 | 0.88 | 0.62 | 0.39 | 0.26 | 0.14 | 0.1 |  |  | 8
8.5 |  | 1.16 | 0.76 | 0.45 |  |  |  |  | 8.5 | 0.78 | 0.63 | [6 empty] | 8.5
9 |  | 0.93 | 0.78 | 0.47 | 0.21 | 0.11 |  |  | 9 | 0.67 | 0.63 | 0.41 | [5 empty] | 9
9.5 |  | 0.76 | 0.76 | 0.48 |  |  |  |  | 9.5 | 0.58 | 0.63 | [6 empty] | 9.5
10 |  | 0.56 | 0.74 | 0.5 | 0.24 | 0.13 |  |  | 10 | 0.46 | 0.62 | 0.43 | 0.29 | 0.18 | 0.12 |  |  | 10
10.5 |  | 0.42 | 0.72 | 0.5 |  |  |  |  | 10.5 | 0.36 | 0.6 | [6 empty] | 10.5
11 |  | 0.29 | 0.65 | 0.49 | 0.27 | 0.15 |  |  | 11 | 0.18 | 0.58 | 0.44 | [5 empty] | 11
11.5 |  | 0.19 | 0.61 | 0.49 |  |  |  |  | 11.5 | 0.87 | 0.54 | [6 empty] | 11.5
12 |  |  | 0.54 | 0.47 | 0.3 | 0.17 |  |  | 12 |  | 0.5 | 0.44 | 0.34 | 0.2 | 0.14 |  |  | 12
12.5 |  |  | 0.45 | 0.43 |  |  |  |  | 12.5 |  | 0.45 | [6 empty] | 12.5
13 |  |  | 0.38 | 0.38 | 0.31 | 0.18 |  |  | 13 |  | 0.4 | 0.45 | [5 empty] | 13
13.5 |  |  | 0.32 | 0.34 |  |  |  |  | 13.5 |  | 0.35 | [6 empty] | 13.5
14 |  |  | 0.25 | 0.29 | 0.32 | 0.19 |  |  | 14 |  | 0.3 | 0.45 | 0.34 | 0.22 | 0.15 |  |  | 14
14.5 |  |  | 0.2 | 0.25 |  |  |  |  | 14.5 |  | 0.24 | [6 empty] | 14.5
15 |  |  | 0.15 | 0.23 | 0.33 | 0.22 |  |  | 15 |  | 0.19 | 0.44 | [5 empty] | 15
15.5 |  |  |  | 0.22 |  |  |  |  | 15.5 |  | 0.16 | [6 empty] | 15.5
16 |  |  |  | 0.19 | 0.31 | 0.21 |  |  | 16 |  | 0.13 | 0.42 | 0.35 | 0.23 | 0.17 |  |  | 16
16.5 |  |  |  | 0.18 |  |  |  |  | 16.5 |  | 0.1 | [6 empty] | 16.5
17 |  |  |  | 0.16 | 0.28 | 0.23 |  |  | 17 |  | 0.084 | 0.39 | [5 empty] | 17
17.5 |  |  |  | 0.13 |  |  |  |  | 17.5 |  | 0.068 | [6 empty] | 17.5
18 |  |  |  | 0.11 | 0.19 | 0.24 |  |  | 18 |  |  | 0.34 | 0.34 | 0.23 | 0.18 |  |  | 18
18.5 |  |  |  | 0.09 |  |  |  |  | 18.5 | [8 empty] | 18.5
19 |  |  |  | 0.07 | 0.15 | 0.23 |  |  | 19 |  |  | 0.28 | [5 empty] | 19
19.5 | [8 empty] | 19.5 | [8 empty] | 19.5
20 |  |  |  |  | 0.13 | 0.22 |  |  | 20 |  |  | 0.19 | 0.3 | 0.24 | 0.17 |  |  | 20
20.5 | [8 empty] | 20.5 | [8 empty] | 20.5
21 |  |  |  |  | 0.13 | 0.15 |  |  | 21 |  |  | 0.14 | [5 empty] | 21
21.5 | [8 empty] | 21.5 | [8 empty] | 21.5
22 |  |  |  |  | 0.11 | 0.12 |  |  | 22 |  |  | 0.1 | 0.24 | 0.23 | 0.18 |  |  | 22
22.5 | [8 empty] | 22.5 | [8 empty] | 22.5
23 |  |  |  |  | 0.08 | 0.08 |  |  | 23 |  |  | 0.078 | [5 empty] | 23
23.5 | [8 empty] | 23.5 | [8 empty] | 23.5
24 | [5 empty] | 0.091 |  |  | 24 |  |  | 0.057 | 0.13 | 0.21 | 0.17 |  |  | 24
24.5 | [8 empty] | 24.5 | [8 empty] | 24.5
25 | [5 empty] | 0.09 |  |  | 25 |  |  | 0.043 | [5 empty] | 25
25.5 | [8 empty] | 25.5 | [8 empty] | 25.5
26 | [5 empty] | 0.09 |  |  | 26 |  |  |  | 0.081 | 0.17 | 0.17 |  |  | 26
26.5 | [8 empty] | 26.5 | [8 empty] | 26.5
27 | [5 empty] | 0.07 |  |  | 27 | [8 empty] | 27
27.5 | [8 empty] | 27.5 | [8 empty] | 27.5
28 | [5 empty] | 0.05 |  |  | 28 |  |  |  | 0.051 | 0.13 | 0.15 |  |  | 28
[9 empty] | 28.5 | [8 empty] | 28.5
[9 empty] | 29 | [8 empty] | 29
[9 empty] | 29.5 | [8 empty] | 29.5
[9 empty] | 30 |  |  |  | 0.03 | 0.086 | 0.12 |  |  | 30
[9 empty] | 30.5 | [8 empty] | 30.5
[9 empty] | 31 | [8 empty] | 31
[9 empty] | 31.5 | [8 empty] | 31.5
[9 empty] | 32 |  |  |  |  | 0.057 | 0.1 |  |  | 32
[9 empty] | 32.5 | [8 empty] | 32.5
[9 empty] | 33 | [8 empty] | 33
[9 empty] | 33.5 | [8 empty] | 33.5
[9 empty] | 34 |  |  |  |  | 0.037 | 0.075 |  |  | 34
[9 empty] | 34.5 | [8 empty] | 34.5
[9 empty] | 35 | [8 empty] | 35
[9 empty] | 35.5 | [8 empty] | 35.5
[9 empty] | 36 | [5 empty] | 0.049 |  |  | 36
[9 empty] | 36.5 | [8 empty] | 36.5
[9 empty] | 37 | [8 empty] | 37
[9 empty] | 37.5 | [8 empty] | 37.5
[9 empty] | 38 | [5 empty] | 0.036 |  |  | 38
[9 empty] | 38.5 | [8 empty] | 38.5
[9 empty] | 39 | [8 empty] | 39
[9 empty] | 39.5 | [8 empty] | 39.5
[9 empty] | 40 | [5 empty] | 0.025 |  |  | 40

## Laser Pulse LARF
Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | Laser Pulse 904nm (1) / EQ2 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Laser Pulse 450nm (2) / EQ3 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Laser Pulse 660nm (3) / EQ4 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Laser Pulse 830nm (4) / EQ5 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 830nm (5) / EQ6 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 830nm (6) / EQ7 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | LARF |  |  |  | Peak (mW) |  | Power | 70WpK | OWNER | LARF |  |  |  | Peak (mW) |  | Power | 70mW | OWNER | LARF |  |  |  | Peak (mW) |  | Power | 30Mw | OWNER | LARF |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | LARF |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | LARF |  |  |  | Peak (mW) |  | Power | 30mW
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 450 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830
0 | 0.1 | 0.094 | 0.043 | 0.04 | 0.018 | 0.01 | Pulse duration | 60ns | 0 | 0.27 | 0.1 | 0.05 | 0.02 | 0.02 | 0.01 |  |  | 0 | 0.19 | 0.15 | 0.04 | 0.028 | 0.028 | 0.023 |  |  | 0 | 0.25 | 0.1 | 0.043 | 0.03 | 0.019 | 0.015 |  |  |  | 0 | [9 empty] | 0 | 0.16 | 0.1 | 0.07 | 0.029 | 0.02 | 0.01
0.5 | 0.12 | 0.13 | 0.055 | [5 empty] | 0.5 | 0.34 | 0.11 |  |  |  |  | 1min |  | 0.5 | 0.34 | 0.28 | 0.07 |  |  |  | 1min |  | 0.5 | 0.55 | 0.16 |  |  |  |  | 1min | CW | 150HZ | 0.5 | [6 empty] | 1min | CW | 150HZ | 0.5 | 0.32 | 0.15 |  |  |  |  | 1min | CW | 150HZ
1 | 0.14 | 0.19 | 0.073 |  |  |  | 1min |  | 1 | 0.45 | 0.13 | 0.06 |  |  |  | Maximum | 65.78 | 1 | 0.74 | 0.44 | 0.11 | 0.05 |  |  | Maximum | 34.07 | 1 | 0.9 | 0.26 | 0.82 | 0.06 |  |  | Maximum | 32.6 | 15.77 | 1 | [6 empty] | Maximum | 885,8u | - | 1 | 0.57 | 0.2 | 0.1 |  |  |  | Maximum | 41.83 | 20.33
1.5 | 0.17 | 0.25 | 0.093 | 0.077 |  |  | Maximum | 43.14 | 1.5 | 0.57 | 0.16 |  |  |  |  | Mean | 62.37 | 1.5 | 1.23 | 0.59 | 0.16 |  | 0.068 |  | Mean | 33.06 | 1.5 | 1.3 | 0.36 |  |  |  |  | Mean | 31.45 | 15 | 1.5 | [6 empty] | Mean | 614,1u | - | 1.5 | 0.83 | 0.25 |  |  |  |  | Mean | 40.28 | 19.53
2 | 0.24 | 0.35 | 0.11 |  | 0.033 | 0.017 | Mean | 41.93 | 2 | 0.76 | 0.22 | 0.09 | 0.03 | 0.039 | 0.019 | Std Dev | 6.63 | 2 | 1.86 | 0.74 | 0.23 | 0.1 |  | 0.06 | Std Dev | 2.27 | 2 | 1.73 | 0.51 | 0.15 | 0.09 | 0.04 | 0.032 | Std Dev | 3.36 | 1.82 | 2 | [6 empty] | Std Dev | 185u | - | 2 | 1.05 | 0.28 | 0.14 | 0.06 | 0.04 | 0.018 | Std Dev | 4.09 | 1.93
2.5 | 0.004 | 0.41 | 0.14 |  |  |  | Std Dev | 4.76 | 2.5 | 0.99 | 0.29 | [6 empty] | 2.5 | 2.26 | 0.89 | 0.28 | [5 empty] | 2.5 | 2.24 | 0.64 | [7 empty] | 2.5 | [9 empty] | 2.5 | 1.24 | 0.32
3 | 0.01 | 0.48 | 0.17 | 0.13 |  |  |  |  | 3 | 1.27 | 0.39 | 0.14 | [5 empty] | 3 | 2.49 | 1 | 0.33 | 0.16 | 0.11 |  |  |  | 3 | 2.42 | 0.8 | 0.23 | 0.13 |  |  | 3min | CW | 150HZ | 3 | [6 empty] | 3min | CW | 150HZ | 3 | 1.39 | 0.34 | 0.13 |  |  |  | 3min | CW | 150HZ
3.5 | 0.06 | 0.54 | 0.2 |  |  |  | 3min |  | 3.5 | 1.49 | 0.47 | [6 empty] | 3.5 | 2.42 | 1.04 | 0.38 | [5 empty] | 3.5 | 2.03 | 0.92 |  |  |  |  | Maximum | 33.91 | 15.94 | 3.5 | [6 empty] | Maximum | 5.262 | - | 3.5 | 1.54 | 0.43 |  |  |  |  | Maximum | 42.17 | 21.53
4 | 0.3 | 0.63 | 0.25 |  | 0.069 | 0.033 | Maximum | 45.73 | 4 | 1.69 | 0.58 | 0.24 | 0.05 | 0.095 | 0.025 |  |  | 4 | 2.19 | 1.08 | 0.47 | 0.22 |  | 0.1 |  |  | 4 | 1.42 | 0.98 | 0.33 | 0.17 | 0.085 | 0.057 | Mean | 32.69 | 15.21 | 4 | [6 empty] | Mean | 4.513 | - | 4 | 1.65 | 0.55 | 0.26 | 0.08 | 0.058 | 0.041 | Mean | 40.56 | 20.5
4.5 | 0.74 | 0.68 | 0.29 | 0.2 |  |  | Mean | 44.46 | 4.5 | 1.86 | 0.66 | [6 empty] | 4.5 | 1.83 | 1.07 | 0.52 |  | 0.18 |  |  |  | 4.5 | 0.85 | 0.98 |  |  |  |  | Std Dev | 2.18 | 1.05 | 4.5 | [6 empty] | Std Dev | 0.53 | - | 4.5 | 1.51 | 0.66 |  |  |  |  | Std Dev | 2.28 | 1.35
5 | 1.05 | 0.73 | 0.36 |  |  |  | Std Dev | 2.38 | 5 | 1.85 | 0.7 | 0.31 | [5 empty] | 5 | 1.48 | 1 | 0.56 | 0.29 |  |  |  |  | 5 | 0.42 | 0.96 | 0.41 | 0.21 | [5 empty] | 5 | [9 empty] | 5 | 1.41 | 0.79 | 0.38
5.5 | 0.87 | 0.76 | 0.39 | [5 empty] | 5.5 | 1.73 | 0.73 | [6 empty] | 5.5 | 1.02 | 0.91 | 0.56 | [5 empty] | 5.5 | 0.19 | 0.86 | [7 empty] | 5.5 | [9 empty] | 5.5 | 1.25 | 0.85
6 | 0.68 | 0.78 | 0.41 | 0.25 | 0.11 | 0.059 |  |  | 6 | 1.55 | 0.83 | 0.38 | 0.13 | 0.12 | 0.04 |  |  | 6 | 0.63 | 0.82 | 0.57 | 0.33 | 0.22 | 0.15 |  |  | 6 |  | 0.7 | 0.46 | 0.24 | 0.13 | 0.082 |  |  |  | 6 | [9 empty] | 6 | 1.13 | 0.89 | 0.47 | 0.082 | 0.049 | 0.033
6.5 | 0.79 | 0.78 | 0.43 | [5 empty] | 6.5 | 1.29 | 0.91 | [6 empty] | 6.5 | 0.4 | 0.69 | 0.58 | [5 empty] | 6.5 |  | 0.61 | [7 empty] | 6.5 | [9 empty] | 6.5 | 0.94 | 0.87
7 | 1.17 | 0.76 | 0.44 | [5 empty] | 7 | 1.14 | 0.93 | 0.43 | [5 empty] | 7 | 0.22 | 0.54 | 0.56 | 0.36 |  |  |  |  | 7 |  | 0.49 | 0.47 | 0.26 | [5 empty] | 7 | [9 empty] | 7 | 0.82 | 0.86 | 0.46
7.5 | 1.01 | 0.72 | 0.45 | 0.28 |  |  |  |  | 7.5 | 0.86 | 0.87 | [6 empty] | 7.5 | 0.1 | 0.41 | 0.54 |  | 0.24 |  |  |  | 7.5 |  | 0.37 | [7 empty] | 7.5 | [9 empty] | 7.5 | 0.66 | 0.85
8 | 0.59 | 0.66 | 0.46 |  | 0.15 | 0.089 |  |  | 8 | 0.66 | 0.82 | 0.44 | 0.17 | 0.13 | 0.08 |  |  | 8 |  | 0.32 | 0.52 | 0.36 |  | 0.17 |  |  | 8 |  | 0.26 | 0.43 | 0.27 | 0.16 | 0.1 |  |  |  | 8 | [9 empty] | 8 | 0.48 | 0.81 | 0.48 | 0.2 | 0.088 | 0.045
8.5 | 0.15 | 0.61 | 0.48 | [5 empty] | 8.5 | 0.4 | 0.79 | [6 empty] | 8.5 |  | 0.23 | 0.48 | [5 empty] | 8.5 |  | 0.17 | [7 empty] | 8.5 | [9 empty] | 8.5 | 0.35 | 0.78
9 |  | 0.55 | 0.49 | 0.34 |  |  |  |  | 9 | 0.26 | 0.72 | 0.5 | [5 empty] | 9 |  | 0.17 | 0.42 | 0.35 | 0.24 |  |  |  | 9 |  | 0.11 | 0.34 | 0.27 | [5 empty] | 9 | [9 empty] | 9 | 0.25 | 0.73 | 0.47
9.5 |  | 0.49 | 0.49 | [5 empty] | 9.5 |  | 0.68 | [6 empty] | 9.5 |  |  | 0.37 | [5 empty] | 9.5 | [9 empty] | 9.5 | [9 empty] | 9.5 | 0.17 | 0.68
10 |  | 0.43 | 0.5 |  | 0.19 | 0.11 |  |  | 10 |  | 0.68 | 0.48 | 0.23 | 0.19 | 0.07 |  |  | 10 |  |  | 0.3 | 0.34 |  | 0.18 |  |  | 10 |  |  | 0.26 | 0.24 | 0.18 | 0.12 |  |  |  | 10 | [9 empty] | 10 |  | 0.61 | 0.45 | 0.28 | 0.17 | 0.064
10.5 |  | 0.36 | 0.5 | 0.34 |  |  |  |  | 10.5 |  | 0.73 | [6 empty] | 10.5 |  |  | 0.25 |  | 0.24 |  |  |  | 10.5 | [9 empty] | 10.5 | [9 empty] | 10.5 |  | 0.57
11 |  | 0.28 | 0.49 | [5 empty] | 11 |  | 0.64 | 0.42 | [5 empty] | 11 |  |  | 0.21 | 0.28 |  |  |  |  | 11 |  |  | 0.18 | 0.21 | [5 empty] | 11 | [9 empty] | 11 |  | 0.51 | 0.42
11.5 |  | 0.16 | 0.48 | [5 empty] | 11.5 |  | 0.65 | [6 empty] | 11.5 |  |  | 0.17 | [5 empty] | 11.5 | [9 empty] | 11.5 | [9 empty] | 11.5 |  | 0.44
12 |  | 0.1 | 0.47 | 0.34 | 0.21 | 0.13 |  |  | 12 |  | 0.44 | 0.42 | 0.26 | 0.21 | 0.1 |  |  | 12 |  |  | 0.13 | 0.22 | 0.21 | 0.17 |  |  | 12 |  |  | 0.11 | 0.16 | 0.17 | 0.13 |  |  |  | 12 | [9 empty] | 12 |  | 0.38 | 0.39 | 0.29 | 0.189 | 0.12
12.5 |  |  | 0.45 | [5 empty] | 12.5 |  | 0.32 | [6 empty] | 12.5 |  |  | 0.1 | [5 empty] | 12.5 | [9 empty] | 12.5 | [9 empty] | 12.5 |  | 0.34
13 |  |  | 0.42 | [5 empty] | 13 |  | 0.22 | 0.35 | [5 empty] | 13 |  |  | 0.08 | 0.17 |  |  |  |  | 13 |  |  | 0.06 | 0.12 | [5 empty] | 13 | [9 empty] | 13 |  | 0.29 | 0.35
13.5 |  |  | 0.39 | 0.34 |  |  |  |  | 13.5 |  | 0.14 | [6 empty] | 13.5 |  |  | 0.06 |  | 0.16 |  |  |  | 13.5 | [9 empty] | 13.5 | [9 empty] | 13.5 |  | 0.26
14 |  |  | 0.35 |  | 0.23 | 0.15 |  |  | 14 |  |  | 0.33 | 0.31 | 0.2 | 0.13 |  |  | 14 |  |  |  | 0.13 |  | 0.15 |  |  | 14 |  |  |  | 0.09 | 0.14 | 0.12 |  |  |  | 14 | [9 empty] | 14 |  | 0.21 | 0.31 | 0.3 | 0.2 | 0.13
14.5 |  |  | 0.31 | [5 empty] | 14.5 | [8 empty] | 14.5 | [8 empty] | 14.5 | [9 empty] | 14.5 | [9 empty] | 14.5 |  | 0.18
15 |  |  | 0.26 | 0.34 |  |  |  |  | 15 |  |  | 0.37 | [5 empty] | 15 |  |  |  | 0.09 | 0.12 |  |  |  | 15 |  |  |  | 0.06 | [5 empty] | 15 | [9 empty] | 15 |  | 0.15 | 0.26
15.5 |  |  | 0.2 | [5 empty] | 15.5 | [8 empty] | 15.5 | [8 empty] | 15.5 | [9 empty] | 15.5 | [9 empty] | 15.5
16 |  |  | 0.16 |  | 0.22 | 0.16 |  |  | 16 |  |  | 0.28 | 0.26 | 0.19 | 0.14 |  |  | 16 |  |  |  | 0.06 |  | 0.11 |  |  | 16 |  |  |  | 0.04 | 0.09 | 0.1 |  |  |  | 16 | [9 empty] | 16 |  |  | 0.21 | 0.27 | 0.196 | 0.14
16.5 |  |  | 0.14 | 0.33 |  |  |  |  | 16.5 | [8 empty] | 16.5 |  |  |  |  | 0.07 |  |  |  | 16.5 | [9 empty] | 16.5 | [9 empty] | 16.5
17 |  |  | 0.11 | [5 empty] | 17 |  |  | 0.12 | [5 empty] | 17 |  |  |  | 0.04 |  |  |  |  | 17 | [9 empty] | 17 | [9 empty] | 17 |  |  | 0.18
17.5 |  |  | 0.08 | [5 empty] | 17.5 | [8 empty] | 17.5 | [8 empty] | 17.5 | [9 empty] | 17.5 | [9 empty] | 17.5
18 |  |  | 0.069 | 0.29 | 0.21 | 0.16 |  |  | 18 |  |  | 0.07 | 0.25 | 0.19 | 0.15 |  |  | 18 |  |  |  | 0.024 | 0.05 | 0.07 |  |  | 18 |  |  |  |  | 0.05 | 0.078 |  |  |  | 18 | [9 empty] | 18 |  |  | 0.14 | 0.24 | 0.189 | 0.14
18.5 |  |  | 0.054 | [5 empty] | 18.5 | [8 empty] | 18.5 | [8 empty] | 18.5 | [9 empty] | 18.5 | [9 empty] | 18.5
19 | [8 empty] | 19 |  |  | 0.04 | [5 empty] | 19 | [8 empty] | 19 | [9 empty] | 19 | [9 empty] | 19 |  |  | 0.11
19.5 |  |  |  | 0.2 |  |  |  |  | 19.5 | [8 empty] | 19.5 |  |  |  |  | 0.03 |  |  |  | 19.5 | [9 empty] | 19.5 | [9 empty] | 19.5
20 |  |  |  |  | 0.2 | 0.16 |  |  | 20 |  |  |  | 0.25 | 0.16 | 0.16 |  |  | 20 | [5 empty] | 0.04 |  |  | 20 |  |  |  |  | 0.03 | 0.05 |  |  |  | 20 | [9 empty] | 20 |  |  | 0.09 | 0.2 | 0.177 | 0.138
20.5 | [8 empty] | 20.5 | [8 empty] | 20.5 | [8 empty] | 20.5 | [9 empty] | 20.5 | [9 empty] | 20.5
21 |  |  |  | 0.13 |  |  |  |  | 21 | [8 empty] | 21 | [8 empty] | 21 | [9 empty] | 21 | [9 empty] | 21 |  |  | 0.07
21.5 | [8 empty] | 21.5 | [8 empty] | 21.5 | [8 empty] | 21.5 | [9 empty] | 21.5 | [9 empty] | 21.5
22 |  |  |  |  | 0.19 | 0.15 |  |  | 22 |  |  |  | 0.13 | 0.14 | 0.12 |  |  | 22 | [5 empty] | 0.02 |  |  | 22 |  |  |  |  | 0.014 | 0.035 |  |  |  | 22 | [9 empty] | 22 |  |  |  | 0.15 | 0.157 | 0.13
22.5 |  |  |  | 0.084 |  |  |  |  | 22.5 | [8 empty] | 22.5 | [8 empty] | 22.5 | [9 empty] | 22.5 | [9 empty] | 22.5
23 | [8 empty] | 23 | [8 empty] | 23 | [8 empty] | 23 | [9 empty] | 23 | [9 empty] | 23
23.5 | [8 empty] | 23.5 | [8 empty] | 23.5 | [8 empty] | 23.5 | [9 empty] | 23.5 | [9 empty] | 23.5
24 |  |  |  | 0.047 | 0.17 | 0.14 |  |  | 24 |  |  |  | 0.057 | 0.07 | 0.11 |  |  | 24 | [8 empty] | 24 | [5 empty] | 0.019 |  |  |  | 24 | [9 empty] | 24 |  |  |  | 0.12 | 0.14 | 0.12
24.5 | [8 empty] | 24.5 | [8 empty] | 24.5 | [18 empty] | 24.5 | [9 empty] | 24.5
25 | [8 empty] | 25 | [8 empty] | 25 | [18 empty] | 25 | [9 empty] | 25
25.5 | [8 empty] | 25.5 | [8 empty] | 25.5 | [18 empty] | 25.5 | [9 empty] | 25.5
26 |  |  |  |  | 0.12 | 0.13 |  |  | 26 |  |  |  | 0.038 | 0.04 | 0.14 |  |  | 26 | [18 empty] | 26 | [9 empty] | 26 |  |  |  | 0.08 | 0.114 | 0.115
26.5 | [8 empty] | 26.5 | [8 empty] | 26.5 | [18 empty] | 26.5 | [9 empty] | 26.5
27 | [8 empty] | 27 | [8 empty] | 27 | [18 empty] | 27 | [9 empty] | 27
27.5 | [8 empty] | 27.5 | [8 empty] | 27.5 | [18 empty] | 27.5 | [9 empty] | 27.5
28 |  |  |  |  | 0.082 | 0.12 |  |  | 28 |  |  |  | 0.012 | 0.02 | 0.09 |  |  | 28 | [18 empty] | 28 | [9 empty] | 28 |  |  |  | 0.06 | 0.091 | 0.105
28.5 | [8 empty] | 28.5 | [8 empty] | 28.5 | [18 empty] | 28.5 | [9 empty] | 28.5
29 | [8 empty] | 29 | [8 empty] | 29 | [18 empty] | 29 | [9 empty] | 29
29.5 | [8 empty] | 29.5 | [8 empty] | 29.5 | [18 empty] | 29.5 | [9 empty] | 29.5
30 |  |  |  |  | 0.049 | 0.1 |  |  | 30 | [5 empty] | 0.05 |  |  | 30 | [18 empty] | 30 | [9 empty] | 30 |  |  |  | 0.03 | 0.074 | 0.09
30.5 | [8 empty] | 30.5 | [8 empty] | 30.5 | [18 empty] | 30.5 | [9 empty] | 30.5
31 | [8 empty] | 31 | [8 empty] | 31 | [18 empty] | 31 | [9 empty] | 31
31.5 | [8 empty] | 31.5 | [8 empty] | 31.5 | [18 empty] | 31.5 | [9 empty] | 31.5
32 | [5 empty] | 0.081 |  |  | 32 | [5 empty] | 0.02 |  |  | 32 | [18 empty] | 32 | [9 empty] | 32 |  |  |  |  | 0.055 | 0.07
32.5 | [8 empty] | 32.5 | [8 empty] | 32.5 | [18 empty] | 32.5 | [9 empty] | 32.5
33 | [8 empty] | 33 | [8 empty] | 33 | [18 empty] | 33 | [9 empty] | 33
33.5 | [8 empty] | 33.5 | [8 empty] | 33.5 | [18 empty] | 33.5 | [9 empty] | 33.5
34 | [5 empty] | 0.056 |  |  | 34 | [5 empty] | 0.02 |  |  | 34 | [18 empty] | 34 | [9 empty] | 34 |  |  |  |  | 0.041 | 0.06
34.5
35
35.5
[47 empty] | 36 |  |  |  |  | 0.029 | 0.05
36.5
37
37.5
[47 empty] | 38 | [5 empty] | 0.04
38.5
39
39.5
[47 empty] | 40 | [5 empty] | 0.03
40.5
41
41.5
[47 empty] | 42 | [5 empty] | 0.02
42.5
43
43.5
[47 empty] | 44 | [5 empty] | 0.018

## Laser Pulse CER
Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | Laser Pulse 830nm (1) / EQ8 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 830nm (2) / EQ9 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 830nm (3) / EQ10 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 660nm (1) / EQ11 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 660nm (2) / EQ12 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 904nm (3) / EQ13 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | CER |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | CER |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | CER |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | CER |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | CER |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | CER |  |  |  | Peak (mW) |  | Power | 70WpK
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904
0 | 0.215 | 0.15 | 0.048 | 0.03 | 0.032 | 0.018 |  |  |  | 0 | [9 empty] | 0 | 0.13 | 0.08 | 0.03 | 0.02 | 0.017 | 0.012 |  |  |  | 0 | 0.1 | 0.028 | 0.02 | 0.006 | 0.009 | 0.004 |  |  |  | 0 | 0.07 | 0.03 | 0.02 | 0.008 | 0.008 | 0.008 |  |  |  | 0 | 0.85 | 0.06 | 0.045 | 0.015 | 0.01 | 0.01 | Pulse duration | 60ns
0.5 | 0.47 | 0.31 | 0.093 | [6 empty] | 0.5 | [9 empty] | 0.5 | 0.3 | 0.15 | [7 empty] | 0.5 | 0.15 | 0.04 | [7 empty] | 0.5 | 0.1 | 0.06 | [7 empty] | 0.5 | 0.36 | 0.09
1 | 0.85 | 0.51 | 0.18 | 0.07 |  |  | 1min | CW | 150Hz | 1 | [6 empty] | 1min | CW | 150Hz | 1 | 0.56 | 0.22 | 0.068 | 0.04 |  |  | 1min | CW | 150Hz | 1 | 0.2 | 0.057 | 0.036 |  |  |  | 1min | CW | 150Hz | 1 | 0.2 | 0.08 |  |  |  |  | 1min | CW | 150Hz | 1 | 0.57 | 0.18 | 0.063 |  |  |  | 1min | PW
1.5 | 1.32 | 0.69 | 0.28 |  | 0.1 | 0.067 | Maximum | 34.07 | 19 | 1.5 | [6 empty] | Maximum | 34.53 | 16.21 | 1.5 | 0.79 | 0.25 |  |  |  |  | Maximum | 30.9 | 12.94 | 1.5 | 0.28 | 0.075 |  |  |  |  | Maximum | 19.14 | 9.301 | 1.5 | 0.34 | 0.11 |  |  |  |  | Maximum | 17.72 | 8.604 | 1.5 | 0.72 | 0.35 |  |  |  |  | Maximum | 53.16
2 | 1.81 | 0.85 | 0.39 | 0.18 |  |  | Mean | 33.63 | 18.47 | 2 | [6 empty] | Mean | 33.77 | 15.95 | 2 | 0.93 | 0.27 | 0.1 | 0.06 | 0.039 | 0.047 | Mean | 30.4 | 12.73 | 2 | 0.38 | 0.095 | 0.046 | 0.011 | 0.013 |  | Mean | 18.08 | 8.941 | 2 | 0.4 | 0.14 | 0.044 |  |  |  | Mean | 16.93 | 8.22 | 2 | 0.93 | 0.41 | 0.087 | 0.03 | 0.018 | 0.015 | Mean | 52.08
2.5 | 2.09 | 0.97 | 0.43 |  |  |  | Std Dev | 826u | 485u | 2.5 | [6 empty] | Std Dev | 2.83 | 1.03 | 2.5 | 0.98 | 0.3 |  |  |  |  | Std Dev | 1,77m | 678u | 2.5 | 0.5 | 0.11 |  |  |  |  | Std Dev | 835u | 338u | 2.5 | 0.48 | 0.18 |  |  |  |  | Std Dev | 484u | 776u | 2.5 | 1.1 | 0.46 |  |  |  |  | Std Dev | 3,57m
3 | 1.95 | 1.01 | 0.48 | 3 | 0.21 | 0.136 |  |  |  | 3 | [9 empty] | 3 | 1.12 | 0.34 | 0.12 | 0.079 | [5 empty] | 3 | 0.66 | 0.14 | 0.055 |  |  | 0.01 |  |  |  | 3 | 0.48 | 0.22 |  | 0.022 | 0.016 | 0.012 |  |  |  | 3 | 1.15 | 0.49 | 0.1
3.5 | 1.56 | 0.98 | 0.51 |  |  |  | 3min | CW | 150Hz | 3.5 | [6 empty] | 3min | CW | 150Hz | 3.5 | 1.26 | 0.45 |  |  |  |  | 3min | CW | 150Hz | 3.5 | 0.81 | 0.17 | [7 empty] | 3.5 | 0.42 | 0.25 | [7 empty] | 3.5 | 1.09 | 0.5 |  |  |  |  | 3min | PW
4 | 1 | 0.86 | 0.54 | 0.34 |  |  | Maximum | 35.34 | 15.25 | 4 | [6 empty] | Maximum | 33.91 | 16.58 | 4 | 1.18 | 0.55 | 0.19 | 0.13 | 0.12 | 0.11 | Maximum | 32.97 | 16.13 | 4 | 0.94 | 0.21 | 0.075 | 0.022 | 0.019 |  |  |  |  | 4 | 0.3 | 0.34 | 0.083 | [6 empty] | 4 | 0.99 | 0.51 | 0.12 | 0.049 | 0.029 | 0.019 | Maximum | 53.74
4.5 | 0.55 | 0.65 | 0.55 |  | 0.2 | 0.146 | Mean | 34.94 | 15.07 | 4.5 | [6 empty] | Mean | 33.53 | 16.36 | 4.5 | 0.93 | 0.6 |  |  |  |  | Mean | 32.43 | 15.96 | 4.5 | 0.92 | 0.22 | [7 empty] | 4.5 | 0.22 | 0.38 | [7 empty] | 4.5 | 0.83 | 0.52 |  |  |  |  | Mean | 52.92
5 | 0.207 | 0.45 | 0.55 | 0.32 |  |  | Std Dev | 736u | 380u | 5 | [6 empty] | Std Dev | 1,04m | 802u | 5 | 0.7 | 0.58 | 0.3 | 0.2 |  |  | Std Dev | 977u | 630u | 5 | 0.74 | 0.24 | 0.086 | [6 empty] | 5 | 0.14 | 0.37 | [7 empty] | 5 | 0.63 | 0.51 | 0.16 |  |  |  | Std Dev | 2.38
5.5 |  | 0.28 | 0.52 | [6 empty] | 5.5 | [9 empty] | 5.5 | 0.43 | 0.5 | [7 empty] | 5.5 | 0.53 | 0.26 | [7 empty] | 5.5 | 0.1 | 0.34 | [7 empty] | 5.5 | 0.43 | 0.49
6 |  | 0.153 | 0.45 | 0.3 | 0.19 | 0.139 |  |  |  | 6 | [9 empty] | 6 | 0.28 | 0.4 | 0.34 | 0.22 | 0.15 | 0.099 |  |  |  | 6 | 0.33 | 0.29 | 0.1 | 0.036 | 0.026 | 0.016 |  |  |  | 6 | 0.06 | 0.31 | 0.11 | 0.043 | 0.025 | 0.018 |  |  |  | 6 | 0.3 | 0.44 | 0.21 | 0.065 | 0.037 | 0.027
6.5 |  |  | 0.35 | [6 empty] | 6.5 | [9 empty] | 6.5 | 0.15 | 0.31 | [7 empty] | 6.5 | 0.22 | 0.32 | [7 empty] | 6.5 |  | 0.27 | [7 empty] | 6.5 | 0.21 | 0.39
7 |  |  | 0.22 | 0.28 | [5 empty] | 7 | [9 empty] | 7 |  | 0.2 | 0.29 | 0.19 | [5 empty] | 7 | 0.11 | 0.28 | 0.12 | [6 empty] | 7 |  | 0.23 | [7 empty] | 7 | 0.11 | 0.34 | 0.27
7.5 |  |  | 0.13 |  | 0.19 | 0.135 |  |  |  | 7.5 | [9 empty] | 7.5 |  | 0.13 | [7 empty] | 7.5 |  | 0.25 | [7 empty] | 7.5 |  | 0.2 | [7 empty] | 7.5 | 0.46 | 0.29
8 |  |  | 0.072 | 0.26 | [5 empty] | 8 | [9 empty] | 8 |  | 0.09 | 0.2 | 0.17 | 0.11 | 0.082 |  |  |  | 8 |  | 0.19 | 0.13 | 0.041 | 0.028 |  |  |  |  | 8 |  | 0.17 | 0.16 | [6 empty] | 8 |  | 0.25 | 0.34 | 0.089 | 0.046 | 0.03
8.5 |  |  | 0.035 | [6 empty] | 8.5 | [9 empty] | 8.5 | [9 empty] | 8.5 |  | 0.16 | [7 empty] | 8.5 |  | 0.14 | [7 empty] | 8.5 |  | 0.21
9 |  |  |  | 0.16 | 0.17 | 0.14 |  |  |  | 9 | [9 empty] | 9 |  |  | 0.11 | 0.13 | [5 empty] | 9 |  | 0.15 | 0.15 |  |  | 0.017 |  |  |  | 9 |  | 0.12 |  | 0.064 | 0.035 | 0.019 |  |  |  | 9 |  | 0.18 | 0.33
9.5 | [9 empty] | 9.5 | [9 empty] | 9.5 | [9 empty] | 9.5 |  | 0.13 | [7 empty] | 9.5 |  | 0.1 | [7 empty] | 9.5 |  | 0.15
10 |  |  |  | 0.056 | [5 empty] | 10 | [9 empty] | 10 |  |  | 0.05 | 0.08 | 0.08 | 0.069 |  |  |  | 10 |  | 0.12 | 0.11 | 0.071 | 0.05 |  |  |  |  | 10 |  | 0.09 | 0.1 | [6 empty] | 10 |  | 0.12 | 0.31 | 0.14 | 0.061 | 0.036
10.5 |  |  |  |  | 0.068 | 0.145 |  |  |  | 10.5 | [9 empty] | 10.5 | [9 empty] | 10.5 |  | 0.11 | [7 empty] | 10.5 |  | 0.09 | [7 empty] | 10.5 |  | 0.09
11 |  |  |  | 0.021 | [5 empty] | 11 | [9 empty] | 11 |  |  | 0.018 | 0.047 | [5 empty] | 11 |  | 0.11 | 0.087 | [6 empty] | 11 |  | 0.08 | [7 empty] | 11 |  | 0.06 | 0.29
11.5 | [9 empty] | 11.5 | [9 empty] | 11.5 | [9 empty] | 11.5 |  | 0.097 | [7 empty] | 11.5 |  | 0.07 | [7 empty] | 11.5
12 |  |  |  |  | 0.017 | 0.077 |  |  |  | 12 | [9 empty] | 12 |  |  |  | 0.02 | 0.03 | 0.038 |  |  |  | 12 |  | 0.087 | 0.075 | 0.074 | 0.051 | 0.037 |  |  |  | 12 |  | 0.05 | 0.06 | 0.053 | 0.028 | 0.024 |  |  |  | 12 |  |  | 0.24 | 0.16 | 0.089 | 0.049
12.5 | [9 empty] | 12.5 | [9 empty] | 12.5 | [9 empty] | 12.5 |  | 0.079 | [7 empty] | 12.5 |  | 0.05 | [7 empty] | 12.5
13 | [9 empty] | 13 | [9 empty] | 13 | [9 empty] | 13 |  | 0.071 | 0.069 | [6 empty] | 13 |  | 0.03 | [7 empty] | 13 |  |  | 0.15
13.5 | [5 empty] | 0.019 |  |  |  | 13.5 | [9 empty] | 13.5 | [9 empty] | 13.5 |  | 0.056 | [7 empty] | 13.5 | [9 empty] | 13.5
[10 empty] | 14 | [9 empty] | 14 |  |  |  |  | 0.007 | 0.012 |  |  |  | 14 |  | 0.045 | 0.073 | 0.079 | 0.048 |  |  |  |  | 14 |  |  | 0.06 | [6 empty] | 14 |  |  | 0.06 | 0.17 | 0.1 | 0.061
[10 empty] | 14.5 | [9 empty] | 14.5 | [9 empty] | 14.5 |  | 0.035 | [7 empty] | 14.5 | [9 empty] | 14.5
[10 empty] | 15 | [9 empty] | 15 | [9 empty] | 15 |  |  | 0.065 |  |  | 0.034 |  |  |  | 15 |  |  |  | 0.04 | 0.028 | 0.021 |  |  |  | 15 |  |  | 0.03
[10 empty] | 15.5 | [9 empty] | 15.5 | [9 empty] | 15.5 | [9 empty] | 15.5 | [9 empty] | 15.5
[10 empty] | 16 | [9 empty] | 16 | [5 empty] | 0.003 |  |  |  | 16 |  |  | 0.062 | 0.071 | 0.055 |  |  |  |  | 16 |  |  | 0.05 | [6 empty] | 16 |  |  |  | 0.156 | 0.1 | 0.66
[10 empty] | 16.5 | [19 empty] | 16.5 | [9 empty] | 16.5 | [9 empty] | 16.5
[10 empty] | 17 | [19 empty] | 17 |  |  | 0.048 | [6 empty] | 17 | [9 empty] | 17
[10 empty] | 17.5 | [19 empty] | 17.5 | [9 empty] | 17.5 | [9 empty] | 17.5
[10 empty] | 18 | [19 empty] | 18 |  |  | 0.033 | 0.043 | 0.032 | 0.035 |  |  |  | 18 |  |  | 0.003 | 0.037 | 0.033 | 0.017 |  |  |  | 18 |  |  |  | 0.1 | 0.1 | 0.065
[10 empty] | 18.5 | [19 empty] | 18.5 | [9 empty] | 18.5 | [9 empty] | 18.5
[10 empty] | 19 | [19 empty] | 19 |  |  | 0.025 | [6 empty] | 19 | [9 empty] | 19 |  |  |  | 0.06
[10 empty] | 19.5 | [19 empty] | 19.5 | [9 empty] | 19.5 | [9 empty] | 19.5
[10 empty] | 20 | [19 empty] | 20 |  |  |  | 0.04 | 0.029 |  |  |  |  | 20 |  |  | 0.002 | [6 empty] | 20 |  |  |  | 0.03 | 0.09 | 0.064
[10 empty] | 20.5 | [19 empty] | 20.5 | [9 empty] | 20.5 | [9 empty] | 20.5
[10 empty] | 21 | [19 empty] | 21 | [5 empty] | 0.029 |  |  |  | 21 |  |  |  | 0.035 | 0.028 | 0.017 |  |  |  | 21
[10 empty] | 21.5 | [19 empty] | 21.5 | [9 empty] | 21.5 | [9 empty] | 21.5
[10 empty] | 22 | [19 empty] | 22 |  |  |  | 0.041 | 0.025 |  |  |  |  | 22 | [9 empty] | 22 |  |  |  |  | 0.08 | 0.058
[10 empty] | 22.5 | [19 empty] | 22.5 | [9 empty] | 22.5 | [9 empty] | 22.5
[10 empty] | 23 | [19 empty] | 23 | [9 empty] | 23 | [9 empty] | 23
[10 empty] | 23.5 | [19 empty] | 23.5 | [9 empty] | 23.5 | [9 empty] | 23.5
[10 empty] | 24 | [19 empty] | 24 |  |  |  | 0.039 | 0.025 | 0.018 |  |  |  | 24 |  |  |  | 0.02 | 0.025 | 0.015 |  |  |  | 24 |  |  |  |  | 0.03 | 0.059
[10 empty] | 24.5 | [19 empty] | 24.5 | [9 empty] | 24.5 | [9 empty] | 24.5
[10 empty] | 25 | [19 empty] | 25 | [9 empty] | 25 | [9 empty] | 25
[10 empty] | 25.5 | [19 empty] | 25.5 | [9 empty] | 25.5 | [9 empty] | 25.5
[10 empty] | 26 | [19 empty] | 26 |  |  |  | 0.021 | 0.025 |  |  |  |  | 26 | [9 empty] | 26 |  |  |  |  | 0.013 | 0.041
[10 empty] | 26.5 | [19 empty] | 26.5 | [9 empty] | 26.5 | [9 empty] | 26.5
[10 empty] | 27 | [19 empty] | 27 | [5 empty] | 0.015 |  |  |  | 27 |  |  |  | 0.012 | 0.013 | 0.009 |  |  |  | 27
[10 empty] | 27.5 | [19 empty] | 27.5 | [9 empty] | 27.5 | [9 empty] | 27.5
[10 empty] | 28 | [19 empty] | 28 |  |  |  | 0.01 | 0.012 |  |  |  |  | 28 | [9 empty] | 28 | [5 empty] | 0.018
[10 empty] | 28.5 | [19 empty] | 28.5 | [9 empty] | 28.5 | [9 empty] | 28.5
[10 empty] | 29 | [19 empty] | 29 | [9 empty] | 29 | [9 empty] | 29
[10 empty] | 29.5 | [19 empty] | 29.5 | [9 empty] | 29.5 | [9 empty] | 29.5
[10 empty] | 30 | [19 empty] | 30 |  |  |  | 0.01 | 0.01 | 0.017 |  |  |  | 30 |  |  |  | 0.006 | 0.01 |  |  |  |  | 30 | [5 empty] | 0.008
[10 empty] | 30.5 | [19 empty] | 30.5 | [9 empty] | 30.5
[10 empty] | 31 | [19 empty] | 31 | [9 empty] | 31
[10 empty] | 31.5 | [19 empty] | 31.5 | [9 empty] | 31.5
[10 empty] | 32 | [19 empty] | 32 |  |  |  |  | 0.005 |  |  |  |  | 32
[10 empty] | 32.5 | [19 empty] | 32.5 | [9 empty] | 32.5
[10 empty] | 33 | [19 empty] | 33 | [5 empty] | 0.014 |  |  |  | 33 |  |  |  |  | 0.008
[10 empty] | 33.5 | [19 empty] | 33.5
[10 empty] | 34 | [19 empty] | 34
34.5
35
35.5
[30 empty] | 36 | [5 empty] | 0.008

## Laser Pulse Multi
Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | Laser Pulse 660nm (1) / EQ14 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 660nm (2) / EQ15 |  |  |  | Mean (mW) |  | Spot beam size |  |  | DEVICE 1 | Laser Pulse 904nm (1) / EQ16 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Laser Pulse 904nm (2) / EQ17 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Laser Pulse 904nm (3) / EQ18 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | Laser Pulse 830nm (1) / EQ19 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | Multidisciplinar |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | Multidisciplinar |  |  |  | Peak (mW) |  | Power | 30mW |  | OWNER | Multidisciplinar |  |  |  | Peak (mW) |  | Power |  | OWNER | Multidisciplinar |  |  |  | Peak (mW) |  | Power |  | OWNER | Multidisciplinar |  |  |  | Peak (mW) |  | Power |  | OWNER | Multidisciplinar |  |  |  | Peak (mW) |  | Power
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660 |  | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 830
0 | 0.23 | 0.09 | 0.04 | 0.028 | 0.023 | 0.021 |  |  |  | 0 | 0.12 | 0.06 | 0.03 | 0.03 | 0.02 | 0.023 |  |  |  | 0 | 0.063 | 0.085 | 0.048 | 0.03 | 0.032 | 0.024 |  |  | 0 | 0.038 | 0.05 | 0.056 | 0.02 | 0.017 | 0.013 |  |  | 0 | 0.06 | 0.072 | 0.03 | 0.036 | 0.026 | 0.017 |  |  | 0 | 0.046 | 0.035 | 0.023 | 0.016 | 0.015 | 0.012
0.5 | 0.4 | 0.14 | 0.05 | 0.036 | [5 empty] | 0.5 | 0.19 | 0.08 | 0.048 | [6 empty] | 0.5 | 0.14 | 0.13 | 0.079 | 0.047 | 0.049 |  |  |  | 0.5 | 0.11 | 0.095 | [6 empty] | 0.5 | 0.19 | 0.12 | [6 empty] | 0.5 | 0.086 | 0.05
1 | 0.73 | 0.21 | 0.07 | 0.046 | 0.037 | 0.03 | 1min | CW | 150Hz | 1 | 0.26 | 0.1 | 0.069 | 0.047 |  |  | 1min | CW | 150Hz | 1 | 0.26 | 0.21 | 0.12 | 0.071 | 0.07 | 0.051 | 1min | CW | 1 | 0.26 | 0.16 | 0.11 |  | 0.035 |  | 1min | CW | 1 | 0.38 | 0.2 | 0.08 | 0.076 | 0.044 |  | 1min | CW | 1 | 0.16 | 0.068 | 0.044 | 0.02 | 0.023 |  | 1min | CW | 150Hz
1.5 | 1.19 | 0.24 | 0.09 | 0.056 |  |  | Maximum | 35.12 | 19.39 | 1.5 | 0.28 | 0.13 | 0.11 |  |  |  | Maximum | 28.54 | 17.46 | 1.5 | 0.42 | 0.28 | 0.17 | 0.1 | 0.099 |  | Maximum | 51.07 | 1.5 | 0.41 | 0.23 |  |  |  |  | Maximum | 42.63 | 1.5 | 0.54 | 0.28 |  |  |  |  | Maximum | 44.35 | 1.5 | 0.24 | 0.09 |  |  |  |  | Maximum | 32.36 | 14.93
2 | 1.73 | 0.33 | 0.13 | 0.07 | 0.053 | 0.04 | Mean | 34.28 | 18.47 | 2 | 0.41 | 0.16 | 0.13 | 0.063 | 0.039 | 0.038 | Mean | 27.13 | 16.92 | 2 | 0.51 | 0.37 | 0.23 | 0.14 | 0.12 | 0.088 | Mean | 50.04 | 2 | 0.49 | 0.34 | 0.2 | 0.088 | 0.07 | 0.052 | Mean | 41.86 | 2 | 0.69 | 0.38 | 0.17 | 0.13 | 0.074 | 0.043 | Mean | 43.08 | 2 | 0.34 | 0.1 | 0.068 | 0.033 | 0.039 | 0.022 | Mean | 31.25 | 14.64
2.5 | 2.04 | 0.47 | 0.16 | 0.087 |  |  | Std Dev | 0.416 | 1.74 | 2.5 | 0.56 | 0.2 | 0.15 |  |  |  | Std Dev | 3.56 | 1.46 | 2.5 | 0.61 | 0.44 | 0.27 | 0.18 | 0.15 |  | Std Dev | 4.44 | 2.5 | 0.55 | 0.41 |  |  |  |  | Std Dev | 3.96 | 2.5 | 0.83 | 0.45 |  |  |  |  | Std Dev | 4.57 | 2.5 | 0.4 | 0.13 |  |  |  |  | Std Dev | 2.04 | 1.29
3 | 2.16 | 0.59 | 0.21 | 0.11 | 0.071 | 0.051 |  |  |  | 3 | 0.74 | 0.25 | 0.18 | 0.087 | [5 empty] | 3 | 0.71 | 0.51 | 0.34 | 0.2 | 0.16 | 0.12 |  |  | 3 | 0.57 | 0.46 | 0.27 | [5 empty] | 3 | 0.93 | 0.5 | 0.27 | 0.2 | 0.11 |  |  |  | 3 | 0.44 | 0.17 | 0.096 | 0.053 | 0.05
3.5 | 1.95 | 0.7 | 0.26 | 0.13 | [5 empty] | 3.5 | 0.91 | 0.29 | 0.22 | [6 empty] | 3.5 | 0.79 | 0.56 | 0.35 | 0.22 | 0.16 |  | 3min | CW | 3.5 | 0.59 | 0.49 |  |  |  |  | 3min | CW | 3.5 | 0.95 | 0.54 |  |  |  |  | 3min | CW | 3.5 | 0.41 | 0.19 |  |  |  |  | 3min | CW | 150Hz
4 | 1.57 | 0.79 | 0.29 | 0.15 | 0.099 | 0.066 |  |  |  | 4 | 1.09 | 0.29 | 0.24 | 0.11 | 0.065 | 0.055 |  |  |  | 4 | 0.83 | 0.6 | 0.37 | 0.22 | 0.16 | 0.13 | Maximum | 51.33 | 4 | 0.59 | 0.51 | 0.29 | 0.16 | 0.15 | 0.11 | Maximum | 43.38 | 4 | 0.79 | 0.58 | 0.38 | 0.26 | 0.16 | 0.093 | Maximum | 44.64 | 4 | 0.35 | 0.23 | 0.12 | 0.076 | 0.058 | 0.041 | Maximum | 34.34 | 14.93
4.5 | 0.98 | 0.84 | 0.29 | 0.18 | [5 empty] | 4.5 | 1.29 | 0.36 | 0.29 | [6 empty] | 4.5 | 0.73 | 0.63 | 0.38 | 0.22 | 0.15 |  | Mean | 50.81 | 4.5 | 0.57 | 0.52 |  |  |  |  | Mean | 42.59 | 4.5 | 0.64 | 0.61 |  |  |  |  | Mean | 43.93 | 4.5 | 0.29 | 0.25 |  |  |  |  | Mean | 32.73 | 14.77
5 | 0.44 | 0.86 | 0.32 | 0.2 | 0.13 | 0.085 |  |  |  | 5 | 1.46 | 0.4 | 0.25 | 0.13 | [5 empty] | 5 | 0.57 | 0.64 | 0.4 | 0.22 | 0.15 | 0.11 | Std Dev | 2.66 | 5 | 0.53 | 0.53 | 0.32 |  |  |  | Std Dev | 2.43 | 5 | 0.53 | 0.64 | 0.42 | 0.29 | 0.2 |  | Std Dev | 2.71 | 5 | 0.2 | 0.28 | 0.16 | 0.08 | 0.069 |  | Std Dev | 1.64 | 0.759
5.5 | 0.23 | 0.8 | 0.37 | 0.22 | [5 empty] | 5.5 | 1.57 | 0.48 | 0.27 | [6 empty] | 5.5 | 0.38 | 0.65 | 0.42 | 0.24 | 0.15 |  |  |  | 5.5 | 0.5 | 0.55 | [6 empty] | 5.5 | 0.44 | 0.66 | [6 empty] | 5.5 | 0.13 | 0.3
6 |  | 0.71 | 0.39 | 0.24 | 0.15 | 0.1 |  |  |  | 6 | 1.62 | 0.54 | 0.3 | 0.17 | 0.094 | 0.074 |  |  |  | 6 | 0.12 | 0.65 | 0.44 | 0.25 | 0.16 | 0.11 |  |  | 6 | 0.42 | 0.57 | 0.33 | 0.17 | 0.17 | 0.12 |  |  | 6 | 0.36 | 0.68 | 0.44 | 0.33 | 0.22 | 0.15 |  |  | 6 | 0.083 | 0.29 | 0.19 | 0.1 | 0.069 | 0.043
6.5 |  | 0.59 | 0.42 | 0.26 | [5 empty] | 6.5 | 1.58 | 0.57 | 0.32 | [6 empty] | 6.5 | 0.039 | 0.65 | 0.45 | 0.26 | 0.16 |  |  |  | 6.5 | 0.31 | 0.58 | [6 empty] | 6.5 | 0.21 | 0.69 | [6 empty] | 6.5 | 0.061 | 0.3
7 |  | 0.46 | 0.43 | 0.28 | 0.17 | 0.11 |  |  |  | 7 | 1.59 | 0.63 | 0.33 | 0.19 | [5 empty] | 7 |  | 0.63 | 0.47 | 0.27 | 0.17 | 0.11 |  |  | 7 | 0.17 | 0.58 | 0.36 | [5 empty] | 7 | 0.13 | 0.69 | 0.46 | 0.33 | 0.22 |  |  |  | 7 | 0.034 | 0.3 | 0.2 | 0.13 | 0.07
7.5 |  | 0.34 | 0.41 | 0.29 | [5 empty] | 7.5 | 1.42 | 0.67 | 0.33 | [6 empty] | 7.5 |  | 0.58 | 0.48 | 0.28 | 0.17 |  |  |  | 7.5 | 0.1 | 0.56 | [6 empty] | 7.5 | 0.06 | 0.68 | [6 empty] | 7.5 |  | 0.29
8 |  | 0.29 | 0.4 | 0.24 | 0.19 | 0.13 |  |  |  | 8 | 1.31 | 0.71 | 0.35 | 0.22 | 0.12 | 0.092 |  |  |  | 8 |  | 0.49 | 0.48 | 0.33 | 0.18 | 0.12 |  |  | 8 | 0.04 | 0.55 | 0.38 | 0.17 | 0.15 | 0.11 |  |  | 8 |  | 0.65 | 0.47 | 0.34 | 0.22 | 0.16 |  |  | 8 |  | 0.26 | 0.19 | 0.15 | 0.07 | 0.038
8.5 |  | 0.18 | 0.37 | 0.24 | [5 empty] | 8.5 | 1.05 | 0.7 | 0.36 | [6 empty] | 8.5 |  | 0.37 | 0.48 | 0.34 | 0.19 |  |  |  | 8.5 |  | 0.51 | [6 empty] | 8.5 |  | 0.6 | [6 empty] | 8.5 |  | 0.23
9 |  | 0.1 | 0.33 | 0.29 | 0.21 | 0.14 |  |  |  | 9 | 0.77 | 0.69 | 0.38 | 0.25 | [5 empty] | 9 |  | 0.22 | 0.48 | 0.35 | 0.2 | 0.12 |  |  | 9 |  | 0.46 | 0.39 | [5 empty] | 9 |  | 0.52 | 0.47 | 0.34 | 0.22 |  |  |  | 9 |  | 0.2 | 0.16 | 0.15 | 0.075
9.5 |  | 0.05 | 0.28 | 0.29 | [5 empty] | 9.5 | 0.51 | 0.68 | 0.39 | [6 empty] | 9.5 |  | 0.1 | 0.48 | 0.35 | 0.22 |  |  |  | 9.5 |  | 0.41 | [6 empty] | 9.5 |  | 0.43 | [6 empty] | 9.5 |  | 0.16
10 |  |  | 0.24 | 0.29 | 0.21 | 0.15 |  |  |  | 10 | 0.3 | 0.65 | 0.38 | 0.24 | 0.15 | 0.11 |  |  |  | 10 |  | 0.044 | 0.47 | 0.35 | 0.23 | 0.13 |  |  | 10 |  | 0.37 | 0.39 | 0.21 | 0.14 | 0.11 |  |  | 10 |  | 0.35 | 0.46 | 0.35 | 0.22 | 0.15 |  |  | 10 |  | 0.11 | 0.19 | 0.15 | 0.075 | 0.031
10.5 |  |  | 0.26 | 0.28 | [5 empty] | 10.5 | 0.2 | 0.65 | 0.36 | [6 empty] | 10.5 |  |  | 0.45 | 0.35 | 0.23 |  |  |  | 10.5 |  | 0.28 | [6 empty] | 10.5 |  | 0.27 | [6 empty] | 10.5 |  | 0.081
11 |  |  | 0.2 | 0.27 | 0.21 | 0.15 |  |  |  | 11 | 0.11 | 0.61 | 0.34 | 0.25 | [5 empty] | 11 |  |  | 0.4 | 0.35 | 0.23 | 0.14 |  |  | 11 |  | 0.18 | 0.39 | [5 empty] | 11 |  | 0.2 | 0.44 | 0.34 | 0.23 |  |  |  | 11 |  | 0.057 | 0.14 | 0.14 | 0.081
11.5 |  |  | 0.16 | 0.24 | [5 empty] | 11.5 |  | 0.57 | 0.31 | [6 empty] | 11.5 |  |  | 0.33 | 0.35 | 0.23 |  |  |  | 11.5 |  | 0.1 | [6 empty] | 11.5 |  | 0.17 | [6 empty] | 11.5 |  | 0.033
12 |  |  | 0.12 | 0.22 | 0.2 | 0.16 |  |  |  | 12 |  | 0.5 | 0.3 | 0.27 | 0.18 | 0.13 |  |  |  | 12 |  |  | 0.22 | 0.36 | 0.23 | 0.15 |  |  | 12 |  | 0.049 | 0.37 | 0.23 | 0.12 | 0.1 |  |  | 12 |  | 0.14 | 0.4 | 0.33 | 0.23 | 0.16 |  |  | 12 |  |  | 0.11 | 0.13 | 0.079 | 0.045
12.5 |  |  | 0.09 | 0.19 | [5 empty] | 12.5 |  | 0.42 | 0.28 | [6 empty] | 12.5 |  |  | 0.14 | 0.36 | 0.23 |  |  |  | 12.5 | [8 empty] | 12.5 |  | 0.11 | [6 empty] | 12.5
13 |  |  | 0.06 | 0.17 | 0.18 | 0.15 |  |  |  | 13 |  | 0.34 | 0.25 | 0.29 | [5 empty] | 13 |  |  | 0.085 | 0.35 | 0.24 | 0.16 |  |  | 13 |  |  | 0.28 | [5 empty] | 13 |  | 0.09 | 0.32 | 0.31 | 0.23 |  |  |  | 13 |  |  | 0.082 | 0.11 | 0.1
13.5 |  |  | 0.03 | 0.14 | [5 empty] | 13.5 |  | 0.27 | 0.24 | [6 empty] | 13.5 |  |  | 0.048 | 0.31 | 0.25 |  |  |  | 13.5 | [8 empty] | 13.5 |  | 0.07 | [6 empty] | 13.5
14 |  |  |  | 0.12 | 0.15 | 0.14 |  |  |  | 14 |  | 0.27 | 0.27 | 0.25 | 0.18 | 0.13 |  |  |  | 14 |  |  |  | 0.26 | 0.26 | 0.17 |  |  | 14 |  |  | 0.13 | 0.25 | 0.11 | 0.09 |  |  | 14 |  |  | 0.19 | 0.29 | 0.22 | 0.17 |  |  | 14 |  |  | 0.051 | 0.1 | 0.1 | 0.044
14.5 |  |  |  | 0.1 | [5 empty] | 14.5 |  | 0.2 | 0.25 | [6 empty] | 14.5 |  |  |  | 0.18 | 0.25 |  |  |  | 14.5 | [8 empty] | 14.5 | [8 empty] | 14.5
15 |  |  |  | 0.079 | 0.12 | 0.12 |  |  |  | 15 |  | 0.15 | 0.22 | 0.23 | [5 empty] | 15 |  |  |  | 0.12 | 0.23 | 0.17 |  |  | 15 |  |  | 0.046 | [5 empty] | 15 |  |  | 0.12 | 0.24 | 0.21 |  |  |  | 15 |  |  | 0.034 | 0.08 | 0.096
15.5 |  |  |  | 0.063 | [5 empty] | 15.5 |  | 0.1 | 0.2 | [6 empty] | 15.5 |  |  |  | 0.088 | 0.19 |  |  |  | 15.5 | [8 empty] | 15.5 | [8 empty] | 15.5
16 |  |  |  | 0.049 | 0.095 | 0.1 |  |  |  | 16 |  | 0.6 | 0.18 | 0.21 | 0.19 | 0.14 |  |  |  | 16 |  |  |  | 0.06 | 0.16 | 0.16 |  |  | 16 |  |  |  | 0.22 | 0.11 | 0.082 |  |  | 16 |  |  | 0.093 | 0.17 | 0.19 | 0.16 |  |  | 16 |  |  | 0.017 | 0.055 | 0.07 | 0.054
16.5 |  |  |  | 0.034 | [5 empty] | 16.5 |  |  | 0.16 | [6 empty] | 16.5 |  |  |  | 0.041 | 0.11 |  |  |  | 16.5 | [8 empty] | 16.5 | [8 empty] | 16.5
17 |  |  |  | 0.023 | 0.071 | 0.085 |  |  |  | 17 |  |  | 0.13 | 0.19 | [5 empty] | 17 |  |  |  | 0.027 | 0.08 | 0.18 |  |  | 17 |  |  |  | 0.16 |  |  |  |  | 17 |  |  | 0.073 | 0.11 | 0.18 |  |  |  | 17 |  |  |  | 0.047 | 0.044
17.5 | [9 empty] | 17.5 |  |  | 0.11 | [6 empty] | 17.5 |  |  |  |  | 0.058 |  |  |  | 17.5 | [8 empty] | 17.5 | [8 empty] | 17.5
18 |  |  |  |  | 0.047 | 0.073 |  |  |  | 18 |  |  | 0.099 | 0.17 | 0.16 | 0.13 |  |  |  | 18 |  |  |  |  | 0.043 | 0.2 |  |  | 18 |  |  |  | 0.089 | 0.11 | 0.08 |  |  | 18 |  |  | 0.055 | 0.08 | 0.16 | 0.15 |  |  | 18 |  |  |  | 0.035 | 0.026 | 0.056
18.5 | [9 empty] | 18.5 |  |  | 0.084 | [6 empty] | 18.5 |  |  |  |  | 0.032 |  |  |  | 18.5 | [8 empty] | 18.5 | [8 empty] | 18.5
19 |  |  |  |  | 0.031 | 0.055 |  |  |  | 19 |  |  | 0.072 | 0.14 | [5 empty] | 19 | [5 empty] | 0.2 |  |  | 19 |  |  |  | 0.037 |  |  |  |  | 19 |  |  | 0.038 | 0.06 | 0.13 |  |  |  | 19 |  |  |  | 0.017 | 0.027
19.5 | [9 empty] | 19.5 |  |  | 0.06 | [6 empty] | 19.5 | [8 empty] | 19.5 | [8 empty] | 19.5 | [8 empty] | 19.5
20 |  |  |  |  | 0.016 | 0.045 |  |  |  | 20 |  |  | 0.053 | 0.12 | 0.14 | 0.11 |  |  |  | 20 | [5 empty] | 0.16 |  |  | 20 |  |  |  |  | 0.085 | 0.08 |  |  | 20 |  |  |  | 0.051 | 0.1 | 0.13 |  |  | 20 |  |  |  |  | 0.018 | 0.036
20.5 | [9 empty] | 20.5 |  |  | 0.043 | [6 empty] | 20.5 | [8 empty] | 20.5 | [8 empty] | 20.5 | [8 empty] | 20.5
21 | [5 empty] | 0.029 |  |  |  | 21 |  |  | 0.03 | 0.1 | [5 empty] | 21 | [5 empty] | 0.1 |  |  | 21 |  |  |  |  | 0.056 |  |  |  | 21 |  |  |  | 0.046 | 0.077 |  |  |  | 21
21.5 | [9 empty] | 21.5 | [9 empty] | 21.5 | [8 empty] | 21.5 | [8 empty] | 21.5 | [8 empty] | 21.5
22 | [5 empty] | 0.017 |  |  |  | 22 |  |  |  | 0.083 | 0.11 | 0.1 |  |  |  | 22 | [5 empty] | 0.05 |  |  | 22 |  |  |  |  | 0.03 | 0.08 |  |  | 22 |  |  |  | 0.039 | 0.056 | 0.11 |  |  | 22 | [5 empty] | 0.024
[10 empty] | 22.5 | [9 empty] | 22.5 | [8 empty] | 22.5 | [8 empty] | 22.5 | [8 empty] | 22.5
[10 empty] | 23 |  |  |  | 0.061 | [5 empty] | 23 | [5 empty] | 0.033 |  |  | 23 | [8 empty] | 23 |  |  |  |  | 0.044 |  |  |  | 23
[10 empty] | 23.5 | [9 empty] | 23.5 | [8 empty] | 23.5 | [8 empty] | 23.5 | [8 empty] | 23.5
[10 empty] | 24 |  |  |  | 0.048 | 0.092 | 0.089 |  |  |  | 24 | [5 empty] | 0.021 |  |  | 24 | [5 empty] | 0.047 |  |  | 24 |  |  |  |  | 0.037 | 0.079 |  |  | 24 | [5 empty] | 0.013
[10 empty] | 24.5 | [18 empty] | 24.5 | [8 empty] | 24.5
[10 empty] | 25 |  |  |  | 0.037 | [14 empty] | 25 | [8 empty] | 25
[10 empty] | 25.5 | [18 empty] | 25.5 | [8 empty] | 25.5
[10 empty] | 26 |  |  |  | 0.028 | 0.069 | 0.073 | [12 empty] | 26 | [5 empty] | 0.015 |  |  | 26 | [5 empty] | 0.048
[10 empty] | 26.5 | [27 empty] | 26.5
[10 empty] | 27 | [27 empty] | 27
[10 empty] | 27.5 | [27 empty] | 27.5
[10 empty] | 28 |  |  |  |  | 0.04 | 0.054 | [21 empty] | 28 | [5 empty] | 0.031
28.5
29
29.5
[10 empty] | 30 |  |  |  |  | 0.027 | 0.038
30.5
31
31.5
[10 empty] | 32 | [5 empty] | 0.027

## HTM CSE
Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | HTM 904nm (1) / EQ20 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | CSE |  |  |  | Peak (mW) |  | Power
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904
0
0.5
1 | [6 empty] | 1min | CW
1.5 | [6 empty] | Maximum | 8.014
2 | [6 empty] | Mean | 7.639
2.5 | [6 empty] | Std Dev | 306u
3
3.5 | [6 empty] | 3min | CW
4 | [6 empty] | Maximum | 8.978
4.5 | [6 empty] | Mean | 8.177
5 | [6 empty] | Std Dev | 472u
5.5
6
6.5
7
7.5
8
8.5
9
9.5
10
10.5
11
11.5
12
12.5
13
13.5
14
14.5
15
15.5
16
16.5
17
17.5
18
18.5
19
19.5
20
20.5
21
21.5
22

## HTM Supere
Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | HTM 904nm (1) / EQ23 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | Supere |  |  |  | Peak (mW) |  | Power
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 904
0 | 0.005 | 0.004 | 0.004 | 0.0023 | 0.0025 | 0.0015
0.5 | 0.01
1 | 0.017 | 0.009 |  |  |  |  | 1min | 5KHz
1.5 | 0.026 | [5 empty] | Maximum | 11.61
2 | 0.037 | 0.019 | 0.016 | 0.007 | 0.008 | 0.004 | Mean | 11.43
2.5 | 0.048 | [5 empty] | Std Dev | 593u
3 | 0.059 | 0.033
3.5 | 0.064 | [5 empty] | 3min | 5KHz
4 | 0.059 | 0.043 | 0.032 | 0.021 | 0.017 | 0.011 | Maximum | 11.8
4.5 | 0.053 | [5 empty] | Mean | 11.68
5 | 0.049 | 0.038 |  |  |  |  | Std Dev | 82,6u
5.5 | 0.047
6 | 0.044 | 0.033 | 0.025 | 0.022 | 0.018 | 0.016
6.5 | 0.041
7 | 0.039 | 0.03
7.5 | 0.036
8 | 0.034 | 0.028 | 0.021 | 0.018 | 0.015 | 0.013
8.5 | 0.033
9 | 0.029 | 0.025
9.5 | 0.029
10 | 0.027 | 0.023 | 0.018 | 0.016 | 0.013 | 0.011
10.5 | 0.023
11 | 0.018 | 0.02
11.5 | 0.013
12 | 0.008 | 0.02 | 0.015 | 0.013 | 0.011 | 0.01
12.5
13 |  | 0.019
13.5
14 |  | 0.016 | 0.015 | 0.011 | 0.009 | 0.009
14.5
15 |  | 0.01
15.5
16 |  | 0.004 | 0.01 | 0.011 | 0.009 | 0.008
16.5
17
17.5
18 |  |  | 0.003 | 0.009 | 0.007 | 0.007
18.5
19
19.5
20 |  |  |  | 0.006 | 0.004 | 0.007
20.5
21
21.5
22 |  |  |  | 0.0012 | 0.001 | 0.006
22.5
23
23.5
24
24.5 | [5 empty] | 0.004
25
25.5
26
26.5
27 | [5 empty] | 0.002

## DMC Therapy XT First
Power and beam characterization | [6 empty] | Parameters (Manual) |  | Power and beam characterization | [6 empty] | Parameters (Manual)
DEVICE 1 | DMC Therapy XT 808nm (1) / EQ24 |  |  |  | Mean (mW) |  | Spot beam size |  | DEVICE 1 | DMC Therapy XT 660nm (1) / EQ25 |  |  |  | Mean (mW) |  | Spot beam size
OWNER | First |  |  |  | Peak (mW) |  | Power | 100 | OWNER | First |  |  |  | Peak (mW) |  | Power | 100
Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 808 | Distance between points (mm) | 0mm (mW) | 3mm (mW) | 6mm (mW) | 9mm (mW) | 12mm (mW) | 15mm (mW) | Wavelength | 660
0 | 0.258 | 0.114 | 0.133 | 0.07 | 0.046 | 0.024 |  |  | 0 | 0.25 | 0.159 | 0.13 | 0.16 | 0.06 | 0.069
0.5 | 0.56 | 0.281 | 0.158 | 0.1 | 0.08 |  |  |  | 0.5 | 0.35
1 | 1.13 | 0.33 | 0.191 | 0.15 | 0.1 |  | 1min |  | 1 | 1.34 | 0.42 |  |  |  |  | 1min
1.5 | 0.65 | 0.5 | 0.227 | 0.2 | 0.15 |  | Maximum | 99.83 | 1.5 | 1.8 | [5 empty] | Maximum | 103.4
2 | 2.09 | 0.66 | 0.46 | 0.22 | 0.18 | 0.12 | Mean | 95.52 | 2 | 2.5 | 0.74 | 0.28 | 0.28 | 0.09 | 0.08 | Mean | 100.6
2.5 | 2.33 | 0.88 | 0.56 | 0.24 |  |  | Std Dev | 13.7 | 2.5 | 2.94 | [5 empty] | Std Dev | 2.52
3 | 2.69 | 1.1 | 0.69 | 0.24 | 0.166 |  |  |  | 3 | 4.5 | 1.77
3.5 | 3.37 | 1.6 | 0.87 | 0.21 |  |  |  |  | 3.5 | 5.94
4 | 3.95 | 2.22 | 1.02 | 0.23 | 0.1 | 0.077 |  |  | 4 | 6.83 | 3.38 | 0.62 | 0.4 | 0.14 | 0.1
4.5 | 4.63 | 2.54 | 1.16 | 0.33 |  |  |  |  | 4.5 | 7.25
5 | 5.11 | 2.77 | 1.17 | 0.39 | 0.088 |  |  |  | 5 | 7.46 | 5.2
5.5 | 5.22 | 2.88 | 1.32 | 0.45 |  |  |  |  | 5.5 | 7.1
6 | 5.14 | 2.82 | 1.38 | 0.53 | 0.1 | 0.057 |  |  | 6 | 6.8 | 6.3 | 1.44 | 0.9 | 0.16 | 0.12
6.5 | 4.9 | 2.62 | 1.44 | 0.58 |  |  |  |  | 6.5 | 5.89
7 | 4.49 | 2.32 | 1.48 | 0.66 | 0.16 |  |  |  | 7 | 4.45 | 6.7
7.5 | 4.15 | 1.94 | 1.45 | 0.69 |  |  |  |  | 7.5 | 2.54
8 | 2.8 | 1.5 | 1.38 | 0.72 | 0.25 | 0.08 |  |  | 8 | 1.3 | 6.66 | 2.26 | 1.14 | 0.27 | 0.18
8.5 | 1.18 | 1.12 | 1.17 | 0.73 |  |  |  |  | 8.5 | 0.23
9 | 0.4 | 0.8 | 1.01 | 0.75 | 0.27 |  |  |  | 9 | 0.144 | 4.99
9.5 | 0.22 | 0.53 | 0.85 | 0.77 |  |  |  |  | 9.5 | 0.02
10 |  | 0.37 | 0.7 | 0.78 | 0.36 | 0.18 |  |  | 10 |  | 3.7 | 2.41 | 1.24 | 0.5 | 0.31
10.5 |  | 0.24 | 0.63 | 0.77 |  |  |  |  | 10.5
11 |  | 0.177 | 0.56 | 0.7 | 0.38 |  |  |  | 11 |  | 2.58
11.5 |  |  | 0.52 | 0.68 |  |  |  |  | 11.5
12 |  |  | 0.45 | 0.64 | 0.4 | 0.24 |  |  | 12 |  | 1.32 | 2.19 | 1.21 | 0.68 | 0.43
12.5 |  |  | 0.41 | 0.6 |  |  |  |  | 12.5
13 |  |  | 0.36 | 0.53 | 0.4 |  |  |  | 13 |  | 0.45
13.5 |  |  | 0.26 | 0.49 |  |  |  |  | 13.5
14 |  |  | 0.156 | 0.43 | 0.39 | 0.27 |  |  | 14 |  | 0.13 | 1.19 | 1.14 | 0.68 | 0.42
14.5 |  |  |  | 0.38 |  |  |  |  | 14.5
15 |  |  |  | 0.3 | 0.38 |  |  |  | 15
15.5 |  |  |  | 0.25 |  |  |  |  | 15.5
16 |  |  |  | 0.194 | 0.41 | 0.25 |  |  | 16 |  |  | 0.27 | 0.82 | 0.73 | 0.45
16.5 |  |  |  | 0.15 |  |  |  |  | 16.5
17 |  |  |  | 0.11 | 0.44 |  |  |  | 17
17.5 |  |  |  | 0.1 |  |  |  |  | 17.5
18 |  |  |  | 0.099 | 0.42 | 0.24 |  |  | 18 |  |  |  | 0.37 | 0.65 | 0.46
18.5 |  |  |  | 0.09 |  |  |  |  | 18.5
19 |  |  |  |  | 0.37 |  |  |  | 19
19.5 | [8 empty] | 19.5
20 |  |  |  |  | 0.31 | 0.26 |  |  | 20 |  |  |  | 0.11 | 0.46 | 0.46
20.5 | [8 empty] | 20.5
21 |  |  |  |  | 0.26 |  |  |  | 21
21.5 | [8 empty] | 21.5
22 |  |  |  |  | 0.24 | 0.2 |  |  | 22 |  |  |  |  | 0.26 | 0.42
22.5 | [8 empty] | 22.5
23 |  |  |  |  | 0.18 |  |  |  | 23
23.5 | [8 empty] | 23.5
24 |  |  |  |  | 0.11 | 0.18 |  |  | 24 |  |  |  |  | 0.08 | 0.41
24.5 | [8 empty] | 24.5
25 |  |  |  |  | 0.08 |  |  |  | 25
25.5 | [8 empty] | 25.5
26 | [5 empty] | 0.15 |  |  | 26 | [5 empty] | 0.36
26.5 | [8 empty] | 26.5
27 | [8 empty] | 27
27.5 | [8 empty] | 27.5
28 | [5 empty] | 0.066 |  |  | 28 | [5 empty] | 0.25
28.5 | [8 empty] | 28.5
29 | [8 empty] | 29
29.5 | [8 empty] | 29.5
30 | [5 empty] | 0.037 |  |  | 30 | [5 empty] | 0.13
30.5
31
31.5
[9 empty] | 32 | [5 empty] | 0.47