## data Average - TO BE USED
[21 empty] | u(Paroscientific) | [18 empty] | u(MBW) | [6 empty] | f dry sphere | 5223000000 | Hz |  |  | measured 3/2019 | 0.0035 | -0.008 | 3.34
[21 empty] | 55.2 | Pa | [17 empty] | 15 | Pa | T0 | 29 | df/dT | -82200 | Hz/K | kx(T0,p0) | 113.05 |  |  |  |  | 1.126 | 1 | 1.031 |  | Vperm |  |  |  | capillary diffusion | xDry | [27 empty] | a2sccm | a200sccm | BronkIN | [5 empty] | From invoice: the company sais that's the expanded uncertanty |  |  |  |  | 0.0023 | 0.00023 |  |  | phi
|  |  |  | The value of a highlighted cell is nominal. | [37 empty] | p0 | 108000 | df/dp | -13.2 | Hz/Pa | lit dx/df | -4.3289e-08 | 1/Hz |  |  | MFC offset | 0 | 0 | -6 | sccm | 0.024 | sccm |  |  | 5.1e-11 | 0.43 |  | argon |  |  |  |  | zero adjust | 0 | 0 | 10 | ppm | [15 empty] | full scale | 2 | 200 | 2000 | sccm |  |  |  |  | To divided by 2 |  | Length |  | 5.07 | 0.5 | 0.06 | m | 9.5e-12 |  | 950
|  |  |  | The input variables are in red font. | [18 empty] | The input variables are in red font. | [18 empty] | f0 | 10570155 | Hz |  |  | fit dx/df | -4.28e-08 | 1/Hz |  |  | MFC slope adjust | 1.087 | 0.952 | 1 | [5 empty] | mol/s | ppm | [6 empty] | slope adjust | 1 | 1 | 1 | [27 empty] | Vperm |  | 0.023 | 0.0022682 | 0.00027219 | sccm | mol/(s m Pa) |  | 1e-14 mol/(s m Pa)
parameters |  |  |  | saturator temperature and pressure | [18 empty] | MBW temperature and pressure | [18 empty] | microwave hygrometer | [7 empty] | gas flow rates |  |  | 0 | [6 empty] | humidity values | [16 empty] | u(difference) |  |  | difference | ratio |  |  | u(ratio) |  |  | mass-flow |  |  | ratio | [6 empty] | PE: full range (plaine echelle) |  | Length-saturation | [6 empty] | humidity values
[7 empty] | H2O vapor | Silicone | [17 empty] | H2O vapor | [23 empty] | Alicat 2 | Alicat 200 | Bronkhorst 2000 | Alicat 2 adj | Alicat 200 adj | Bronk 2000 adj |  | Ar perm |  |  | nominal | diffusion | xDry |  | from MBW | from Tmirror |  |  |  |  | adjusted | adjusted | adjusted | [14 empty] | Alicat2 | Alicat200 | Bronkhorst 2000 | [9 empty] | Tube length | V perm | V wet | Ar perm. |  |  | Saturation | nominal | diffusion | xDry | from MBW
general |  |  | h | C | K |  | Pa | Pa |  |  |  | m3/mol | [5 empty] | 1e-4 | 1e-4 | 1e-4 | 1e-4 | 1e-4 | C | K |  | Pa | Pa |  |  |  | m3/mol | [5 empty] | 1e-4 | 1e-4 | 1e-4 | 1e-4 | 1e-4 | C | K | Pa | Hz | Hz | Hz |  | ppm | sccm | sccm | sccm | sccm | sccm | sccm | sccm | sccm | mol/s | mol/s | ppm | ppm | ppm | ppm | ppm | ppm | ppm |  | ppm | ppm | ppm | ppm | ppm |  | ppm | ppm | ppm | ppm | ppm | ppm | ppm | [6 empty] | accuracy | accuracy | accuracy | [9 empty] | m | sccm | sccm | sccm | mol/s | m3/s | length | ppm | ppm | ppm | ppm
Rgas | 8.3145 | J / (mol K) | t | T | T | theta | pVlit | p | Zw0 | Z | ln(Z/Zw0) | Baw | virial | Poynting | Raoult | fw | x | u(fw) | u(pVlit)/pVlit | slope*u(T)/T | u(p)/p | u(x)/x | Tmirror | Tmirror | thetaMirror | pVlit | p | Zw0 | Z | ln(Z/Zw0) | Baw | virial | Poynting | Raoult | fw | x | u(fw) | u(pVlit)/pVlit | slope*u(Tmirror)/Tmirror | u(p)/p | u(x)/x | Tspheres | Tspheres | p | f1-f2 | f1-f2 - f0 | adjust for T,p | x CNAM | xPitre | V0 | V0 | Vdry | V0 small | V0 big | Vdry | V0 | V0-Vperm | V0-Vperm | Vdry | x gen | x gen | x gen | u(gen) | x MBW | x MBW | u(MBW) | u(MBW)/u(gen) | CNAM Pitre | CNAM | MBW | CNAM | CNAM Pitre |  | CNAM-MBW | CNAM - gen | gen-MBW | CNAM - MBW | CNAM - gen | MBW - gen | CNAM Pitre / CNAM | CNAM/MBW | CNAM / gen | gen/MBW | CNAM / MBW | CNAM / gen | 1 + u(MBW/gen) | sccm | sccm | sccm | u(a2)/a2 | u(a200)/a200 | u(BR)/BR |  | u_difference+ | u_ratio | u_ratio+ |  |  | L_tube |  | V0+Vh2o | Vwet-perm | Vwet-Vperm | Vwet-Vperm | m | x gen | x gen | x gen | x MBW
Patm | 101325 | Pa |  | 28.96 | 302.11 | 0.53313 | 3999.8471 | 108880 | 0.9981 | 0.99936 | 0.001252 | -2.5177e-05 | 0.0021038 | 0.00075219 | 2.5113e-05 | 1.0041 | 0.036888 | 5.7566 | 2 | 12.1401 | 5.0698 | 14.4991 | -44.555 | 228.595 | 0.83685 | 7.5902 | 108590.5 | 0.99997 | 0.998 | -0.0019773 | -4.8149e-05 | 0.0055125 | 0.0011223 | 0.00015397 | 1.0048 | 7.0235e-05 | 7.75 | 10 | 117.7613 | 1.3813 | 118.447 | 29 |  | 108000 | 10570011.5 | -143.5 | -143.5 | 6.1418e-06 |  | 0.1 | 0 | 49.9 | 0.1087 | 0 | 43.9 | 0.1087 | 0.0847 | 5.77e-08 | 2.9906e-05 | 73.7502 | 73.7502 | 74.1802 | 0.10755 | 70.25 | 70.2349 | 0.83209 | 7.7365 |  | 6.1418 | 70.25 | 6.1418 |  |  | -64.0931 | -68.0384 | 3.9302 |  |  | 0.83901 | 0 | 0.087428 | 0.082796 | 1.0559 |  |  | 1.0119 | 0.0024 |  | 2.2495 | 0.024 | 0 | 0.04508 |  | 0.83936 | 0.011933 | 0.026803 |  |  | 5.07 | 0.023 | 0.11271 | 0.08971 | 6.1113e-08 | 1.4952e-09 | 0.0040433 | 75.2279 | 75.2279 | 75.6579 | 70.25
mH2O | 0.018015 | kg/mol |  | 28.729 | 301.879 | 0.53349 | 3946.698 | 108880 | 0.99812 | 0.99935 | 0.0012334 | -2.5229e-05 | 0.0021108 | 0.00075315 | 2.5198e-05 | 1.0041 | 0.036398 | 5.7605 | 2 | 12.1613 | 5.0698 | 14.5184 | -36.9375 | 236.2125 | 0.86474 | 18.0797 | 108582 | 0.99995 | 0.99821 | -0.0017364 | -4.4949e-05 | 0.0049782 | 0.001086 | 0.00012179 | 1.0045 | 0.00016725 | 7.4843 | 10 | 110.1701 | 1.3814 | 110.8845 | [8 empty] | 0.2 | 0 | 49.8 | 0.2174 | 0 | 43.8 | 0.2174 | 0.1934 | 1.3175e-07 | 2.9838e-05 | 166.0257 | 166.0257 | 166.4557 | 0.24167 | 167.35 |  | 1.8557 | 7.6785 |  |  | 167.35 | [5 empty] | -0.89432 |  |  | 1.8713 |  |  |  | 0.99466 |  |  | 1.0112 | 0.0028 |  | 2.249 | 0.014 | 0 | 0.045161 |  | 1.8714 | 0.011183 | 0.017918 |  |  | 5.07 | 0.023 | 0.22531 | 0.20231 | 1.3782e-07 | 3.3719e-09 | 0.0091255 | 167.3494 | 167.3494 | 167.7794 | 167.35
TRefAlicat | 298.15 | K |  | 28.72 | 301.87 | 0.5335 | 3944.6397 | 108880 | 0.99812 | 0.99935 | 0.0012327 | -2.5231e-05 | 0.0021111 | 0.00075318 | 2.5201e-05 | 1.0041 | 0.036379 | 5.7606 | 2 | 12.1622 | 5.0698 | 14.5192 | -27.675 | 245.475 | 0.89865 | 48.3055 | 108585.5 | 0.9999 | 0.99845 | -0.0014562 | -4.1383e-05 | 0.0044082 | 0.0010447 | 9.3381e-05 | 1.0041 | 0.00044668 | 7.1786 | 10 | 101.931 | 1.3814 | 102.6809 | [8 empty] | 0.5 | 0 | 49.5 | 0.5435 | 0 | 43.5 | 0.5435 | 0.5195 | 3.539e-07 | 2.9634e-05 | 445.3384 | 445.3384 | 445.7684 | 0.64722 | 447 |  | 4.5898 | 7.0916 |  |  | 447 | [5 empty] | -1.2316 |  |  | 4.6352 |  |  |  | 0.99724 |  |  | 1.0104 | 0.004 |  | 2.2475 | 0.008 | 0 | 0.045404 |  | 4.6353 | 0.01037 | 0.013097 |  |  | 5.07 | 0.023 | 0.56327 | 0.54027 | 3.6805e-07 | 9.0045e-09 | 0.02437 | 446.2846 | 446.2846 | 446.7146 | 447
sccmAlicat | 6.8123e-07 | (mol/s)/sccm |  | 28.8455 | 301.9955 | 0.53331 | 3973.4253 | 108880 | 0.99811 | 0.99935 | 0.0012427 | -2.5203e-05 | 0.0021073 | 0.00075267 | 2.5155e-05 | 1.0041 | 0.036645 | 5.7585 | 2 | 12.1506 | 5.0698 | 14.5087 | -20.445 | 252.705 | 0.92512 | 98.9355 | 108585 | 0.99984 | 0.9986 | -0.0012348 | -3.882e-05 | 0.0040144 | 0.0010144 | 7.6909e-05 | 1.0039 | 0.00091467 | 6.9525 | 10 | 96.1838 | 1.3814 | 96.9617 | [8 empty] | 1 | 0 | 49 | 1.087 | 0 | 43 | 1.087 | 1.063 | 7.2415e-07 | 2.9293e-05 | 916.8197 | 916.8197 | 917.2497 | 1.3308 | 916 |  | 8.8817 | 6.6739 |  |  | 916 | [5 empty] | 1.2497 |  |  | 8.9808 |  |  |  | 1.0014 |  |  | 1.0098 | 0.006 |  | 2.245 | 0.006 | 0 | 0.045816 |  | 8.9808 | 0.0098041 | 0.011494 |  |  | 5.07 | 0.023 | 1.1268 | 1.1038 | 7.5197e-07 | 1.8397e-08 | 0.04977 | 917.1419 | 917.1419 | 917.5719 | 916
|  |  |  | 28.8215 | 301.9715 | 0.53334 | 3967.9064 | 108880 | 0.99811 | 0.99935 | 0.0012408 | -2.5208e-05 | 0.002108 | 0.00075276 | 2.5164e-05 | 1.0041 | 0.036594 | 5.7589 | 2 | 12.1528 | 5.0698 | 14.5107 | -12.995 | 260.155 | 0.95239 | 198.6098 | 108584 | 0.99974 | 0.99875 | -0.00099077 | -3.636e-05 | 0.0036483 | 0.00098439 | 6.3654e-05 | 1.0037 | 0.0018359 | 6.7313 | 10 | 90.8313 | 1.3814 | 91.6381 | [8 empty] | 2 | 0 | 48 | 2.174 | 0 | 42 | 2.174 | 2.15 | 1.4647e-06 | 2.8612e-05 | 1846.2954 | 1846.2954 | 1846.7254 | 2.6797 | 1840 |  | 16.8614 | 6.2922 |  |  | 1840 | [5 empty] | 6.7254 |  |  | 17.073 |  |  |  | 1.0037 |  |  | 1.0093 | 0.01 |  | 2.24 | 0.005 | 0 | 0.046667 |  | 17.073 | 0.009278 | 0.010539 |  |  | 5.07 | 0.023 | 2.2536 | 2.2306 | 1.5195e-06 | 3.7176e-08 | 0.10058 | 1845.4233 | 1845.4233 | 1845.8533 | 1840
liquid-vapor | 2002 Wagner Eq. (2.5) |  |  | 28.9605 | 302.1105 | 0.53313 | 3999.9628 | 109760 | 0.9981 | 0.99935 | 0.0012468 | -2.5177e-05 | 0.0021215 | 0.0007585 | 2.5323e-05 | 1.0042 | 0.036594 | 5.7624 | 2 | 12.1401 | 5.0292 | 14.4872 | -50.76 | 222.39 | 0.81414 | 3.5823 | 109020 | 0.99998 | 0.99779 | -0.0021961 | -5.0957e-05 | 0.0060219 | 0.0011583 | 0.00018933 | 1.0052 | 3.303e-05 | 7.9939 | 10 | 124.5452 | 1.3759 | 125.209 | [8 empty] | 0.1 | 0 | 99.9 | 0.1087 | 0 | 93.9 | 0.1087 | 0.0847 | 5.77e-08 | 6.3968e-05 | 34.2308 | 34.2308 | 34.6608 | 0.050214 | 33.05 |  | 0.41382 | 8.2411 |  |  | 33.05 | [5 empty] | 1.6108 |  |  | 0.41685 |  |  |  | 1.0487 |  |  | 1.0126 | 0.0024 |  | 2.4995 | 0.024 | 0 | 0.02502 |  | 0.41754 | 0.012604 | 0.027109 |  |  | 5.07 | 0.023 | 0.11268 | 0.089678 | 6.1092e-08 | 1.4946e-09 | 0.0040419 | 34.9156 | 34.9156 | 35.3456 | 33.05
Tc | 647.096 | K |  | 28.718 | 301.868 | 0.5335 | 3944.1825 | 109760 | 0.99812 | 0.99935 | 0.0012273 | -2.5231e-05 | 0.0021289 | 0.00075951 | 2.5413e-05 | 1.0041 | 0.036084 | 5.7666 | 2 | 12.1624 | 5.0292 | 14.5075 | -43.7425 | 229.4075 | 0.83983 | 8.3493 | 109014 | 0.99997 | 0.99801 | -0.0019586 | -4.7795e-05 | 0.0054737 | 0.0011227 | 0.00015064 | 1.0048 | 7.6957e-05 | 7.7365 | 10 | 116.9141 | 1.376 | 117.6038 | [8 empty] | 0.2 | 0 | 99.8 | 0.2174 | 0 | 93.8 | 0.2174 | 0.1934 | 1.3175e-07 | 6.39e-05 | 77.019 | 77.019 | 77.449 | 0.11236 | 77 |  | 0.90555 | 8.0594 |  |  | 77 | [5 empty] | 0.44897 |  |  | 0.91249 |  |  |  | 1.0058 |  |  | 1.0118 | 0.0028 |  | 2.499 | 0.014 | 0 | 0.02504 |  | 0.9126 | 0.01185 | 0.018342 |  |  | 5.07 | 0.023 | 0.22524 | 0.20224 | 1.3778e-07 | 3.3707e-09 | 0.0091227 | 77.6337 | 77.6337 | 78.0637 | 77
pc | 22.064 | Mpa |  | 28.7215 | 301.8715 | 0.5335 | 3944.9827 | 109760 | 0.99812 | 0.99935 | 0.0012275 | -2.523e-05 | 0.0021288 | 0.00075949 | 2.5412e-05 | 1.0041 | 0.036091 | 5.7665 | 2 | 12.162 | 5.0292 | 14.5072 | -34.945 | 238.205 | 0.87203 | 22.4807 | 109014 | 0.99994 | 0.99826 | -0.0016824 | -4.4153e-05 | 0.004868 | 0.0010811 | 0.00011528 | 1.0044 | 0.00020712 | 7.4309 | 10 | 108.3096 | 1.376 | 109.0325 | [8 empty] | 0.5 | 0 | 99.5 | 0.5435 | 0 | 93.5 | 0.5435 | 0.5195 | 3.539e-07 | 6.3695e-05 | 206.8432 | 206.8432 | 207.2732 | 0.3007 | 207.05 |  | 2.2575 | 7.5076 |  |  | 207.05 | [5 empty] | 0.22321 |  |  | 2.2775 |  |  |  | 1.0011 |  |  | 1.011 | 0.004 |  | 2.4975 | 0.008 | 0 | 0.025101 |  | 2.2775 | 0.010999 | 0.013601 |  |  | 5.07 | 0.023 | 0.56312 | 0.54012 | 3.6794e-07 | 9.0019e-09 | 0.024363 | 207.2874 | 207.2874 | 207.7174 | 207.05
liquid1 | -7.8595 |  |  | 28.8365 | 301.9865 | 0.53332 | 3971.3549 | 109760 | 0.99811 | 0.99935 | 0.0012368 | -2.5205e-05 | 0.0021253 | 0.00075902 | 2.5369e-05 | 1.0042 | 0.036333 | 5.7645 | 2 | 12.1515 | 5.0292 | 14.4976 | -28.054 | 245.096 | 0.89726 | 46.4694 | 109009 | 0.9999 | 0.99843 | -0.0014737 | -4.1523e-05 | 0.0044474 | 0.0010504 | 9.4733e-05 | 1.0041 | 0.00042805 | 7.2031 | 10 | 102.2482 | 1.376 | 102.9974 | [8 empty] | 1 | 0 | 99 | 1.087 | 0 | 93 | 1.087 | 1.063 | 7.2415e-07 | 6.3355e-05 | 425.8902 | 425.8902 | 426.3202 | 0.61806 | 428.35 |  | 4.4119 | 7.1383 |  |  | 428.35 | [5 empty] | -2.0298 |  |  | 4.455 |  |  |  | 0.99526 |  |  | 1.0104 | 0.006 |  | 2.495 | 0.006 | 0 | 0.025202 |  | 4.455 | 0.010401 | 0.012008 |  |  | 5.07 | 0.023 | 1.1265 | 1.1035 | 7.5174e-07 | 1.8392e-08 | 0.049756 | 426.0488 | 426.0488 | 426.4788 | 428.35
liquid2 | 1.8441 |  |  | 28.721 | 301.871 | 0.5335 | 3944.8684 | 109760 | 0.99812 | 0.99935 | 0.0012275 | -2.523e-05 | 0.0021288 | 0.0007595 | 2.5412e-05 | 1.0041 | 0.03609 | 5.7665 | 2 | 12.1621 | 5.0292 | 14.5073 | -21.12 | 252.03 | 0.92265 | 92.6921 | 109055 | 0.99985 | 0.99858 | -0.001262 | -3.9052e-05 | 0.004067 | 0.0010215 | 7.8619e-05 | 1.0039 | 0.00085328 | 6.9853 | 10 | 96.6965 | 1.3755 | 97.4726 | [8 empty] | 2 | 0 | 98 | 2.174 | 0 | 92 | 2.174 | 2.15 | 1.4647e-06 | 6.2673e-05 | 854.2744 | 854.2744 | 854.7044 | 1.2399 | 854.35 |  | 8.3276 | 6.7161 |  |  | 854.35 | [5 empty] | 0.35442 |  |  | 8.4194 |  |  |  | 1.0004 |  |  | 1.0099 | 0.01 |  | 2.49 | 0.005 | 0 | 0.025408 |  | 8.4194 | 0.0098546 | 0.011051 |  |  | 5.07 | 0.023 | 2.2525 | 2.2295 | 1.5188e-06 | 3.7158e-08 | 0.10056 | 853.8859 | 853.8859 | 854.3159 | 854.35
liquid3 | -11.7866 |  |  | 28.875 | 302.025 | 0.53326 | 3980.2181 | 111520 | 0.99811 | 0.99934 | 0.0012294 | -2.5196e-05 | 0.0021594 | 0.00077148 | 2.5777e-05 | 1.0042 | 0.03584 | 5.7757 | 2 | 12.1479 | 4.9498 | 14.4717 | -78 | 195.15 | 0.71442 | 0.075658 | 109888.75 | 1 | 0.99658 | -0.0034229 | -6.5985e-05 | 0.0089684 | 0.0013305 | 0.00054131 | 1.0074 | 6.9362e-07 | 9.2699 | 10 | 162.0448 | 1.365 | 162.6232 | [8 empty] | 0.015 | 0 | 199.985 | 0.016305 | 0 | 193.985 | 0.016305 | -0.007695 | -5.2421e-09 | 0.00013215 | -1.4746 | 0.038005 | 0.468 | 0.00067728 | 0.695 |  | 0.011302 | 16.6877 |  |  | 0.695 | [5 empty] | -0.227 |  |  | 0.011323 |  |  |  | 0.67339 |  |  | 1.0163 | 0.00206 |  | 2.9999 | 0.13733 | 0 | 0.015001 |  | 0.1378 | 0.016327 | 0.1383 |  |  | 5.07 | 0.023 | 0.016889 | -0.0061106 | -4.1628e-09 | -1.0184e-10 | -0.00027549 | #NUM! | #NUM! | #NUM! | 0.695
liquid4 | 22.6807 |  |  | 28.8675 | 302.0175 | 0.53327 | 3978.4902 | 111520 | 0.99811 | 0.99934 | 0.0012288 | -2.5198e-05 | 0.0021597 | 0.00077151 | 2.578e-05 | 1.0042 | 0.035825 | 5.7758 | 2 | 12.1486 | 4.9498 | 14.4724 | -72.675 | 200.475 | 0.73391 | 0.17451 | 109887.5 | 1 | 0.99686 | -0.0031396 | -6.2641e-05 | 0.0082853 | 0.0012951 | 0.00043177 | 1.0069 | 1.5991e-06 | 8.9744 | 10 | 153.5869 | 1.365 | 154.1795 | [8 empty] | 0.025 | 0 | 199.975 | 0.027175 | 0 | 193.975 | 0.027175 | 0.003175 | 2.1629e-09 | 0.00013214 | 0.60816 | 0.77855 | 1.2086 | 0.0017491 | 1.6 |  | 0.024669 | 14.104 |  |  | 1.6 | [5 empty] | -0.39145 |  |  | 0.024731 |  |  |  | 0.75535 |  |  | 1.0155 | 0.0021 |  | 2.9999 | 0.084 | 0 | 0.015001 |  | 0.087565 | 0.015486 | 0.085416 |  |  | 5.07 | 0.023 | 0.028149 | 0.0051485 | 3.5074e-09 | 8.5809e-11 | 0.00023212 | 0.95085 | 1.0393 | 1.4693 | 1.6
liquid5 | -15.9619 |  |  | 28.876 | 302.026 | 0.53326 | 3980.4485 | 111520 | 0.99811 | 0.99934 | 0.0012295 | -2.5196e-05 | 0.0021594 | 0.00077148 | 2.5776e-05 | 1.0042 | 0.035842 | 5.7757 | 2 | 12.1478 | 4.9498 | 14.4716 | -63.875 | 209.275 | 0.76613 | 0.63326 | 109938 | 0.99999 | 0.99727 | -0.0027287 | -5.7581e-05 | 0.0072961 | 0.0012412 | 0.00030489 | 1.0061 | 5.7955e-06 | 8.5622 | 10 | 140.8855 | 1.3644 | 141.5058 | [8 empty] | 0.05 | 0 | 199.95 | 0.05435 | 0 | 193.95 | 0.05435 | 0.03035 | 2.0675e-08 | 0.00013213 | 5.8163 | 5.8163 | 6.2463 | 0.0090394 | 5.8 |  | 0.082073 | 9.0795 |  |  | 5.8 | [5 empty] | 0.44631 |  |  | 0.08257 |  |  |  | 1.077 |  |  | 1.0142 | 0.0022 |  | 2.9997 | 0.044 | 0 | 0.015003 |  | 0.093561 | 0.014224 | 0.046242 |  |  | 5.07 | 0.023 | 0.056298 | 0.033298 | 2.2684e-08 | 5.5497e-10 | 0.0015012 | 6.1525 | 6.1525 | 6.5825 | 5.8
liquid6 | 1.8012 |  |  | 28.9335 | 302.0835 | 0.53317 | 3993.7185 | 111520 | 0.99811 | 0.99934 | 0.0012342 | -2.5183e-05 | 0.0021576 | 0.00077124 | 2.5755e-05 | 1.0042 | 0.035962 | 5.7747 | 2 | 12.1425 | 4.9498 | 14.4668 | -56.62 | 216.53 | 0.79269 | 1.6947 | 109950 | 0.99999 | 0.99756 | -0.0024321 | -5.3793e-05 | 0.0065865 | 0.0011998 | 0.00023372 | 1.0056 | 1.55e-05 | 8.2599 | 10 | 131.4917 | 1.3643 | 132.1369 | [8 empty] | 0.1 | 0 | 199.9 | 0.1087 | 0 | 193.9 | 0.1087 | 0.0847 | 5.77e-08 | 0.00013209 | 16.2877 | 16.2877 | 16.7177 | 0.024185 | 15.52 |  | 0.20508 | 8.4795 |  |  | 15.52 | [5 empty] | 1.1977 |  |  | 0.2065 |  |  |  | 1.0772 |  |  | 1.0133 | 0.0024 |  | 2.9995 | 0.024 | 0 | 0.015005 |  | 0.20789 | 0.013293 | 0.027435 |  |  | 5.07 | 0.023 | 0.11261 | 0.089609 | 6.1045e-08 | 1.4935e-09 | 0.0040391 | 16.6118 | 16.6118 | 17.0418 | 15.52
|  |  |  | 28.7175 | 301.8675 | 0.5335 | 3944.0682 | 111520 | 0.99812 | 0.99934 | 0.0012167 | -2.5231e-05 | 0.0021643 | 0.00077214 | 2.5836e-05 | 1.0042 | 0.035515 | 5.7785 | 2 | 12.1624 | 4.9498 | 14.485 | -49.9275 | 223.2225 | 0.81719 | 3.9716 | 109934.5 | 0.99998 | 0.9978 | -0.0021852 | -5.0569e-05 | 0.0060037 | 0.0011636 | 0.00018567 | 1.0052 | 3.6314e-05 | 7.999 | 10 | 123.6019 | 1.3644 | 124.271 | [8 empty] | 0.2 | 0 | 199.8 | 0.2174 | 0 | 193.8 | 0.2174 | 0.1934 | 1.3175e-07 | 0.00013202 | 36.7083 | 36.7083 | 37.1383 | 0.053795 | 36.35 |  | 0.45172 | 8.3972 |  |  | 36.35 | [5 empty] | 0.78832 |  |  | 0.45492 |  |  |  | 1.0217 |  |  | 1.0125 | 0.0028 |  | 2.999 | 0.014 | 0 | 0.01501 |  | 0.45513 | 0.012511 | 0.018776 |  |  | 5.07 | 0.023 | 0.22512 | 0.20212 | 1.3769e-07 | 3.3687e-09 | 0.0091171 | 37.0008 | 37.0008 | 37.4308 | 36.35
u(liquid-vapor) | 2002 Wagner, Fig. 6.4 |  |  | 28.7425 | 301.8925 | 0.53347 | 3949.7871 | 111520 | 0.99812 | 0.99934 | 0.0012188 | -2.5226e-05 | 0.0021635 | 0.00077204 | 2.5827e-05 | 1.0042 | 0.035566 | 5.778 | 2 | 12.1601 | 4.9498 | 14.4829 | -41.5375 | 231.6125 | 0.8479 | 10.7779 | 109930.5 | 0.99996 | 0.99806 | -0.0019038 | -4.685e-05 | 0.0053587 | 0.0011214 | 0.00014179 | 1.0047 | 9.8506e-05 | 7.6909 | 10 | 114.6609 | 1.3645 | 115.3609 | [8 empty] | 0.5 | 0 | 199.5 | 0.5435 | 0 | 193.5 | 0.5435 | 0.5195 | 3.539e-07 | 0.00013182 | 98.7327 | 98.7327 | 99.1627 | 0.14362 | 98.6 |  | 1.1375 | 7.9201 |  |  | 98.6 | [5 empty] | 0.56274 |  |  | 1.1465 |  |  |  | 1.0057 |  |  | 1.0116 | 0.004 |  | 2.9975 | 0.008 | 0 | 0.015025 |  | 1.1465 | 0.011627 | 0.014113 |  |  | 5.07 | 0.023 | 0.56283 | 0.53983 | 3.6775e-07 | 8.9972e-09 | 0.024348 | 98.947 | 98.947 | 99.377 | 98.6
upV/pV | 0.0002 |  |  | 28.792 | 301.942 | 0.53339 | 3961.1319 | 111520 | 0.99812 | 0.99934 | 0.0012227 | -2.5215e-05 | 0.002162 | 0.00077183 | 2.5808e-05 | 1.0042 | 0.035668 | 5.7772 | 2 | 12.1555 | 4.9498 | 14.4787 | -35.05 | 238.1 | 0.87165 | 22.2261 | 109917 | 0.99994 | 0.99824 | -0.0017001 | -4.4194e-05 | 0.0049152 | 0.0010906 | 0.00011659 | 1.0044 | 0.0002031 | 7.4638 | 10 | 108.4064 | 1.3647 | 109.1307 | [8 empty] | 1 | 0 | 199 | 1.087 | 0 | 193 | 1.087 | 1.063 | 7.2415e-07 | 0.00013148 | 202.5625 | 202.5625 | 202.9925 | 0.29391 | 203.35 |  | 2.2192 | 7.5506 |  |  | 203.35 | [5 empty] | -0.3575 |  |  | 2.2386 |  |  |  | 0.99824 |  |  | 1.011 | 0.006 |  | 2.995 | 0.006 | 0 | 0.01505 |  | 2.2386 | 0.011009 | 0.012538 |  |  | 5.07 | 0.023 | 1.1258 | 1.1028 | 7.5124e-07 | 1.838e-08 | 0.049731 | 202.6454 | 202.6454 | 203.0754 | 203.35
|  |  |  | 28.8785 | 302.0285 | 0.53326 | 3981.0247 | 111520 | 0.99811 | 0.99934 | 0.0012297 | -2.5195e-05 | 0.0021593 | 0.00077147 | 2.5776e-05 | 1.0042 | 0.035848 | 5.7756 | 2 | 12.1476 | 4.9498 | 14.4714 | -28.31 | 244.84 | 0.89632 | 45.2657 | 109987.5 | 0.9999 | 0.99841 | -0.0014956 | -4.1617e-05 | 0.0045024 | 0.001061 | 9.6264e-05 | 1.0042 | 0.00041327 | 7.2401 | 10 | 102.4634 | 1.3638 | 103.2135 | [8 empty] | 2 | 0 | 198 | 2.174 | 0 | 192 | 2.174 | 2.15 | 1.4647e-06 | 0.0001308 | 411.5635 | 411.5635 | 411.9935 | 0.59621 | 413.66 |  | 4.2695 | 7.1611 |  |  | 413.66 | [5 empty] | -1.6665 |  |  | 4.311 |  |  |  | 0.99597 |  |  | 1.0104 | 0.01 |  | 2.99 | 0.005 | 0 | 0.015101 |  | 4.311 | 0.010422 | 0.01156 |  |  | 5.07 | 0.023 | 2.2519 | 2.2289 | 1.5184e-06 | 3.7149e-08 | 0.10049 | 411.3801 | 411.3801 | 411.8101 | 413.66
solid-vapor | 2002 Wagner, Eq. (2.21) |  |  | 28.8855 | 302.0355 | 0.53324 | 3982.6383 | 113280 | 0.99811 | 0.99933 | 0.0012198 | -2.5194e-05 | 0.0021945 | 0.00078406 | 2.6195e-05 | 1.0042 | 0.035306 | 5.7875 | 2 | 12.1469 | 4.8729 | 14.4495 | -60.056 | 213.094 | 0.78011 | 1.0721 | 110988 | 0.99999 | 0.9974 | -0.0025928 | -5.5547e-05 | 0.0069773 | 0.0012306 | 0.00026699 | 1.0059 | 9.7165e-06 | 8.4469 | 10 | 135.8264 | 1.3515 | 136.4624 | [8 empty] | 0.1 | 0 | 299.9 | 0.1087 | 0 | 293.9 | 0.1087 | 0.0847 | 5.77e-08 | 0.00020021 | 10.5443 | 10.5443 | 10.9743 | 0.015857 | 9.73 |  | 0.13278 | 8.3733 |  |  | 9.73 | [5 empty] | 1.2443 |  |  | 0.13372 |  |  |  | 1.1279 |  |  | 1.0137 | 0.0024 |  | 3.4995 | 0.024 | 0 | 0.011669 |  | 0.13586 | 0.013723 | 0.027646 |  |  | 5.07 | 0.023 | 0.11254 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.098947 | 397 | 397 | 397.43 | 413.66
Tn | 273.16 | K |  | 28.717 | 301.867 | 0.53351 | 3943.9539 | 113280 | 0.99812 | 0.99933 | 0.0012062 | -2.5231e-05 | 0.0021998 | 0.00078478 | 2.6259e-05 | 1.0042 | 0.034963 | 5.7906 | 2 | 12.1624 | 4.8729 | 14.4638 | -53.425 | 219.725 | 0.80438 | 2.5614 | 110970 | 0.99999 | 0.99766 | -0.0023331 | -5.2223e-05 | 0.0063591 | 0.0011933 | 0.00021099 | 1.0054 | 2.3207e-05 | 8.1762 | 10 | 127.6367 | 1.3517 | 128.2957 | [8 empty] | 0.2 | 0 | 299.8 | 0.2174 | 0 | 293.8 | 0.2174 | 0.1934 | 1.3175e-07 | 0.00020015 | 23.8328 | 23.8328 | 24.2628 | 0.035093 | 23.2 |  | 0.29765 | 8.4816 |  |  | 23.2 | [5 empty] | 1.0628 |  |  | 0.29971 |  |  |  | 1.0458 |  |  | 1.0129 | 0.0028 |  | 3.499 | 0.014 | 0 | 0.011671 |  | 0.30003 | 0.012911 | 0.019044 |  |  | 5.07 | 0.023 | 0.225 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.099002 | 397 | 397 | 397.43 | 413.66
pn | 0.00061166 | Mpa |  | 28.725 | 301.875 | 0.53349 | 3945.7831 | 113280 | 0.99812 | 0.99933 | 0.0012069 | -2.5229e-05 | 0.0021995 | 0.00078475 | 2.6256e-05 | 1.0042 | 0.034979 | 5.7904 | 2 | 12.1617 | 4.8729 | 14.4631 | -45.182 | 227.968 | 0.83456 | 7.0487 | 110951 | 0.99997 | 0.99793 | -0.0020419 | -4.8424e-05 | 0.0056805 | 0.0011499 | 0.00016049 | 1.005 | 6.3845e-05 | 7.8607 | 10 | 118.4214 | 1.3519 | 119.1103 | [8 empty] | 0.5 | 0 | 299.5 | 0.5435 | 0 | 293.5 | 0.5435 | 0.5195 | 3.539e-07 | 0.00019994 | 64.0408 | 64.0408 | 64.4708 | 0.093245 | 63.88 |  | 0.76088 | 8.16 |  |  | 63.88 | [5 empty] | 0.59077 |  |  | 0.76657 |  |  |  | 1.0092 |  |  | 1.012 | 0.004 |  | 3.4975 | 0.008 | 0 | 0.011678 |  | 0.76661 | 0.011999 | 0.014421 |  |  | 5.07 | 0.023 | 0.56251 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.098999 | 397 | 397 | 397.43 | 413.66
solid1 | -13.9282 |  |  | 28.7315 | 301.8815 | 0.53348 | 3947.2699 | 113280 | 0.99812 | 0.99933 | 0.0012074 | -2.5228e-05 | 0.0021993 | 0.00078472 | 2.6254e-05 | 1.0042 | 0.034993 | 5.7903 | 2 | 12.1611 | 4.8729 | 14.4625 | -38.93 | 234.22 | 0.85745 | 14.4865 | 111412.5 | 0.99996 | 0.99811 | -0.0018461 | -4.5761e-05 | 0.0052452 | 0.0011238 | 0.00013267 | 1.0047 | 0.00013063 | 7.6498 | 10 | 112.0812 | 1.3463 | 112.7942 | [8 empty] | 1 | 0 | 299 | 1.087 | 0 | 293 | 1.087 | 1.063 | 7.2415e-07 | 0.0001996 | 131.0631 | 131.0631 | 131.4931 | 0.19017 | 130.6 |  | 1.4731 | 7.7461 |  |  | 130.6 | [5 empty] | 0.89314 |  |  | 1.4853 |  |  |  | 1.0068 |  |  | 1.0114 | 0.006 |  | 3.495 | 0.006 | 0 | 0.011689 |  | 1.4853 | 0.011372 | 0.012858 |  |  | 5.07 | 0.023 | 1.125 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.098997 | 397 | 397 | 397.43 | 413.66
solid2 | 34.7078 |  |  | 28.8895 | 302.0395 | 0.53324 | 3983.5606 | 113280 | 0.99811 | 0.99933 | 0.0012202 | -2.5193e-05 | 0.0021944 | 0.00078405 | 2.6193e-05 | 1.0042 | 0.035314 | 5.7874 | 2 | 12.1466 | 4.8729 | 14.4492 | -32.38 | 240.77 | 0.88142 | 29.6009 | 111022.5 | 0.99993 | 0.99829 | -0.0016354 | -4.3152e-05 | 0.0047933 | 0.0010892 | 0.00010899 | 1.0044 | 0.00026779 | 7.4079 | 10 | 105.987 | 1.3511 | 106.7237 | [8 empty] | 2 | 0 | 298 | 2.174 | 0 | 292 | 2.174 | 2.15 | 1.4647e-06 | 0.00019892 | 267.4981 | 267.4981 | 267.9281 | 0.38713 | 267.87 |  | 2.8588 | 7.3845 |  |  | 267.87 | [5 empty] | 0.058068 |  |  | 2.8849 |  |  |  | 1.0002 |  |  | 1.0108 | 0.01 |  | 3.49 | 0.005 | 0 | 0.011711 |  | 2.8849 | 0.01077 | 0.011874 |  |  | 5.07 | 0.023 | 2.2508 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.098945 | 397 | 397 | 397.43 | 413.66
|  |  |  | 28.95 | 302.1 | 0.53315 | 3997.5334 | 111520 | 0.99811 | 0.99934 | 0.0012355 | -2.5179e-05 | 0.0021571 | 0.00077117 | 2.5749e-05 | 1.0042 | 0.035996 | 5.7744 | 2 | 12.141 | 4.9498 | 14.4654 | -36.285 | 236.865 | 0.86713 | 19.4246 | 109892 | 0.99995 | 0.99821 | -0.0017378 | -4.4686e-05 | 0.004995 | 0.001096 | 0.00012089 | 1.0045 | 0.00017755 | 7.5054 | 10 | 109.5553 | 1.365 | 110.2749 | [8 empty] | 0 | 1 | 199 | 0 | 0.952 | 193 | 0.952 | 0.928 | 6.3218e-07 | 0.00013148 | 178.6529 | 178.6529 | 179.0829 | 0.25905 | 177.8 |  | 1.9607 | 7.5687 |  |  | 177.8 | [5 empty] | 1.2829 |  |  | 1.9777 |  |  |  | 1.0072 |  |  | 1.0111 |  | 0.204 | 2.995 | 0 | 0.204 | 0.01505 |  | 1.9882 | 0.011122 | 0.2043 |  |  | 5.07 | 0.023 | 0.98627 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.098926 | 397 | 397 | 397.43 | 413.66
u(solid-vapor) | 1994 Wagner, Table 5 |  |  | 28.9235 | 302.0735 | 0.53319 | 3991.4079 | 111520 | 0.99811 | 0.99934 | 0.0012334 | -2.5185e-05 | 0.0021579 | 0.00077128 | 2.5759e-05 | 1.0042 | 0.035941 | 5.7748 | 2 | 12.1435 | 4.9498 | 14.4676 | -29.593 | 243.557 | 0.89163 | 39.654 | 109895 | 0.99991 | 0.99838 | -0.0015333 | -4.2094e-05 | 0.0045745 | 0.0010657 | 9.9688e-05 | 1.0042 | 0.00036236 | 7.2796 | 10 | 103.5532 | 1.3649 | 104.2983 | [8 empty] | 0 | 2 | 198 | 0 | 1.904 | 192 | 1.904 | 1.88 | 1.2807e-06 | 0.0001308 | 361.3738 | 361.3738 | 361.8038 | 0.52344 | 362.75 |  | 3.7834 | 7.2279 |  |  | 362.75 | [5 empty] | -0.9462 |  |  | 3.8195 |  |  |  | 0.99739 |  |  | 1.0105 |  | 0.208 | 2.99 | 0 | 0.104 | 0.015101 |  | 3.8209 | 0.01053 | 0.10453 |  |  | 5.07 | 0.023 | 1.9724 | 2.1948 | 1.4952e-06 | 3.658e-08 | 0.098934 | 397 | 397 | 397.43 | 413.66
upV/pV | 0.001 |  |  | 28.851 | 302.001 | 0.5333 | 3974.691 | 111520 | 0.99811 | 0.99934 | 0.0012275 | -2.5201e-05 | 0.0021602 | 0.00077158 | 2.5786e-05 | 1.0042 | 0.035791 | 5.7761 | 2 | 12.1501 | 4.9498 | 14.4737 | -20.3985 | 252.7515 | 0.92529 | 99.3795 | 109904.5 | 0.99984 | 0.99859 | -0.0012503 | -3.8804e-05 | 0.0040609 | 0.0010265 | 7.775e-05 | 1.0039 | 0.00090778 | 6.9849 | 10 | 96.1486 | 1.3648 | 96.9289 | [8 empty] | 0 | 5 | 195 | 0 | 4.76 | 189 | 4.76 | 4.736 | 3.2263e-06 | 0.00012875 | 906.576 | 906.576 | 907.006 | 1.3128 | 910.5 |  | 8.8254 | 6.7227 |  |  | 910.5 | [5 empty] | -3.494 |  |  | 8.9225 |  |  |  | 0.99616 |  |  | 1.0098 |  | 0.22 | 2.975 | 0 | 0.044 | 0.015256 |  | 8.9226 | 0.0098004 | 0.045078 |  |  | 5.07 | 0.023 | 4.9304 | 4.9074 | 3.3431e-06 | 8.1789e-08 | 0.22126 | 905.7786 | 905.7786 | 906.2086 | 910.5
|  |  |  | 28.82 | 301.97 | 0.53335 | 3967.5617 | 111520 | 0.99811 | 0.99934 | 0.001225 | -2.5208e-05 | 0.0021611 | 0.00077171 | 2.5798e-05 | 1.0042 | 0.035726 | 5.7767 | 2 | 12.153 | 4.9498 | 14.4764 | -13.065 | 260.085 | 0.95213 | 197.3505 | 109946 | 0.99974 | 0.99874 | -0.0010089 | -3.6382e-05 | 0.0036976 | 0.00099704 | 6.4566e-05 | 1.0038 | 0.0018017 | 6.7639 | 10 | 90.879 | 1.3643 | 91.6875 | [8 empty] | 0 | 10 | 190 | 0 | 9.52 | 184 | 9.52 | 9.496 | 6.469e-06 | 0.00012535 | 1814.9611 | 1814.9611 | 1815.3911 | 2.628 | 1806 |  | 16.5588 | 6.3008 |  |  | 1806 | [5 empty] | 9.3911 |  |  | 16.766 |  |  |  | 1.0052 |  |  | 1.0093 |  | 0.24 | 2.95 | 0 | 0.024 | 0.015526 |  | 16.766 | 0.0092823 | 0.025733 |  |  | 5.07 | 0.023 | 9.8601 | 9.8371 | 6.7014e-06 | 1.6395e-07 | 0.44358 | 1813.0673 | 1813.0673 | 1813.4973 | 1806
Zw0(T) |  |  |  | 28.851 | 302.001 | 0.5333 | 3974.691 | 111520 | 0.99811 | 0.99934 | 0.0012275 | -2.5201e-05 | 0.0021602 | 0.00077158 | 2.5786e-05 | 1.0042 | 0.035791 | 5.7761 | 2 | 12.1501 | 4.9498 | 14.4737 | -5.114 | 268.036 | 0.98124 | 397.8663 | 109896 | 0.99959 | 0.99888 | -0.00071115 | -3.3936e-05 | 0.0033385 | 0.00096526 | 5.3301e-05 | 1.0037 | 0.0036336 | 6.5362 | 10 | 85.7454 | 1.3649 | 86.5844 | [8 empty] | 0 | 20 | 180 | 0 | 19.04 | 174 | 19.04 | 19.016 | 1.2954e-05 | 0.00011853 | 3643.6556 | 3643.6556 | 3644.0856 | 5.2744 | 3649.5 |  | 31.599 | 5.9911 |  |  | 3649.5 | [5 empty] | -5.4144 |  |  | 32.0361 |  |  |  | 0.99852 |  |  | 1.0088 |  | 0.28 | 2.9 | 0 | 0.014 | 0.016111 |  | 32.0361 | 0.0087786 | 0.016525 |  |  | 5.07 | 0.023 | 19.7215 | 19.6985 | 1.3419e-05 | 3.2831e-07 | 0.88815 | 3627.6956 | 3627.6956 | 3628.1256 | 3649.5
zTsol | 1899 | [6 empty] | 111520 | [41 empty] | 0 | 50 | 150 | 0 | 47.6 | 144 | 47.6 | 47.576 | 3.241e-05 | 9.8098e-05 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | 9020 | [5 empty] | 9020 | [17 empty] | 0.4 | 2.75 | 0 | 0.008 | 0.018333 | [6 empty] | 5.07 | 0.023 | 47.6 | 47.577 | 3.2411e-05 | [5 empty] | 9020
Tzsol | 4114 | K | [5 empty] | 108000
zTliq | 124 |  |  | 28.9335 | 302.0835 | 0.53317 | 3993.7185 | 111520 | 0.99811 | 0.99934 | 0.0012342 | -2.5183e-05 | 0.0021576 | 0.00077124 | 2.5755e-05 | 1.0042 | 0.035962 | 5.7747 | 2 | 12.1425 | 4.9498 | 14.4668 |  | 273.15 | 0.57788 | 611.2128 | 109950 | 0.99942 | 0.99896 | -0.00045214 | -3.2453e-05 | 0.0031281 | 0.00086731 | 4.7016e-05 | 1.0036 | 0.005579 | 6.3997 | 10 | 72.4876 | 1.3643 | 73.4662 | [8 empty] | 0.1 | 0 | 199.9 | 0.1087 | 0 | 193.9 | 0.1087 | 0.0847 | 5.77e-08 | 0.00013209 | 16.2877 | 16.2877 | 16.7177 | 0.024185 |  |  | 0 | 0 |  |  | 0 | [5 empty] | 16.7177 |  |  | 0.024185 | [6 empty] | 1.0075 | 0.0024 |  | 2.9995 | 0.024 | 0 | 0.015005 |  | 0.024 |  |  |  |  | 0.06 | 0.023 | 0.11261 | 0.089609 | 6.1045e-08 | 1.4935e-09 | 0.0040391 | 16.6118 | 16.6118 | 17.0418
Tzliq | 3350 | K |  | 28.489 | 301.639 | 0.53386 | 3892.13 | 111520 | 0.99814 | 0.99933 | 0.0011984 | -2.5282e-05 | 0.0021714 | 0.0007731 | 2.5924e-05 | 1.0042 | 0.035047 | 5.7826 | 2 | 12.1834 | 4.9498 | 14.5043 |  | 273.15 | 0.57788 | 611.2128 | 109875 | 0.99942 | 0.99896 | -0.00045143 | -3.2453e-05 | 0.0031259 | 0.00086671 | 4.6983e-05 | 1.0036 | 0.0055828 | 6.3984 | 10 | 72.4876 | 1.3652 | 73.4661 | [8 empty] | 0.2 | 0 | 199.8 | 0.2174 | 0 | 193.8 | 0.2174 | 0.1934 | 1.3175e-07 | 0.00013202 | 36.207 | 36.207 | 36.637 | 0.053139 | 38.44 |  | 0.2824 | 5.3144 |  |  | 38.44 | [5 empty] | -1.803 |  |  | 0.28736 |  |  |  | 0.9531 |  |  | 1.0075 | 0.0028 |  | 2.999 | 0.014 | 0 | 0.01501 |  | 0.28275 |  |  |  |  | 0.06 | 0.00027219 | 0.22502 | 0.22475 | 1.5311e-07 | 3.7458e-09 | 0.010145 | 40.4862 | 40.4862 | 40.9162 | 38.44
|  |  |  | 28.7425 | 301.8925 | 0.53347 | 3949.7871 | 111520 | 0.99812 | 0.99934 | 0.0012188 | -2.5226e-05 | 0.0021635 | 0.00077204 | 2.5827e-05 | 1.0042 | 0.035566 | 5.778 | 2 | 12.1601 | 4.9498 | 14.4829 |  | 273.15 | 0.57788 | 611.2128 | 109930.5 | 0.99942 | 0.99896 | -0.00045196 | -3.2453e-05 | 0.0031275 | 0.00086715 | 4.7007e-05 | 1.0036 | 0.00558 | 6.3994 | 10 | 72.4876 | 1.3645 | 73.4661 | [8 empty] | 0.5 | 0 | 199.5 | 0.5435 | 0 | 193.5 | 0.5435 | 0.5195 | 3.539e-07 | 0.00013182 | 98.7327 | 98.7327 | 99.1627 | 0.14362 |  |  | 0 | 0 |  |  | 0 | [5 empty] | 99.1627 |  |  | 0.14362 | [6 empty] | 1.0075 | 0.004 |  | 2.9975 | 0.008 | 0 | 0.015025 |  | 0.008 |  |  |  |  | 0.06 | 0.00027219 | 0.56283 | 0.56256 | 3.8323e-07 | 9.376e-09 | 0.025373 | 93.4114 | 93.4114 | 93.8414
dZar/dp | fit Refprop |  |  | 28.513 | 301.663 | 0.53382 | 3897.5571 | 111520 | 0.99814 | 0.99933 | 0.0012003 | -2.5277e-05 | 0.0021706 | 0.000773 | 2.5915e-05 | 1.0042 | 0.035095 | 5.7821 | 2 | 12.1812 | 4.9498 | 14.5023 |  | 273.15 | 0.57788 | 611.2128 | 109862 | 0.99942 | 0.99896 | -0.00045131 | -3.2453e-05 | 0.0031255 | 0.00086661 | 4.6978e-05 | 1.0036 | 0.0055835 | 6.3982 | 10 | 72.4876 | 1.3653 | 73.466 | [8 empty] | 1 | 0 | 199 | 1.087 | 0 | 193 | 1.087 | 1.063 | 7.2415e-07 | 0.00013148 | 199.1913 | 199.1913 | 199.6213 | 0.2895 | 130.92 |  | 0.96182 | 3.3224 |  |  | 130.92 | [5 empty] | 68.7013 |  |  | 1.0044 |  |  |  | 1.5248 |  |  | 1.0075 | 0.006 |  | 2.995 | 0.006 | 0 | 0.01505 |  | 0.96184 |  |  |  |  | 0.06 | 0.00027219 | 1.1251 | 1.1249 | 7.663e-07 | 1.8748e-08 | 0.050775 | 140.9803 | 140.9803 | 141.4103 | 130.92
| 0.00801 |  |  | 28.543 | 301.693 | 0.53377 | 3904.3502 | 111520 | 0.99813 | 0.99933 | 0.0012027 | -2.527e-05 | 0.0021697 | 0.00077288 | 2.5903e-05 | 1.0042 | 0.035157 | 5.7816 | 2 | 12.1785 | 4.9498 | 14.4997 |  | 273.15 | 0.57788 | 611.2128 | 109859 | 0.99942 | 0.99896 | -0.00045128 | -3.2453e-05 | 0.0031254 | 0.00086658 | 4.6977e-05 | 1.0036 | 0.0055836 | 6.3982 | 10 | 72.4876 | 1.3654 | 73.466 | [8 empty] | 2 | 0 | 198 | 2.174 | 0 | 192 | 2.174 | 2.15 | 1.4647e-06 | 0.0001308 | 403.3449 | 403.3449 | 403.7749 | 0.58546 | 176.84 |  | 1.2992 | 2.2191 |  |  | 176.84 | [5 empty] | 226.9349 |  |  | 1.425 |  |  |  | 2.2833 |  |  | 1.0075 | 0.01 |  | 2.99 | 0.005 | 0 | 0.015101 |  | 1.2992 |  |  |  |  | 0.06 | 0.00027219 | 2.2504 | 2.2502 | 1.5329e-06 | 3.7503e-08 | 0.10156 | 181.6796 | 181.6796 | 182.1096 | 176.84
| -1.54 |  |  | 28.851 | 302.001 | 0.5333 | 3974.691 | 111520 | 0.99811 | 0.99934 | 0.0012275 | -2.5201e-05 | 0.0021602 | 0.00077158 | 2.5786e-05 | 1.0042 | 0.035791 | 5.7761 | 2 | 12.1501 | 4.9498 | 14.4737 |  | 273.15 | 0.57788 | 611.2128 | 109904.5 | 0.99942 | 0.99896 | -0.00045171 | -3.2453e-05 | 0.0031268 | 0.00086695 | 4.6996e-05 | 1.0036 | 0.0055813 | 6.3989 | 10 | 72.4876 | 1.3648 | 73.4661 | [8 empty] | 0 | 5 | 195 | 0 | 4.76 | 189 | 4.76 | 4.736 | 3.2263e-06 | 0.00012875 | 906.576 | 906.576 | 907.006 | 1.3128 |  |  | 0 | 0 |  |  | 0 | [5 empty] | 907.006 |  |  | 1.3128 | [6 empty] | 1.0075 |  | 0.22 | 2.975 | 0 | 0.044 | 0.015256 |  | 0.044 |  |  |  |  | 0.06 | 0.00027219 | 4.9304 | 4.9301 | 3.3585e-06 | 8.2168e-08 | 0.22229 | 215.24 | 215.24 | 215.67
| -108 |  |  | 28.464 | 301.614 | 0.5339 | 3886.4838 | 111520 | 0.99814 | 0.99933 | 0.0011964 | -2.5288e-05 | 0.0021722 | 0.00077321 | 2.5933e-05 | 1.0042 | 0.034996 | 5.783 | 2 | 12.1858 | 4.9498 | 14.5064 |  | 273.15 | 0.57788 | 611.2128 | 109866 | 0.99942 | 0.99896 | -0.00045135 | -3.2453e-05 | 0.0031256 | 0.00086664 | 4.698e-05 | 1.0036 | 0.0055833 | 6.3983 | 10 | 72.4876 | 1.3653 | 73.4661 | [8 empty] | 0 | 10 | 190 | 0 | 9.52 | 184 | 9.52 | 9.496 | 6.469e-06 | 0.00012535 | 1776.5659 | 1776.5659 | 1776.9959 | 2.5778 | 227.2 |  | 1.6691 | 0.64751 |  |  | 227.2 | [5 empty] | 1549.7959 |  |  | 3.071 |  |  |  | 7.8213 |  |  | 1.0075 |  | 0.24 | 2.95 | 0 | 0.024 | 0.015526 |  | 1.6693 |  |  |  |  | 0.06 | 0.00027219 | 9.8532 | 9.8529 | 6.7121e-06 | 1.6421e-07 | 0.44481 | 224.4496 | 224.4496 | 224.8796 | 227.2
| -211000 |  |  | 28.851 | 302.001 | 0.5333 | 3974.691 | 111520 | 0.99811 | 0.99934 | 0.0012275 | -2.5201e-05 | 0.0021602 | 0.00077158 | 2.5786e-05 | 1.0042 | 0.035791 | 5.7761 | 2 | 12.1501 | 4.9498 | 14.4737 |  | 273.15 | 0.57788 | 611.2128 | 109896 | 0.99942 | 0.99896 | -0.00045163 | -3.2453e-05 | 0.0031265 | 0.00086688 | 4.6992e-05 | 1.0036 | 0.0055817 | 6.3988 | 10 | 72.4876 | 1.3649 | 73.4661 | [8 empty] | 0 | 20 | 180 | 0 | 19.04 | 174 | 19.04 | 19.016 | 1.2954e-05 | 0.00011853 | 3643.6556 | 3643.6556 | 3644.0856 | 5.2744 |  |  | 0 | 0 |  |  | 0 | [5 empty] | 3644.0856 |  |  | 5.2744 | [6 empty] | 1.0075 |  | 0.28 | 2.9 | 0 | 0.014 | 0.016111 |  | 0.014 |  |  |  |  | 0.06 | 0.00027219 | 19.7215 | 19.7212 | 1.3435e-05 | 3.2869e-07 | 0.88918 | 237.7476 | 237.7476 | 238.1776
B12 | 2002 Hodges, Eq. (11) |  |  | 28.8295 | 301.9795 | 0.53333 | 3969.7453 | 111520 | 0.99811 | 0.99934 | 0.0012258 | -2.5206e-05 | 0.0021608 | 0.00077167 | 2.5794e-05 | 1.0042 | 0.035746 | 5.7765 | 2 | 12.1521 | 4.9498 | 14.4756 |  | 273.15 | 0.57788 | 611.2128 | 109849 | 0.99942 | 0.99896 | -0.00045119 | -3.2453e-05 | 0.0031252 | 0.00086651 | 4.6972e-05 | 1.0036 | 0.0055841 | 6.398 | 10 | 72.4876 | 1.3655 | 73.466 | [8 empty] | 0.1 | 0 | 199.9 | 0.1087 | 0 | 193.9 | 0.1087 | 0.0847 | 5.77e-08 | 0.00013209 | 16.1862 | 16.1862 | 16.6162 | 0.024053 | 21.2 |  | 0.15575 | 6.4752 |  |  | 21.2 | [5 empty] | -4.5838 |  |  | 0.15759 |  |  |  | 0.78378 |  |  | 1.0075 | 0.0024 |  | 2.9995 | 0.024 | 0 | 0.015005 |  | 0.15759 |  |  |  |  | 0.5 | 0.0022682 | 0.11259 | 0.11032 | 7.5152e-08 | 1.8386e-09 | 0.0049743 | 20.3257 | 20.3257 | 20.7557 | 21.2
a1 | 96.1591 | cm3/mol |  | 28.799 | 301.949 | 0.53338 | 3962.7385 | 111520 | 0.99812 | 0.99934 | 0.0012233 | -2.5213e-05 | 0.0021618 | 0.0007718 | 2.5805e-05 | 1.0042 | 0.035683 | 5.777 | 2 | 12.1549 | 4.9498 | 14.4781 |  | 273.15 | 0.57788 | 611.2128 | 109845 | 0.99942 | 0.99896 | -0.00045115 | -3.2453e-05 | 0.003125 | 0.00086647 | 4.6971e-05 | 1.0036 | 0.0055843 | 6.3979 | 10 | 72.4876 | 1.3656 | 73.466 | [8 empty] | 0.2 | 0 | 199.8 | 0.2174 | 0 | 193.8 | 0.2174 | 0.1934 | 1.3175e-07 | 0.00013202 | 36.8887 | 36.8887 | 37.3187 | 0.05403 | 42.25 |  | 0.31039 | 5.7448 |  |  | 42.25 | [5 empty] | -4.9313 |  |  | 0.31506 |  |  |  | 0.88328 |  |  | 1.0075 | 0.0028 |  | 2.999 | 0.014 | 0 | 0.01501 |  | 0.31071 |  |  |  |  | 0.5 | 0.0022682 | 0.22516 | 0.22289 | 1.5184e-07 | 3.7148e-09 | 0.010051 | 40.9916 | 40.9916 | 41.4216 | 42.25
a2 | -211.074 | cm3/mol |  | 28.7425 | 301.8925 | 0.53347 | 3949.7871 | 111520 | 0.99812 | 0.99934 | 0.0012188 | -2.5226e-05 | 0.0021635 | 0.00077204 | 2.5827e-05 | 1.0042 | 0.035566 | 5.778 | 2 | 12.1601 | 4.9498 | 14.4829 |  | 273.15 | 0.57788 | 611.2128 | 109930.5 | 0.99942 | 0.99896 | -0.00045196 | -3.2453e-05 | 0.0031275 | 0.00086715 | 4.7007e-05 | 1.0036 | 0.00558 | 6.3994 | 10 | 72.4876 | 1.3645 | 73.4661 | [8 empty] | 0.5 | 0 | 199.5 | 0.5435 | 0 | 193.5 | 0.5435 | 0.5195 | 3.539e-07 | 0.00013182 | 98.7327 | 98.7327 | 99.1627 | 0.14362 |  |  | 0 | 0 |  |  | 0 | [5 empty] | 99.1627 |  |  | 0.14362 | [6 empty] | 1.0075 | 0.004 |  | 2.9975 | 0.008 | 0 | 0.015025 |  | 0.008 |  |  |  |  | 0.5 | 0.0022682 | 0.56283 | 0.56056 | 3.8187e-07 | 9.3427e-09 | 0.025283 | 102.736 | 102.736 | 103.166
a3 | -96.4425 | cm3/mol |  | 28.649 | 301.799 | 0.53361 | 3928.4352 | 111520 | 0.99813 | 0.99934 | 0.0012112 | -2.5246e-05 | 0.0021664 | 0.00077243 | 2.5862e-05 | 1.0042 | 0.035374 | 5.7797 | 2 | 12.1687 | 4.9498 | 14.4908 |  | 273.15 | 0.57788 | 611.2128 | 109875 | 0.99942 | 0.99896 | -0.00045143 | -3.2453e-05 | 0.0031259 | 0.00086671 | 4.6983e-05 | 1.0036 | 0.0055828 | 6.3984 | 10 | 72.4876 | 1.3652 | 73.4661 | [8 empty] | 1 | 0 | 199 | 1.087 | 0 | 193 | 1.087 | 1.063 | 7.2415e-07 | 0.00013148 | 200.8282 | 200.8282 | 201.2582 | 0.29164 | 210.22 |  | 1.5444 | 5.2956 |  |  | 210.22 | [5 empty] | -8.9618 |  |  | 1.5717 |  |  |  | 0.95737 |  |  | 1.0075 | 0.006 |  | 2.995 | 0.006 | 0 | 0.01505 |  | 1.5444 |  |  |  |  | 0.5 | 0.0022682 | 1.1255 | 1.1232 | 7.6515e-07 | 1.872e-08 | 0.050675 | 204.6591 | 204.6591 | 205.0891 | 210.22
a4 | -12.6006 | cm3/mol |  | 28.614 | 301.764 | 0.53366 | 3920.4684 | 111520 | 0.99813 | 0.99934 | 0.0012084 | -2.5254e-05 | 0.0021675 | 0.00077258 | 2.5876e-05 | 1.0042 | 0.035302 | 5.7803 | 2 | 12.1719 | 4.9498 | 14.4937 |  | 273.15 | 0.57788 | 611.2128 | 109888 | 0.99942 | 0.99896 | -0.00045156 | -3.2453e-05 | 0.0031263 | 0.00086681 | 4.6989e-05 | 1.0036 | 0.0055821 | 6.3987 | 10 | 72.4876 | 1.365 | 73.4661 | [8 empty] | 2 | 0 | 198 | 2.174 | 0 | 192 | 2.174 | 2.15 | 1.4647e-06 | 0.0001308 | 405.0716 | 405.0716 | 405.5016 | 0.58772 | 420.5 |  | 3.0892 | 5.2563 |  |  | 420.5 | [5 empty] | -14.9984 |  |  | 3.1447 |  |  |  | 0.96433 |  |  | 1.0075 | 0.01 |  | 2.99 | 0.005 | 0 | 0.015101 |  | 3.0893 |  |  |  |  | 0.5 | 0.0022682 | 2.2507 | 2.2485 | 1.5317e-06 | 3.7475e-08 | 0.10146 | 405.6706 | 405.6706 | 406.1006 | 420.5
b1 | -0.31 |  |  | 28.851 | 302.001 | 0.5333 | 3974.691 | 111520 | 0.99811 | 0.99934 | 0.0012275 | -2.5201e-05 | 0.0021602 | 0.00077158 | 2.5786e-05 | 1.0042 | 0.035791 | 5.7761 | 2 | 12.1501 | 4.9498 | 14.4737 |  | 273.15 | 0.57788 | 611.2128 | 109904.5 | 0.99942 | 0.99896 | -0.00045171 | -3.2453e-05 | 0.0031268 | 0.00086695 | 4.6996e-05 | 1.0036 | 0.0055813 | 6.3989 | 10 | 72.4876 | 1.3648 | 73.4661 | [8 empty] | 0 | 5 | 195 | 0 | 4.76 | 189 | 4.76 | 4.736 | 3.2263e-06 | 0.00012875 | 906.576 | 906.576 | 907.006 | 1.3128 |  |  | 0 | 0 |  |  | 0 | [5 empty] | 907.006 |  |  | 1.3128 | [6 empty] | 1.0075 |  | 0.22 | 2.975 | 0 | 0.044 | 0.015256 |  | 0.044 |  |  |  |  | 0.5 | 0.0022682 | 4.9304 | 4.9281 | 3.3572e-06 | 8.2135e-08 | 0.2222 | 813.6726 | 813.6726 | 814.1026
b2 | -0.82 |  |  | 28.586 | 301.736 | 0.53371 | 3914.105 | 111520 | 0.99813 | 0.99934 | 0.0012062 | -2.5261e-05 | 0.0021684 | 0.0007727 | 2.5887e-05 | 1.0042 | 0.035245 | 5.7808 | 2 | 12.1745 | 4.9498 | 14.4961 |  | 273.15 | 0.57788 | 611.2128 | 109882 | 0.99942 | 0.99896 | -0.0004515 | -3.2453e-05 | 0.0031261 | 0.00086677 | 4.6986e-05 | 1.0036 | 0.0055824 | 6.3986 | 10 | 72.4876 | 1.3651 | 73.4661 | [8 empty] | 0 | 10 | 190 | 0 | 9.52 | 184 | 9.52 | 9.496 | 6.469e-06 | 0.00012535 | 1789.6398 | 1789.6398 | 1790.0698 | 2.5949 | 1212.76 |  | 8.9097 | 3.4335 |  |  | 1212.76 | [5 empty] | 577.3098 |  |  | 9.2799 |  |  |  | 1.476 |  |  | 1.0075 |  | 0.24 | 2.95 | 0 | 0.024 | 0.015526 |  | 8.9097 |  |  |  |  | 0.5 | 0.0022682 | 9.8555 | 9.8533 | 6.7124e-06 | 1.6422e-07 | 0.44465 | 1209.5347 | 1209.5347 | 1209.9647 | 1212.76
b3 | -2.24 |  |  | 28.844 | 301.994 | 0.53331 | 3973.0802 | 111520 | 0.99811 | 0.99934 | 0.0012269 | -2.5203e-05 | 0.0021604 | 0.00077161 | 2.5788e-05 | 1.0042 | 0.035776 | 5.7762 | 2 | 12.1508 | 4.9498 | 14.4743 |  | 273.15 | 0.57788 | 611.2128 | 109840 | 0.99942 | 0.99896 | -0.0004511 | -3.2453e-05 | 0.0031249 | 0.00086643 | 4.6968e-05 | 1.0036 | 0.0055846 | 6.3978 | 10 | 72.4876 | 1.3656 | 73.466 | [8 empty] | 0 | 20 | 180 | 0 | 19.04 | 174 | 19.04 | 19.016 | 1.2954e-05 | 0.00011853 | 3642.1285 | 3642.1285 | 3642.5585 | 5.2724 | 1571.15 |  | 11.5426 | 2.1893 |  |  | 1571.15 | [5 empty] | 2071.4085 |  |  | 12.6897 |  |  |  | 2.3184 |  |  | 1.0075 |  | 0.28 | 2.9 | 0 | 0.014 | 0.016111 |  | 11.5426 |  |  |  |  | 0.5 | 0.0022682 | 19.7212 | 19.7189 | 1.3433e-05 | 3.2865e-07 | 0.8891 | 1566.4449 | 1566.4449 | 1566.8749 | 1571.15
b4 | -4.6
rhoCondensed | constant values
rhoSolid | 917 | kg/m3
rhoLiquid | 1000 | kg/m3
Raoult
kH0 | 0.00043 | 1/MPa
kH1 | -1.02e-05 | 1/MPa/K
kH2 | 1.25e-07 | 1/MPa/K2
p0clathrate | 3633 | MPa
Tclathrate | 1661 | K
xas* | 0.0036
(T/p)(dp/dT) u(T)/T
aSlopeSol0 | 22.6
aSlopeSol1 | -0.072 | 1/K
aSlopeSol2 | 0.00056 | 1/K2
aSlopeLiq0 | 19.8
aSlopeLiq1 | -0.087 | 1/K
aSlopeLiq2 | 0.00022 | 1/K2
u(Tsaturator) | 0.021 | K
u(Tmirror) | 0.1 | K
Tube (from 1/3/2021)
d_in | 0.004 | m
d_out | 0.006 | m
phi | 1100 | 1e-14 mol/(s m Pa)
k_cap | 8.8 | Pa/sccm
Note on AZ3: See report "Humidity from frequency" Eq. (32) and Excel spreadsheet "humidity from frequency".
Note on BK3: Ar permeation out of silicone tube. Value is a free parameter because the effective Ar pressure outside the tube is unknown.
Note on BP3: H2O diffusion along 1 m capillary. See calculation of 17 April 2020. No free parameter.
Note on BQ3: MBW value with zero flow from the humid MFC
Note on AZ4: fit to match MBW values