## Sheet1
Rev. |  |  |  |  | New York State Department Of Environmental Conservation | [25 empty] | Reviewed By:
2014-01-20 |  |  |  |  | TEI MODEL 42C NITRIC OXIDES ANALYZER
Weekly Check Sheet
Station: |  |  | Pinnacle State Park | [11 empty] | Station #: |  |  |  | Pinnacle | [7 empty] | Week Of: |  |  |  | 2020-01-05
Operator: |  |  | John Spicer | [10 empty] | Performed By: |  |  |  |  | John B Spicer | [5 empty] | Day of Week: | [5 empty] | Thursday
|  | Date of Testing: | [5 empty] | 2020-01-09 | [7 empty] | Time: |  |  |  | OOC: |  | 13:44 |  |  | IC: |  | 15:05 |  | hours EST
[12 empty] | Analyzer Information | [10 empty] | ** if Audit or Calibration dates unknown leave blank.
Model #: |  |  |  | 42C | [7 empty] | ENCON #: | [11 empty] | Serial #: | [6 empty] | 63346-339
Last 2 Calibration Dates: | [8 empty] | 2018-12-10 | [5 empty] | 2019-12-12 |  |  |  |  | Last 2 Audit Dates: | [6 empty] | 2019-09-23 |  |  |  |  | 2019-12-12
[12 empty] | Calibrator Information | [11 empty] | ** taken from calibrator certification sheet
Calibrator #: | [5 empty] | T290U | [5 empty] | NO Primary Tank: | [7 empty] | CC408646 |  |  |  |  | NO Concentration, ppb: | [7 empty] | 9900
Calibrator Certification Period: | [11 empty] | 2019-08-21 | [5 empty] | thru |  | 2020-02-27 |  |  |  |  | Total NO2 Impurity, ppb:
Zero and Span Checks
Enter Certified QA | [12 empty] | Instrument Response (taken from Data Logger)
|  | Value Here ↓ | [8 empty] | Zero |  | pre - |  |  | Span |  |  |  |  | % Diff |  |  |  | Zero** |  |  | Post - |  | Date
[13 empty] | Span | [14 empty] | Span
| NO: |  |  | 46.9 |  |  | (ppb) |  |  | -0.02 | [5 empty] | 47.9 |  |  |  |  | 2.1322 |  |  |  | 0.03 |  |  |  |  | 2020-01-09
| NOx: |  |  | 46.8 |  |  | (ppb) |  |  | -0.09 | [5 empty] | 47.8 |  |  |  |  | 2.1368 |  |  |  | 0.02
** If warranted
NO2 Checks
Calibrator # | [5 empty] | GPT | [5 empty] | Calibrator Output | [6 empty] | Instrument Response | [6 empty] | % Diff |  |  |  | Date
[12 empty] | 123.4 | [6 empty] | 115.9 | [6 empty] | -6.0778 |  |  |  | 2020-01-09
Scheduled Maintenance
Procedure | [12 empty] | Frequency | [8 empty] | Last Done | [6 empty] | By:
Replace Particulate Filter | [12 empty] | As Needed | [8 empty] | 2019-12-31 | [6 empty] | JBS
Clean Cooler Fan / Fins | [12 empty] | Quarterly | [8 empty] | 2019-12-31 | [6 empty] | JBS
Check Converter Efficiency | [12 empty] | Every Audit or Cal. | [8 empty] | 2020-01-20 | [6 empty] | JBS
Exhaust Scrubber | [12 empty] | Quarterly | [8 empty] | 2019-11-14 | [6 empty] | JBS
Change Dryer Drierite | [12 empty] | As Needed | [8 empty] | na | [6 empty] | JBS
[12 empty] | Diagnostic Checks and Comments on other side. | [20 empty] | Page 1 of 2
Diagnostic Checks
Range: NO | [5 empty] | 500 |  | 200 | [6 empty] | Voltages | PMT |  |  | volts |  | -1061 |  |  |  | Pressure |  |  |  |  | mmHg |  | 278.6
ppb
Range: NOx | [5 empty] | 500 |  | 200 | [7 empty] | +5 V |  |  | volts |  | 4.6 |  |  |  | Flow |  |  |  |  | lpm |  | 1.048
ppb
Range: NO2 | [5 empty] | 500 |  | 200 | [7 empty] | +15 V |  |  | volts |  | 14.8 |  |  |  | Ozonator Flow |  |  |  |  | lpm |  | 0.085
ppb
Avg. Time (sec) | [5 empty] | 60 |  | 60 | [7 empty] | - 15 V |  |  | volts |  | -14.8 |  |  |  | PRE Background | [6 empty] | 0.69
secs
Alarm On? | [5 empty] | (Y / N) |  | n | [7 empty] | Battery |  |  | volts | [5 empty] | NO Background | [6 empty] | 0.72
Ozonator On? | [5 empty] | (Y / N) |  | y | [6 empty] | Temperatures | Internal |  |  | °C |  | 31.8 |  |  |  | NOX Background | [6 empty] | 0.85
PMT Supply On? | [5 empty] | (Y / N) |  | y | [7 empty] | Chamber |  |  | °C |  | 50.7 |  |  |  | NO Coefficient | [6 empty] | 0.965
Auto Mode On? | [7 empty] | y | [7 empty] | Cooler |  |  | °C |  | -7.2 |  |  |  | NOX Coefficient | [6 empty] | 0.989
Service Mode Off? | [5 empty] | (Y / N) |  | y | [7 empty] | Converter |  |  | °C |  | 316 |  |  |  | NO2 Coefficient | [6 empty] | 1
Temp. Correction On? | [5 empty] | (Y / N) |  | y | [7 empty] | Converter |  |  | °C |  | 0 |  |  |  | Shelter Temp. | [6 empty] | 21.7
Setting
Press. Correct On? | [5 empty] | (Y / N) |  | y | [16 empty] | Shelter BP | [6 empty] | 28.67
Are analyzer diagnostics within tolerances? | [13 empty] | Y / N |  |  | yes |  |  | If no please explain in comments section.
Comments:
Page 2 of 2
Note on K10: note: can press CNTRL + ; simultaneously to enter current date in a cell.
Note on AJ42: Can press CNTRL + ; simultaneously to enter a date in a cell.
Note on AK6: Date entered automatically. Note: week defined as starting on Sunday and ending on Saturday.
Note on AL84: Enter Value in Celsius or Fahrenheit.
Note on AL86: Enter value in mmHg or inHg.