## Slide 1
Sneak peek at VENUS
Vendor-neutral sequences & multi-center reproducibility of quantitative MRI
@agahkarakuzu

## Slide 2
% Difference
30
-30
Lee et al. 2019 MRM 81(1):454-65
vs
Variability in qMRI
“To ensure the comparability of qMRI maps, comparable acquisition techniques should be used in multi-center studies.”
Gracien et al. 2020 Neuroimage 207:116364
Using similar acquisition techniques
https://bit.ly/vfa_inter
T1(s)
0.2
1.8
Teixeira et al. 2019 MRM 81(1):907-20
CSMT
non-CSMT
Reducing variability
Improving the theory aspect
https://bit.ly/ss_ihmt

## Slide 3
% Difference
30
-30
Lee et al. 2019 MRM 81(1):454-65
vs
Variability in qMRI
https://bit.ly/vfa_inter
T1(s)
0.2
1.8
Reducing variability
Implementing consensus
acquisition protocols
Cannot fully remove site-specific user preferences.
Boudreau et al. 2021, in preparation

## Slide 4
Consensus
protocols
Identical
acquisitions
Improving
the theory
😎
T
T
T
Transparency
Technical
Translation
To improve theoretical aspects of a qMRI application, we need access to the details about pulse sequences.
Playing identical
sequences on multiple sites requires a vendor-agnostic solution.
qMRI acquisitions should be user friendly.

## Slide 5
Consensus
protocols
Identical
acquisitions
Improving
the theory
Transparency
Technical
Translation
😎
One-click acquisition UI
RTHAWK
BIDS

## Slide 6
Transparency
Technical
Translation
G1: GE Signa 750
S1: Siemens Prisma
S2: Siemens Skyra
MTsat
with B1+ correction

## Slide 7
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference

## Slide 8
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native

## Slide 9
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native
S2native

## Slide 10
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native
S2native
G1native

## Slide 11
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native
S2native
G1native
S1neutral

## Slide 12
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native
S2native
G1native
S1neutral
S2neutral

## Slide 13
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native
S2native
G1native
S1neutral
S2neutral
G1neutral

## Slide 14
S1
ISMRM-NIST System Phantom
T1 plate
S2
S3
S4
S5
S7
S8
S9
S10
Spherical
ROI (10mm)
T1 reference
sphere (20mm)
1.98s
1.45s
0.98s
0.7s
0.49s
0.35s
0.24s
0.17s
0.12s
0.09s
VFA-VTR (all)
FA-1: 6°
TR-1: 28ms
FA-2: 20°
TR-2: 18ms
1x1x3 mm
TR-1: 25ms
TR-2: 50ms
FA: 60°
Pinc: 117°
∆A/B: 48/240
4x4x4 mm
AFI (rth)
TFL-B1 (S1-2)
Vendor
defaults
were
used.
6x6x8 mm
Reference
S1native
S2native
G1native
S1neutral
S2neutral
G1neutral

## Slide 15
Before B1+ correction
Native
Neutral
T1 (s)
0.5
2.5
G1
S1
S2
N/A
After B1+ correction
G1
S1
S2

## Slide 16
1.17
1.04
1.09
1.28
1.12
1.14
Median

## Slide 17
1.49
1.28
1.35
1.03
1.04
Median

## Slide 18
G1
S1
S2
Native
Neutral
MTR (%)
35
70

## Slide 19
54.3
55.4
54.8
61.6
55.6
55.7
Median

## Slide 20
Before B1+ correction
Native
Neutral
MTsat (s)
1
4.5
G1
S1
S2
After B1+ correction
G1
S1
S2
Rowley et al. 2021, MRM, 86:4, 2192-2207
N/A

## Slide 21
ON
OFF
What do we see
about the MT
pulse on a
vendor UI?
MTC
Here’s what you can know about the MT pulse with our vendor-neutral approach:
Shape: Fermi
Duration: 12ms
B1rms: 3.64 uT
MaxB1: 5uT
Offset: 1.2kHz
Transition: 0.35
∆be4-exc: 1.08ms
Aspoiler: 8 cyc/cm
Exc. Duration: 1.3ms
qMRLab/mt_sat
V1.0.0
Correction files for this protocol are now
available based on, Rowley et al. 2021.

## Slide 22
Consensus
protocols
Identical
acquisitions
Improving
the theory
Transparency
Technical
Translation
😎
One-click acquisition UI
RTHAWK
BIDS
Notes: Going back to one of the most fundamental qMRI application,s T1-mapping, we see several studies published about its reproducibility aspects.
This study by Yoojin and colleaques showed that inter-vendor differences in VFA T1maps can go up to 30% voxelwise, whereas within-vendor results are more acceptable.
To reduce discrepancies, we can improve the theory aspect of the acquisition by leveling the total amount of RF energy deposited to our sample at different flip angles so that the MT saturation is controlled. It is known to improve the accuracy & precision of T1 estimation.
Another solution is that we can make our sequences as similar as possible across multiple centers. Because in reality we often don’t know the implementation differences between stock sequences such as FLASH and SPGR, which obviously has an effect on our estimations.

## Slide 23
Quantitative MRI.
Under one umbrella.
qmrlab.org

## Slide 24
www.neuro.polymtl.ca
THANKS!
agahkarakuzu@gmail.com
@agahkarakuzu
Nikola Stikov (Advisor)
Julien Cohen-Adad
Ilana Leppert
Mathieu Boudreau
Jennifer Campbell
Agah Karakuzu
Tommy Boshkovski
Tanguy Duval
Bruce Pike
Sridar Narayanan
Christine Tardif
Robert Brown
David Rudko
Jean-Francois Cabana
Robert Dougherty
Brian Wandell
Notes: I would like to thank TransMedTech, CONP, Montreal Heart Institute, ISMRM, IVADO and Quebec Bio-Imaging Network for funding my work and to all my colleagues at NeuroPoly research lab. & finally to ISMRM Turkish chapter for organizing this workshop.