## U_test
|  | BP110 group median [P10, P90] | BAHA5P group median [P10, P90] | N (BP110 / BAHA5P) | Mann-Whitney U Test statistic | p – value (two-sided) | Analysis Results File | Plot | Analysis Code File
General characteristics | Count of M , F |  |  | 27, 22 / 14, 22 | * | * 0.188 | analysis_sex.xlsx | - | Analysis_Age_&_Sex
| Age at BCD fitting [years] | 54.0 [41.6, 69.6] | 52.4 [30.9, 67.9] | 49 / 36 | 958 | 0.499 | analysis_age.xlsx | - | Analysis_Age_&_Sex
Hearing Loss Best Ear (AC) | Hearing Loss Air Conduction at 0.5 kHz [dB HL] | 10 [5 , 25] | 10 [5, 20] | 49 / 36 | 842 | 0.898 | analysis_AC_thresholds.xlsx | AC_thresholds.tiff | Analysis_AC_ Thresholds
| Hearing Loss Air Conduction at 1 kHz [dB HL] | 10 [5 , 20] | 10 [0 , 20] | 49 / 36 | 852 | 0.781 | analysis_AC_thresholds.xlsx | AC_thresholds.tiff | Analysis_AC_ Thresholds
| Hearing Loss Air Conduction at 2 kHz [dB HL] | 10 [5, 25] | 10 [5, 20] | 49 / 36 | 739 | 0.194 | analysis_AC_thresholds.xlsx | AC_thresholds.tiff | Analysis_AC_ Thresholds
| Hearing Loss Air Conduction at 4 kHz [dB HL] | 25 [10, 55] | 20 [8, 43] | 49 / 36 | 693 | 0.091 | analysis_AC_thresholds.xlsx | AC_thresholds.tiff | Analysis_AC_ Thresholds
| Air Conduction PTA at 4 frequencies (0.5-4 kHz) [dB HL] | 16 [8, 25] | 13 [6, 24] | 49 / 36 | 655 | 0.043 | analysis_AC_thresholds.xlsx | AC_thresholds.tiff | Analysis_AC_ Thresholds
| Air Conduction PTA at 3 frequencies (1-4 kHz) [dB HL] | 17 [7, 28] | 13 [7, 26] | 49 / 36 | 644 | 0.034 | analysis_AC_thresholds.xlsx | AC_thresholds.tiff | Analysis_AC_ Thresholds
Bone Conduction direct Thresholds | Bone Conduction Direct threshold at 0.5 kHz [dB ref BC direct] | 25 [15, 37] | 33 [20, 43] | 47 / 36 | 1103 | 0.016 | analysis_BCdirect_thresholds.xlsx | BC_direct_thresholds.tiff | Analysis_BC_direct_thresholds
| Bone Conduction Direct threshold at 1 kHz [dB ref BC direct] | 25 [15, 40] | 23 [18, 33] | 47 / 36 | 803 | 0.682 | analysis_BCdirect_thresholds.xlsx | BC_direct_thresholds.tiff | Analysis_BC_direct_thresholds
| Bone Conduction Direct threshold at 2 kHz [dB ref BC direct] | 30 [23, 45] | 30 [20, 43] | 47 / 36 | 686 | 0.134 | analysis_BCdirect_thresholds.xlsx | BC_direct_thresholds.tiff | Analysis_BC_direct_thresholds
| Bone Conduction Direct threshold at 4 kHz [dB ref BC direct] | 40 [28, 67] | 35 [25, 53] | 47 / 36 | 677 | 0.115 | analysis_BCdirect_thresholds.xlsx | BC_direct_thresholds.tiff | Analysis_BC_direct_thresholds
| Bone Conduction Direct PTA at 4 frequencies (0.5-4 kHz) [dB ref BC direct] | 33 [26, 39] | 30 [24, 41] | 47 / 36 | 690.5 | 0.167 | analysis_BCdirect_thresholds.xlsx | BC_direct_thresholds.tiff | Analysis_BC_direct_thresholds
| Bone Conduction Direct PTA at 3 frequencies (1-4 kHz) [dB ref BC direct] | 34 [26, 40] | 30 [24, 41] | 47 / 36 | 621 | 0.039 | analysis_BCdirect_thresholds.xlsx | BC_direct_thresholds.tiff | Analysis_BC_direct_thresholds
Daily use BCD [hours] | Daily average use of BCD | 0.7 [0.1, 5] | 1.5 [0.1, 4.8] | 47 / 31 | 632 | 0.321 | analysis_usage_BCD.xlsx | usage_BCD.tiff | Analysis_Usage_BCD
Output of the BCD on skull simulator | Output for ISTS 65 dB on skull simulator at 0.5 kHz [dB ref 1 μN] for first-fit settings | 88 [84, 90] | 81 [80, 84] | 49 / 36 | 85 | 0.000 | analysis_output_first_fit_65.xlsx | output_bp_vs_bh5, 65dB.tiff | Analysis_output_BP110_vs_BAHA5P_65dB
| Output for ISTS 65 dB on skull simulator at 1 kHz [dB ref 1 μN] for first-fit settings | 91 [89, 93] | 85 [80, 87] | 49 / 36 | 39 | 0.000 | analysis_output_first_fit_65.xlsx | output_bp_vs_bh5, 65dB.tiff | Analysis_output_BP110_vs_BAHA5P_65dB
| Output for ISTS 65 dB on skull simulator at 2 kHz [dB ref 1 μN] for first-fit settings | 82 [80, 84] | 86 [80, 90] | 49 / 36 | 1435 | 0.000 | analysis_output_first_fit_65.xlsx | output_bp_vs_bh5, 65dB.tiff | Analysis_output_BP110_vs_BAHA5P_65dB
| Output for ISTS 65 dB on skull simulator at 4 kHz [dB ref 1 μN] for first-fit settings | 74 [70, 75] | 80 [75, 83] | 49 / 36 | 1625 | 0.000 | analysis_output_first_fit_65.xlsx | output_bp_vs_bh5, 65dB.tiff | Analysis_output_BP110_vs_BAHA5P_65dB
Compression Ratio | Compression Ratio at 0.5 kHz | 1.05 [0.99, 1.16] | 0.95 [0.9, 1.08] | 49 / 36 | 225 | 0.000 | analysis_compressio_ratio.xlsx | compression_ratio.tiff | Analysis_compression_ratio
| Compression Ratio at 1 kHz | 1.03 [1.01, 1.1] | 0.95 [0.92, 1.00] | 49 / 36 | 70 | 0.000 | analysis_compressio_ratio.xlsx | compression_ratio.tiff | Analysis_compression_ratio
| Compression Ratio at 2 kHz | 0.96 [0.94, 0.99] | 1.00 [0.95, 1.04] | 49 / 36 | 1311 | 0.000 | analysis_compressio_ratio.xlsx | compression_ratio.tiff | Analysis_compression_ratio
| Compression Ratio at 4 kHz | 0.93 [0.91, 0.95] | 0.92 [0.89, 0.93] | 49 / 36 | 509 | 0.000 | analysis_compressio_ratio.xlsx | compression_ratio.tiff | Analysis_compression_ratio
Signal in front, no noise | Difference Sensation Level Bone Conduction versus Air Conduction at 0.5 kHz [dB] | -33 [-48, -23] | -45 [-55, -34] | 35 / 35 | 304 | 0.000 | analysis_difference_sensation_level_65dB_in_front.xlsx | sl_S65_in_front.tiff | Analysis_sensation_level_difference_S65dB_in_front
| Difference Sensation Level Bone Conduction versus Air Conduction at 1 kHz [dB] | -13 [-25, -3] | -20 [-30, -13] | 35 / 35 | 300 | 0.000 | analysis_difference_sensation_level_65dB_in_front.xlsx | sl_S65_in_front.tiff | Analysis_sensation_level_difference_S65dB_in_front
| Difference Sensation Level Bone Conduction versus Air Conduction at 2 kHz [dB] | -19 [-28, -7] | -12 [-22, -3] | 35 / 35 | 831 | 0.011 | analysis_difference_sensation_level_65dB_in_front.xlsx | sl_S65_in_front.tiff | Analysis_sensation_level_difference_S65dB_in_front
| Difference Sensation Level Bone Conduction versus Air Conduction at 4 kHz [dB] | -18 [-35, -8] | -14 [-25, 4] | 35 / 35 | 822 | 0.014 | analysis_difference_sensation_level_65dB_in_front.xlsx | sl_S65_in_front.tiff | Analysis_sensation_level_difference_S65dB_in_front
| SII AC_path | 0.992 [0.965, 1] | 1 [0,968, 1] | 35 / 35 | 424 | 0.022 | analysis_SII_Sfront65_Silence | BP110_SII_Sfront65_Silence.tiff & BAHA5P_SII_Sfront65_Silence.tiff | Analysis_SII_Sfront65_silence
| SII BC_path | 0.870 [0.820, 0.905] | 0.834 [0.789, 0.878] | 35 / 35 | 897 | 0.001 | analysis_SII_Sfront65_Silence.xlsx | BP110_SII_Sfront65_Silence.tiff & BAHA5P_SII_Sfront65_Silence.tiff | Analysis_SII_Sfront65_silence
| SII AC&BC_path | 0.992 [0.965, 1] | 1 [0,980, 1] | 35 / 35 | 409 | 0.014 | analysis_SII_Sfront65_Silence.xlsx | BP110_SII_Sfront65_Silence.tiff & BAHA5P_SII_Sfront65_Silence.tiff | Analysis_SII_Sfront65_silence
| Difference between SII Combination path and SII Air Conduction path | 0.000 [0.000, 0.007] | 0.000 [0.000, 0.006] | 35 / 35 | 670 | 0.334 | analysis_SII_Sfront65_Silence.xlsx | BP110_SII_Sfront65_Silence.tiff & BAHA5P_SII_Sfront65_Silence.tiff | Analysis_SII_Sfront65_silence
| CVC score AC_path | 0.98 [0.98, 0.98] | 0.98 [0.98, 0.98] | 35 / 35 | 595 | 0.317 | analysis_CVC_Sfront65_Silence.xlsx | BP110_CVC_Sfront65_Silence.tiff & BAHA5P_CVC_Sfront65_Silence.tiff | Analysis_CVC_Sfront65_silence
| CVC score BC_path | 0.97 [0.96, 0.97] | 0.96 [0.96, 0.97] | 35 / 35 | 807 | 0.011 | analysis_CVC_Sfront65_Silence.xlsx | BP110_CVC_Sfront65_Silence.tiff & BAHA5P_CVC_Sfront65_Silence.tiff | Analysis_CVC_Sfront65_silence
| CVC score AC&BC_path | 0.98 [0.98, 0.98] | 0.98 [0.98, 0.98] | 35 / 35 | 595 | 0.317 | analysis_CVC_Sfront65_Silence.xlsx | BP110_CVC_Sfront65_Silence.tiff & BAHA5P_CVC_Sfront65_Silence.tiff | Analysis_CVC_Sfront65_silence
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [0.00, 0.00] | 0.00 [0.00, 0.00] | 35 / 35 | † | † | analysis_CVC_Sfront65_Silence.xlsx | BP110_CVC_Sfront65_Silence.tiff & BAHA5P_CVC_Sfront65_Silence.tiff | Analysis_CVC_Sfront65_silence
Signal at best ear side, no noise | Difference Sensation Level Bone Conduction versus Air Conduction at 0.5 kHz [dB] | -37 [-52, -26] | -49 [ -58, -38] | 35 / 35 | 303 | 0.000 | analysis_difference_sensation_level_65dB_best_ear.xlsx | sl_S65dB_best_ear_side.tiff | Analysis_sensation_level_difference_S65dB_best_ear_side
| Difference Sensation Level Bone Conduction versus Air Conduction at 1 kHz [dB] | -21 [-32, -11] | -28 [-38, -21] | 35 / 35 | 302 | 0.000 | analysis_difference_sensation_level_65dB_best_ear.xlsx | sl_S65dB_best_ear_side.tiff | Analysis_sensation_level_difference_S65dB_best_ear_side
| Difference Sensation Level Bone Conduction versus Air Conduction at 2 kHz [dB] | -29 [-38, -16] | -22 [-32, -12] | 35 / 35 | 825 | 0.013 | analysis_difference_sensation_level_65dB_best_ear.xlsx | sl_S65dB_best_ear_side.tiff | Analysis_sensation_level_difference_S65dB_best_ear_side
| Difference Sensation Level Bone Conduction versus Air Conduction at 4 kHz [dB] | -25 [-42, -15] | -21 [-32, -3] | 35 / 35 | 819 | 0.015 | analysis_difference_sensation_level_65dB_best_ear.xlsx | sl_S65dB_best_ear_side.tiff | Analysis_sensation_level_difference_S65dB_best_ear_side
| SII AC_path | 0.996 [0.973, 1] | 1 [0.976, 1] | 35 / 35 | 415 | 0.009 | analysis_SII_Sbest65_Silence.xlsx | BP110_SII_Sbest65_Silence.tiff & BAHA5P_SII_Sbest65_Silence.tiff | Analysis_SII_Sbest65_silence
| SII BC_path | 0.847, [0.791, 0.890] | 0.814 [0.765, 0.869] | 35 / 35 | 828 | 0.011 | analysis_SII_Sbest65_Silence.xlsx | BP110_SII_Sbest65_Silence.tiff & BAHA5P_SII_Sbest65_Silence.tiff | Analysis_SII_Sbest65_silence
| SII AC&BC_path | 0.996 [0.973, 1] | 1 [0.978, 1] | 35 / 35 | 414 | 0.009 | analysis_SII_Sbest65_Silence.xlsx | BP110_SII_Sbest65_Silence.tiff & BAHA5P_SII_Sbest65_Silence.tiff | Analysis_SII_Sbest65_silence
| Difference between SII Combination path and SII Air Conduction path | 0.000 [0.000, 0.000] | 0.000 [0.000, 0.000] | 35 / 35 | 578 | 0.154 | analysis_SII_Sbest65_Silence.xlsx | BP110_SII_Sbest65_Silence.tiff & BAHA5P_SII_Sbest65_Silence.tiff | Analysis_SII_Sbest65_silence
| CVC score AC_path | 0.98 [0,98, 0,98] | 0.98 [0,98, 0,98] | 35 / 35 | 595 | 0.317 | analysis_CVC_Sbest65_Silence.xlsx | BP110_CVC_Sbest65_Silence.tiff & BAHA5P_CVC_Sbest65_Silence.tiff | Analysis_CVC_Sbest65_silence
| CVC score BC_path | 0.97 [0.96, 0.97] | 0.96 [0.95, 0.97] | 35 / 35 | 815 | 0.009 | analysis_CVC_Sbest65_Silence.xlsx | BP110_CVC_Sbest65_Silence.tiff & BAHA5P_CVC_Sbest65_Silence.tiff | Analysis_CVC_Sbest65_silence
| CVC score AC&BC_path | 0.98 [0,98, 0,98] | 0.98 [0,98, 0,98] | 35 / 35 | 595 | 0.317 | analysis_CVC_Sbest65_Silence.xlsx | BP110_CVC_Sbest65_Silence.tiff & BAHA5P_CVC_Sbest65_Silence.tiff | Analysis_CVC_Sbest65_silence
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [0.00, 0.00] | 0.00 [0.00, 0.00] | 35 / 35 | † | † | analysis_CVC_Sbest65_Silence.xlsx | BP110_CVC_Sbest65_Silence.tiff & BAHA5P_CVC_Sbest65_Silence.tiff | Analysis_CVC_Sbest65_silence
Signal at BCD side, no noise | Difference Sensation Level Bone Conduction versus Air Conduction at 0.5 kHz [dB] | -27 [-42, -16] | -39 [-48, -28] | 35 / 35 | 302 | 0.000 | analysis_difference_sensation_level_S65dB_BCDside.xlsx | sl_S65_BCDside.tiff | Analysis_sensation_level_difference_S65dB_BCD_side
| Difference Sensation Level Bone Conduction versus Air Conduction at 1 kHz [dB] | -6 [-18, 4] | -13 [-23, -6] | 35 / 35 | 300 | 0.000 | analysis_difference_sensation_level_S65dB_BCDside.xlsx | sl_S65_BCDside.tiff | Analysis_sensation_level_difference_S65dB_BCD_side
| Difference Sensation Level Bone Conduction versus Air Conduction at 2 kHz [dB] | -12 [-21, 1] | -5 [-15, 4] | 35 / 35 | 829 | 0.011 | analysis_difference_sensation_level_S65dB_BCDside.xlsx | sl_S65_BCDside.tiff | Analysis_sensation_level_difference_S65dB_BCD_side
| Difference Sensation Level Bone Conduction versus Air Conduction at 4 kHz [dB] | -13 [-30, -3] | -9 [-20, 9] | 35 / 35 | 821 | 0.014 | analysis_difference_sensation_level_S65dB_BCDside.xlsx | sl_S65_BCDside.tiff | Analysis_sensation_level_difference_S65dB_BCD_side
| SII AC_path | 0.955 [0.903, 0.968] | 0.966 [0.929, 0.968] | 35 / 35 | 426 | 0.024 | analysis_SII_Sbcd65_Silence.xlsx | BP110_SII_Sbcd65_Silence.tiff & BAHA5P_SII_Sbcd65_Silence.tiff | Analysis_SII_Sbcd65_silence
| SII BC_path | 0.897 [0.830, 0.924] | 0.859 [0.809, 0.889] | 35 / 35 | 870 | 0.002 | analysis_SII_Sbcd65_Silence.xlsx | BP110_SII_Sbcd65_Silence.tiff & BAHA5P_SII_Sbcd65_Silence.tiff | Analysis_SII_Sbcd65_silence
| SII AC&BC_path | 0.968 [0.920, 0,980] | 0.971 [0.948, 0,993] | 35 / 35 | 361 | 0.003 | analysis_SII_Sbcd65_Silence.xlsx | BP110_SII_Sbcd65_Silence.tiff & BAHA5P_SII_Sbcd65_Silence.tiff | Analysis_SII_Sbcd65_silence
| Difference between SII Combination path and SII Air Conduction path | 0.008 [0.000, 0.039] | 0.008 [0.000, 0.039] | 35 / 35 | 538 | 0.380 | analysis_SII_Sbcd65_Silence.xlsx | BP110_SII_Sbcd65_Silence.tiff & BAHA5P_SII_Sbcd65_Silence.tiff | Analysis_SII_Sbcd65_silence
| CVC score AC_path | 0.98 [0.97, 0.98] | 0.98 [0.97, 0.98] | 35 / 35 | 543 | 0.235 | analysis_CVC_Sbcd65_Silence.xlsx | BP110_CVC_Sbcd65_Silence.tiff & BAHA5P_CVC_Sbcd65_Silence.tiff | Analysis_CVC_Sbcd65_silence
| CVC score BC_path | 0.97 [0.96, 0.97] | 0.97 [0.96, 0.97] | 35 / 35 | 749 | 0.052 | analysis_CVC_Sbcd65_Silence.xlsx | BP110_CVC_Sbcd65_Silence.tiff & BAHA5P_CVC_Sbcd65_Silence.tiff | Analysis_CVC_Sbcd65_silence
| CVC score AC&BC_path | 0.98 [0.97, 0.98] | 0.98 [0.98, 0.98] | 35 / 35 | 490 | 0.006 | analysis_CVC_Sbcd65_Silence.xlsx | BP110_CVC_Sbcd65_Silence.tiff & BAHA5P_CVC_Sbcd65_Silence.tiff | Analysis_CVC_Sbcd65_silence
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [0.00, 0.00] | 0.00 [0.00, 0.01] | 35 / 35 | 560 | 0.235 | analysis_CVC_Sbcd65_Silence.xlsx | BP110_CVC_Sbcd65_Silence.tiff & BAHA5P_CVC_Sbcd65_Silence.tiff | Analysis_CVC_Sbcd65_silence
Signal in front, noise in front | SII Air Conduction path | 0,488 [0.476, 0.488] | 0,488 [0.476, 0.488] | 35 / 35 | 458 | 0.027 | analysis_SII_Sfront65_Nfront65.xlsx | BP110_SII_Sfront65_Nfront65.tiff & BAHA5P_SII_Sfront65_Nfront65.tiff | Analysis_SII_Sfront65_Nfront65
| SII Bone Conduction path | 0,254 [0.218, 0.307] | 0,372 [0.340, 0.391] | 35 / 35 | 10 | 0.000 | analysis_SII_Sfront65_Nfront65.xlsx | BP110_SII_Sfront65_Nfront65.tiff & BAHA5P_SII_Sfront65_Nfront65.tiff | Analysis_SII_Sfront65_Nfront66
| SII Combination path | 0,488 [0.476, 0.488] | 0,489 [0.481, 0.490] | 35 / 35 | 472 | 0.069 | analysis_SII_Sfront65_Nfront65.xlsx | BP110_SII_Sfront65_Nfront65.tiff & BAHA5P_SII_Sfront65_Nfront65.tiff | Analysis_SII_Sfront65_Nfront67
| Difference between SII Combination path and SII Air Conduction path | 0,000 [0.000, 0.003] | 0,000 [0.000, 0.003] | 35 / 35 | 588 | 0.727 | analysis_SII_Sfront65_Nfront65.xlsx | BP110_SII_Sfront65_Nfront65.tiff & BAHA5P_SII_Sfront65_Nfront65.tiff | Analysis_SII_Sfront65_Nfront68
| CVC score of the Air Conduction path | 0,86 [0.85, 0.86] | 0,86 [0.85, 0.86] | 35 / 35 | 595 | 0.744 | analysis_CVC_Sfront65_Nfront65.xlsx | BP110_CVC_Sfront65_Nfront65.tiff & BAHA5P_CVC_Sfront65_Nfront65.tiff | Analysis_CVC_Sfront65_Nfront65
| CVC score of the Bone Conduction path | 0,65 [0.60, 0.72] | 0.78 [0.75, 0.80] | 35 / 35 | 10 | 0.000 | analysis_CVC_Sfront65_Nfront65.xlsx | BP110_CVC_Sfront65_Nfront65.tiff & BAHA5P_CVC_Sfront65_Nfront65.tiff | Analysis_CVC_Sfront65_Nfront66
| CVC score of the Combination path | 0.86 [0.85, 0.86] | 0.86 [0.86, 0.86] | 35 / 35 | 543 | 0,136 | analysis_CVC_Sfront65_Nfront65.xlsx | BP110_CVC_Sfront65_Nfront65.tiff & BAHA5P_CVC_Sfront65_Nfront65.tiff | Analysis_CVC_Sfront65_Nfront67
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [0.00, 0.00] | 0.00 [0.00, 0.00] | 35 / 35 | 560 | 0.079 | analysis_CVC_Sfront65_Nfront65.xlsx | BP110_CVC_Sfront65_Nfront65.tiff & BAHA5P_CVC_Sfront65_Nfront65.tiff | Analysis_CVC_Sfront65_Nfront68
Signal in front, noise BCD side | SII Air Conduction path | 0.651 [0.626, 0.651] | 0.651 [0.627, 0.651] | 35 / 35 | 439 | 0.014 | analysis_SII_Sfront65_Nbcd65.xlsx | BP110_SII_Sfront65_Nbcd65.tiff & BAHA5P_SII_Sfront65_Nbcd65.tiff | Analysis_SII_Sfront65_Nbcd65
| SII Bone Conduction path | 0.115 [0.102, 0.138] | 0.247 [0.205, 0.281] | 35 / 35 | 0 | 0.000 | analysis_SII_Sfront65_Nbcd65.xlsx | BP110_SII_Sfront65_Nbcd65.tiff & BAHA5P_SII_Sfront65_Nbcd65.tiff | Analysis_SII_Sfront65_Nbcd66
| SII Combination path | 0.627 [0.572, 0.651] | 0.631 [0.587, 0.650] | 35 / 35 | 623 | 0.906 | analysis_SII_Sfront65_Nbcd65.xlsx | BP110_SII_Sfront65_Nbcd65.tiff & BAHA5P_SII_Sfront65_Nbcd65.tiff | Analysis_SII_Sfront65_Nbcd67
| Difference between SII Combination path and SII Air Conduction path | -0.007 [-0.060, 0.000] | -0.018 [-0.057, -0.001] | 35 / 35 | 743 | 0.126 | analysis_SII_Sfront65_Nbcd65.xlsx | BP110_SII_Sfront65_Nbcd65.tiff & BAHA5P_SII_Sfront65_Nbcd65.tiff | Analysis_SII_Sfront65_Nbcd68
| CVC score of the Air Conduction path | 0.93 [0.92, 0.93] | 0.93 [0.92, 0.93] | 35 / 35 | 427 | 0.008 | analysis_CVC_Sfront65_Nbcd65.xlsx | BP110_CVC_Sfront65_Nbcd65.tiff & BAHA5P_CVC_Sfront65_Nbcd65.tiff | Analysis_CVC_Sfront65_Nbcd65
| CVC score of the Bone Conduction path | 0.39 [0.36, 0.45] | 0,64 [0.57, 0.69] | 35 / 35 | 0 | 0.000 | analysis_CVC_Sfront65_Nbcd65.xlsx | BP110_CVC_Sfront65_Nbcd65.tiff & BAHA5P_CVC_Sfront65_Nbcd65.tiff | Analysis_CVC_Sfront65_Nbcd66
| CVC score of the Combination path | 0.92 [0.90, 0.93] | 0.92 [0.90, 0.93] | 35 / 35 | 606 | 0.930 | analysis_CVC_Sfront65_Nbcd65.xlsx | BP110_CVC_Sfront65_Nbcd65.tiff & BAHA5P_CVC_Sfront65_Nbcd65.tiff | Analysis_CVC_Sfront65_Nbcd67
| Difference between CVC score Combination path and CVC score Air Conduction path | -0.01 [-0.02, 0.00] | -0.01 [-0.02, 0.00] | 35 / 35 | 771 | 0.045 | analysis_CVC_Sfront65_Nbcd65.xlsx | BP110_CVC_Sfront65_Nbcd65.tiff & BAHA5P_CVC_Sfront65_Nbcd65.tiff | Analysis_CVC_Sfront65_Nbcd68
Signal in front, noise best ear side | SII Air Conduction path | 0.384 [0.383, 0.384] | 0.384 [0.383, 0.384] | 35 / 35 | 516 | 0.122 | analysis_SII_Sfront65_Nbest65.xlsx | BP110_SII_Sfront65_Nbest65.tiff & BAHA5P_SII_Sfront65_Nbest65.tiff | Analysis_SII_Sfront65_Nbest65
| SII Bone Conduction path | 0.356 [0.356, 0.318] | 0.520 [0.496. 0.547] | 35 / 35 | 1 | 0.000 | analysis_SII_Sfront65_Nbest65.xlsx | BP110_SII_Sfront65_Nbest65.tiff & BAHA5P_SII_Sfront65_Nbest65.tiff | Analysis_SII_Sfront65_Nbest66
| SII Combination path | 0.384 [0.383, 0.398] | 0.384 [0.384, 0.401] | 35 / 35 | 566 | 0.540 | analysis_SII_Sfront65_Nbest65.xlsx | BP110_SII_Sfront65_Nbest65.tiff & BAHA5P_SII_Sfront65_Nbest65.tiff | Analysis_SII_Sfront65_Nbest67
| Difference between SII Combination path and SII Air Conduction path | 0.000 [0.000, 0.015] | 0.000 [0.000, 0.017] | 35 / 35 | 622 | 0.896 | analysis_SII_Sfront65_Nbest65.xlsx | BP110_SII_Sfront65_Nbest65.tiff & BAHA5P_SII_Sfront65_Nbest65.tiff | Analysis_SII_Sfront65_Nbest68
| CVC score of the Air Conduction path | 0.79 [0.79, 0.79] | 0.79 [0.79, 0.79] | 35 / 35 | † | † | analysis_CVC_Sfront65_Nbest65.xlsx | BP110_CVC_Sfront65_Nbest65.tiff & BAHA5P_CVC_Sfront65_Nbest65.tiff | Analysis_CVC_Sfront65_Nbest65
| CVC score of the Bone Conduction path | 0.77 [0.73, 0.83] | 0.88 [0.86, 0.89] | 35 / 35 | 2 | 0.000 | analysis_CVC_Sfront65_Nbest65.xlsx | BP110_CVC_Sfront65_Nbest65.tiff & BAHA5P_CVC_Sfront65_Nbest65.tiff | Analysis_CVC_Sfront65_Nbest66
| CVC score of the Combination path | 0.79 [0.79, 0.80] | 0.79 [0.79, 0.81] | 35 / 35 | 591 | 0.736 | analysis_CVC_Sfront65_Nbest65.xlsx | BP110_CVC_Sfront65_Nbest65.tiff & BAHA5P_CVC_Sfront65_Nbest65.tiff | Analysis_CVC_Sfront65_Nbest67
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [0.00, 0.01] | 0.00 [0.00, 0.02] | 35 / 35 | 591 | 0.736 | analysis_CVC_Sfront65_Nbest65.xlsx | BP110_CVC_Sfront65_Nbest65.tiff & BAHA5P_CVC_Sfront65_Nbest65.tiff | Analysis_CVC_Sfront65_Nbest68
Signal BCD side, noise best ear side | SII Air Conduction path | 0.230 [0.230, 0.230] | 0.230 [0.230, 0.230] | 35 / 35 | † | † | analysis_SII_Sbcd65_Nbest65.xlsx | BP110_SII_Sbcd65_Nbest65.tiff & BAHA5P_SII_Sbcd65_Nbest65.tiff | Analysis_SII_Sbcd65_Nbest65
| SII Bone Conduction path | 0.432 [0.391, 0.502] | 0.590 [0.564, 0.622] | 35 / 35 | 5 | 0.000 | analysis_SII_Sbcd65_Nbest65.xlsx | BP110_SII_Sbcd65_Nbest65.tiff & BAHA5P_SII_Sbcd65_Nbest65.tiff | Analysis_SII_Sbcd65_Nbest66
| SII Combination path | 0.239 [0.230, 0.294] | 0.246 [0.230, 0.294] | 35 / 35 | 569 | 0.608 | analysis_SII_Sbcd65_Nbest65.xlsx | BP110_SII_Sbcd65_Nbest65.tiff & BAHA5P_SII_Sbcd65_Nbest65.tiff | Analysis_SII_Sbcd65_Nbest67
| Difference between SII Combination path and SII Air Conduction path | 0.009 [0.000, 0.064] | 0.016 [0.000, 0.064] | 35 / 35 | 569 | 0.608 | analysis_SII_Sbcd65_Nbest65.xlsx | BP110_SII_Sbcd65_Nbest65.tiff & BAHA5P_SII_Sbcd65_Nbest65.tiff | Analysis_SII_Sbcd65_Nbest68
| CVC score of the Air Conduction path | 0.62 [0.62, 0.62] | 0.62 [0.62, 0.62] | 35 / 35 | † | † | analysis_CVC_Sbcd65_Nbest65.xlsx | BP110_CVC_Sbcd65_Nbest65.tiff & BAHA5P_CVC_Sbcd65_Nbest65.tiff | Analysis_CVC_Sbcd65_Nbest65
| CVC score of the Bone Conduction path | 0.83 [0.80, 0.87] | 0.91 [0.90, 0.92] | 35 / 35 | 8 | 0.000 | analysis_CVC_Sbcd65_Nbest65.xlsx | BP110_CVC_Sbcd65_Nbest65.tiff & BAHA5P_CVC_Sbcd65_Nbest65.tiff | Analysis_CVC_Sbcd65_Nbest66
| CVC score of the Combination path | 0.63 [0.62, 0.70] | 0.64 [0.62, 0.71] | 35 / 35 | 530 | 0.318 | analysis_CVC_Sbcd65_Nbest65.xlsx | BP110_CVC_Sbcd65_Nbest65.tiff & BAHA5P_CVC_Sbcd65_Nbest65.tiff | Analysis_CVC_Sbcd65_Nbest67
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.01 [0.00, 0.08] | 0.02 [0.00, 0.09] | 35 / 35 | 530 | 0.318 | analysis_CVC_Sbcd65_Nbest65.xlsx | BP110_CVC_Sbcd65_Nbest65.tiff & BAHA5P_CVC_Sbcd65_Nbest65.tiff | Analysis_CVC_Sbcd65_Nbest68
Signal BCD side, noise in front | SII Air Conduction path | 0.325 [0.319, 0.325] | 0.325 [0.319, 0.325] | 35 / 35 | 447 | 0,017 | analysis_SII_Sbcd65_Nfront65.xlsx | BP110_SII_Sbcd65_Nfront65.tiff & BAHA5P_SII_Sbcd65_Nfront65.tiff | Analysis_SII_Sbcd65_Nfront65
| SII Bone Conduction path | 0.334 [0.300, 0.379] | 0.444 [0.406, 0.464] | 35 / 35 | 16 | 0.000 | analysis_SII_Sbcd65_Nfront65.xlsx | BP110_SII_Sbcd65_Nfront65.tiff & BAHA5P_SII_Sbcd65_Nfront65.tiff | Analysis_SII_Sbcd65_Nfront66
| SII Combination path | 0.333 [0.325, 0.377] | 0.347 [0.326, 0.385] | 35 / 35 | 498 | 0.176 | analysis_SII_Sbcd65_Nfront65.xlsx | BP110_SII_Sbcd65_Nfront65.tiff & BAHA5P_SII_Sbcd65_Nfront65.tiff | Analysis_SII_Sbcd65_Nfront67
| Difference between SII Combination path and SII Air Conduction path | 0.010 [0.000, 0.053] | 0.022 [0.001, 0.060] | 35 / 35 | 525 | 0.304 | analysis_SII_Sbcd65_Nfront65.xlsx | BP110_SII_Sbcd65_Nfront65.tiff & BAHA5P_SII_Sbcd65_Nfront65.tiff | Analysis_SII_Sbcd65_Nfront68
| CVC score of the Air Conduction path | 0.74 [0.73, 0.74] | 0.74 [0.73, 0.74] | 35 / 35 | 490 | 0,053 | analysis_CVC_Sbcd65_Nfront65.xlsx | BP110_CVC_Sbcd65_Nfront65.tiff & BAHA5P_CVC_Sbcd65_Nfront65.tiff | Analysis_CVC_Sbcd65_Nfront65
| CVC score of the Bone Conduction path | 0.75 [0.71, 0.79] | 0.81 [0.83, 0.85] | 35 / 35 | 19 | 0.000 | analysis_CVC_Sbcd65_Nfront65.xlsx | BP110_CVC_Sbcd65_Nfront65.tiff & BAHA5P_CVC_Sbcd65_Nfront65.tiff | Analysis_CVC_Sbcd65_Nfront66
| CVC score of the Combination path | 0.75 [0.74, 0.79] | 0.76 [0.74, 0.79] | 35 / 35 | 493 | 0.146 | analysis_CVC_Sbcd65_Nfront65.xlsx | BP110_CVC_Sbcd65_Nfront65.tiff & BAHA5P_CVC_Sbcd65_Nfront65.tiff | Analysis_CVC_Sbcd65_Nfront67
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.01 [0.00, 0.05] | 0.02 [0.00, 0.05] | 35 / 35 | 539 | 0.372 | analysis_CVC_Sbcd65_Nfront65.xlsx | BP110_CVC_Sbcd65_Nfront65.tiff & BAHA5P_CVC_Sbcd65_Nfront65.tiff | Analysis_CVC_Sbcd65_Nfront68
Signal best ear side, noise BCD side | SII Air Conduction path | 0.750 [0.725, 0.750] | 0.750 [0.725, 0.750] | 35 / 35 | 439 | 0.015 | analysis_SII_Sbest65_Nbcd65.xlsx | BP110_SII_Sbest65_Nbcd65.tiff & BAHA5P_SII_Sbest65_Nbcd65.tiff | Analysis_SII_Sbest65_Nbcd65
| SII Bone Conduction path | 0.095 [0.084, 0.104] | 0.117 [0.082, 0.141] | 35 / 35 | 248 | 0.000 | analysis_SII_Sbest65_Nbcd65.xlsx | BP110_SII_Sbest65_Nbcd65.tiff & BAHA5P_SII_Sbest65_Nbcd65.tiff | Analysis_SII_Sbest65_Nbcd66
| SII Combination path | 0.725 [0.666, 0.750] | 0.734 [0.679, 0.750] | 35 / 35 | 608 | 0.958 | analysis_SII_Sbest65_Nbcd65.xlsx | BP110_SII_Sbest65_Nbcd65.tiff & BAHA5P_SII_Sbest65_Nbcd65.tiff | Analysis_SII_Sbest65_Nbcd67
| Difference between SII Combination path and SII Air Conduction path | -0.006 [-0.063, 0.000] | -0.016 [-0.071, 0.000] | 35 / 35 | 675 | 0.461 | analysis_SII_Sbest65_Nbcd65.xlsx | BP110_SII_Sbest65_Nbcd65.tiff & BAHA5P_SII_Sbest65_Nbcd65.tiff | Analysis_SII_Sbest65_Nbcd68
| CVC score of the Air Conduction path | 0.95 [0.94, 0.95] | 0.95 [0.94, 0.95] | 35 / 35 | 610 | 0.961 | analysis_CVC_Sbest65_Nbcd65.xlsx | BP110_CVC_Sbest65_Nbcd65.tiff & BAHA5P_CVC_Sbest65_Nbcd65.tiff | Analysis_CVC_Sbest65_Nbcd65
| CVC score of the Bone Conduction path | 0.34 [0.31, 0.37] | 0.40 [0.31, 0.46] | 35 / 35 | 246 | 0.000 | analysis_CVC_Sbest65_Nbcd65.xlsx | BP110_CVC_Sbest65_Nbcd65.tiff & BAHA5P_CVC_Sbest65_Nbcd65.tiff | Analysis_CVC_Sbest65_Nbcd66
| CVC score of the Combination path | 0.94 [0.93, 0.95] | 0.95 [0.93, 0.95] | 35 / 35 | 564 | 0.530 | analysis_CVC_Sbest65_Nbcd65.xlsx | BP110_CVC_Sbest65_Nbcd65.tiff & BAHA5P_CVC_Sbest65_Nbcd65.tiff | Analysis_CVC_Sbest65_Nbcd67
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [-0.02, 0.00] | 0.00 [-0.02, 0.00] | 35 / 35 | 593 | 0.794 | analysis_CVC_Sbest65_Nbcd65.xlsx | BP110_CVC_Sbest65_Nbcd65.tiff & BAHA5P_CVC_Sbest65_Nbcd65.tiff | Analysis_CVC_Sbest65_Nbcd68
Signal best ear side, noise in front | SII Air Conduction path | 0.594 [0.577, 0.594] | 0.594 [0.575, 0.594] | 35 / 35 | 458 | 0.027 | analysis_SII_Sbest65_Nfront65.xlsx | BP110_SII_Sbest65_Nfront65.tiff & BAHA5P_SII_Sbest65_Nfront65.tiff | Analysis_SII_Sbest65_Nfront65
| SII Bone Conduction path | 0.154 [0.138, 0.176] | 0.221 [0.190, 0.240] | 35 / 35 | 16 | 0.000 | analysis_SII_Sbest65_Nfront65.xlsx | BP110_SII_Sbest65_Nfront65.tiff & BAHA5P_SII_Sbest65_Nfront65.tiff | Analysis_SII_Sbest65_Nfront66
| SII Combination path | 0.592 [0.568, 0.594] | 0.594 [0.575, 0.594] | 35 / 35 | 493 | 0,027 | analysis_SII_Sbest65_Nfront65.xlsx | BP110_SII_Sbest65_Nfront65.tiff & BAHA5P_SII_Sbest65_Nfront65.tiff | Analysis_SII_Sbest65_Nfront67
| Difference between SII Combination path and SII Air Conduction path | 0.000 [-0.013, 0.000] | 0.000 [-0.017, 0.000] | 35 / 35 | 588 | 0.734 | analysis_SII_Sbest65_Nfront65.xlsx | BP110_SII_Sbest65_Nfront65.tiff & BAHA5P_SII_Sbest65_Nfront65.tiff | Analysis_SII_Sbest65_Nfront68
| CVC score of the Air Conduction path | 0.91 [0.90, 0.91] | 0.91 [0.90, 0.91] | 35 / 35 | 508 | 0.09 | analysis_CVC_Sbest65_Nfront65.xlsx | BP110_CVC_Sbest65_Nfront65.tiff & BAHA5P_CVC_Sbest65_Nfront65.tiff | Analysis_CVC_Sbest65_Nfront65
| CVC score of the Bone Conduction path | 0.48 [0.45, 0.53] | 0.61 [0.55, 0.63] | 35 / 35 | 15 | 0.000 | analysis_CVC_Sbest65_Nfront65.xlsx | BP110_CVC_Sbest65_Nfront65.tiff & BAHA5P_CVC_Sbest65_Nfront65.tiff | Analysis_CVC_Sbest65_Nfront66
| CVC score of the Combination path | 0.91 [0.90, 0.91] | 0.91 [0.90, 0.91] | 35 / 35 | 523 | 0.194 | analysis_CVC_Sbest65_Nfront65.xlsx | BP110_CVC_Sbest65_Nfront65.tiff & BAHA5P_CVC_Sbest65_Nfront65.tiff | Analysis_CVC_Sbest65_Nfront67
| Difference between CVC score Combination path and CVC score Air Conduction path | 0.00 [-0.01, 0.00] | 0.00 [-0.01, 0.00] | 35 / 35 | 635 | 0.676 | analysis_CVC_Sbest65_Nfront65.xlsx | BP110_CVC_Sbest65_Nfront65.tiff & BAHA5P_CVC_Sbest65_Nfront65.tiff | Analysis_CVC_Sbest65_Nfront68
* Fisher exact test
† Calculation not possible, due to the same values