## Overview
Overview Pyranometer
TH Köln |  | 2022-12-23
Pyranometer | Serial number | Sensibility (acc. Manufacturer) [10^-6V/W/m^2] | Calibration | sensibility (adapted) [10^-6V/W/m^2] | Notes
CM11 (Reference) | 007398 | 5.13 | 2019-11-06 | 5.259 | Blue label on the cable
CMP11 | 127037 | 8.83 | 2020-01-21 | 9.1833
CMP11 | 127038 | 8.94 | 2020-01-21 | 9.3783
Calibration with Logbox SE | 2020-01-21
Socket | Pyranometer | Serial number | Sensibility [10^-6V/W/m^2]
A9 | CM11 | 007398 | 5.259
A10 | CMP11 | 127038 | 9.3783
A11 | CMP11 | 127037 | 9.1833
Performed by: | Eva-Maria Grommes
Albedo measurements
Ghana
Location | Calculated Albedo | Error propagation [+/-] | Albedo Range
Akwatia | 0.19995 | 9.746% | 0.18047 | 0.21944
Accra | 0.26861 | 9.746% | 0.24244 | 0.29479
Kumasi | 0.31022 | 9.746% | 0.27999 | 0.34046
Average Albedo | 0.2596 | 9.746% | 0.2343 | 0.2849
0.0253
The measurement span of the albedo of red soil in three different locations in Ghana including measurement uncertainty is between 0.175 and 0.335.

## Calibration
Calibration Pyranometer
TH Köln |  | 2022-12-23
Calibration mit Logbox SE | 2020-01-21
Socket | Pyranometer | Serial number | Sensibility [10^-6V/W/m^2]
A9 | CM11 | 007398 | 5.259
A10 | CMP11 | 127038 | 9.3783
A11 | CMP11 | 127037 | 9.1833
Measurement data | [5 empty] | A10 | A11
Date | Time | 007398_Avg [V] | 127038_Avg [V] | 127037_Avg [V] | Irradiance [W/m^2] | Sensitivity | Sensitivity |  | Average Sensitivity | Manufac. Sensitivity | Deviation [%] | Time [s] | CM11 (Reference) | CMP11_1 | CMP11_2
2020-01-21 | 12:02 | 1.555 | 2.775 | 2.712 | 295.6836 | 9.385 | 9.172 | A10 | 9.3783 | 8.94 | 4.9026 | 0 | 295.6836 | 310.4027 | 307.1348
2020-01-21 | 12:02 | 1.554 | 2.775 | 2.712 | 295.4934 | 9.3911 | 9.1779 | A11 | 9.1833 | 8.83 | 4.001 | 5 | 295.4934 | 310.4027 | 307.1348
2020-01-21 | 12:02 | 1.555 | 2.777 | 2.715 | 295.6836 | 9.3918 | 9.1821 |  |  |  |  | 10 | 295.6836 | 310.6264 | 307.4745
2020-01-21 | 12:02 | 1.555 | 2.777 | 2.715 | 295.6836 | 9.3918 | 9.1821 |  |  |  |  | 15 | 295.6836 | 310.6264 | 307.4745
2020-01-21 | 12:02 | 1.556 | 2.778 | 2.715 | 295.8737 | 9.3891 | 9.1762 |  |  |  |  | 20 | 295.8737 | 310.7383 | 307.4745
2020-01-21 | 12:03 | 1.556 | 2.779 | 2.718 | 295.8737 | 9.3925 | 9.1864 |  |  |  |  | 25 | 295.8737 | 310.8501 | 307.8143
2020-01-21 | 12:03 | 1.558 | 2.782 | 2.72 | 296.254 | 9.3906 | 9.1813 |  |  |  |  | 30 | 296.254 | 311.1857 | 308.0408
2020-01-21 | 12:03 | 1.558 | 2.782 | 2.721 | 296.254 | 9.3906 | 9.1847 |  |  |  |  | 35 | 296.254 | 311.1857 | 308.154
2020-01-21 | 12:03 | 1.559 | 2.785 | 2.724 | 296.4442 | 9.3947 | 9.1889 |  |  |  |  | 40 | 296.4442 | 311.5213 | 308.4938
2020-01-21 | 12:03 | 1.561 | 2.787 | 2.725 | 296.8245 | 9.3894 | 9.1805 |  |  |  |  | 45 | 296.8245 | 311.745 | 308.607
2020-01-21 | 12:03 | 1.564 | 2.787 | 2.724 | 297.3949 | 9.3714 | 9.1595 |  |  |  |  | 50 | 297.3949 | 311.745 | 308.4938
2020-01-21 | 12:03 | 1.565 | 2.788 | 2.723 | 297.5851 | 9.3687 | 9.1503 |  |  |  |  | 55 | 297.5851 | 311.8568 | 308.3805
2020-01-21 | 12:03 | 1.565 | 2.788 | 2.721 | 297.5851 | 9.3687 | 9.1436 |  |  |  |  | 60 | 297.5851 | 311.8568 | 308.154
2020-01-21 | 12:03 | 1.565 | 2.784 | 2.718 | 297.5851 | 9.3553 | 9.1335 |  |  |  |  | 65 | 297.5851 | 311.4094 | 307.8143
2020-01-21 | 12:03 | 1.562 | 2.783 | 2.719 | 297.0146 | 9.3699 | 9.1544 |  |  |  |  | 70 | 297.0146 | 311.2975 | 307.9275
2020-01-21 | 12:03 | 1.56 | 2.783 | 2.72 | 296.6343 | 9.3819 | 9.1695 |  |  |  |  | 75 | 296.6343 | 311.2975 | 308.0408
2020-01-21 | 12:04 | 1.56 | 2.781 | 2.719 | 296.6343 | 9.3752 | 9.1662 |  |  |  |  | 80 | 296.6343 | 311.0738 | 307.9275
2020-01-21 | 12:04 | 1.559 | 2.78 | 2.718 | 296.4442 | 9.3778 | 9.1687 |  |  |  |  | 85 | 296.4442 | 310.962 | 307.8143
2020-01-21 | 12:04 | 1.559 | 2.781 | 2.718 | 296.4442 | 9.3812 | 9.1687 |  |  |  |  | 90 | 296.4442 | 311.0738 | 307.8143
2020-01-21 | 12:04 | 1.559 | 2.78 | 2.718 | 296.4442 | 9.3778 | 9.1687 |  |  |  |  | 95 | 296.4442 | 310.962 | 307.8143
2020-01-21 | 12:04 | 1.559 | 2.78 | 2.718 | 296.4442 | 9.3778 | 9.1687 |  |  |  |  | 100 | 296.4442 | 310.962 | 307.8143
2020-01-21 | 12:04 | 1.557 | 2.779 | 2.718 | 296.0639 | 9.3865 | 9.1805 |  |  |  |  | 105 | 296.0639 | 310.8501 | 307.8143
2020-01-21 | 12:04 | 1.558 | 2.781 | 2.719 | 296.254 | 9.3872 | 9.1779 |  |  |  |  | 110 | 296.254 | 311.0738 | 307.9275
2020-01-21 | 12:04 | 1.558 | 2.782 | 2.72 | 296.254 | 9.3906 | 9.1813 |  |  |  |  | 115 | 296.254 | 311.1857 | 308.0408
2020-01-21 | 12:04 | 1.557 | 2.782 | 2.721 | 296.0639 | 9.3966 | 9.1906 |  |  |  |  | 120 | 296.0639 | 311.1857 | 308.154
2020-01-21 | 12:04 | 1.559 | 2.78 | 2.718 | 296.4442 | 9.3778 | 9.1687 |  |  |  |  | 125 | 296.4442 | 310.962 | 307.8143
2020-01-21 | 12:04 | 1.561 | 2.778 | 2.715 | 296.8245 | 9.3591 | 9.1468 |  |  |  |  | 130 | 296.8245 | 310.7383 | 307.4745
2020-01-21 | 12:04 | 1.56 | 2.778 | 2.715 | 296.6343 | 9.3651 | 9.1527 |  |  |  |  | 135 | 296.6343 | 310.7383 | 307.4745
2020-01-21 | 12:04 | 1.56 | 2.777 | 2.715 | 296.6343 | 9.3617 | 9.1527 |  |  |  |  | 140 | 296.6343 | 310.6264 | 307.4745
2020-01-21 | 12:05 | 1.561 | 2.777 | 2.714 | 296.8245 | 9.3557 | 9.1435 |  |  |  |  | 145 | 296.8245 | 310.6264 | 307.3613
2020-01-21 | 12:05 | 1.562 | 2.779 | 2.715 | 297.0146 | 9.3564 | 9.141 |  |  |  |  | 150 | 297.0146 | 310.8501 | 307.4745
2020-01-21 | 12:05 | 1.565 | 2.779 | 2.713 | 297.5851 | 9.3385 | 9.1167 |  |  |  |  | 155 | 297.5851 | 310.8501 | 307.248
2020-01-21 | 12:05 | 1.567 | 2.778 | 2.713 | 297.9654 | 9.3232 | 9.1051 |  |  |  |  | 160 | 297.9654 | 310.7383 | 307.248
2020-01-21 | 12:05 | 1.566 | 2.78 | 2.714 | 297.7752 | 9.3359 | 9.1143 |  |  |  |  | 165 | 297.7752 | 310.962 | 307.3613
2020-01-21 | 12:05 | 1.566 | 2.782 | 2.717 | 297.7752 | 9.3426 | 9.1243 |  |  |  |  | 170 | 297.7752 | 311.1857 | 307.701
2020-01-21 | 12:05 | 1.566 | 2.784 | 2.719 | 297.7752 | 9.3493 | 9.131 |  |  |  |  | 175 | 297.7752 | 311.4094 | 307.9275
2020-01-21 | 12:05 | 1.568 | 2.786 | 2.72 | 298.1555 | 9.3441 | 9.1228 |  |  |  |  | 180 | 298.1555 | 311.6331 | 308.0408
2020-01-21 | 12:05 | 1.569 | 2.787 | 2.721 | 298.3457 | 9.3415 | 9.1203 |  |  |  |  | 185 | 298.3457 | 311.745 | 308.154
2020-01-21 | 12:05 | 1.571 | 2.789 | 2.724 | 298.726 | 9.3363 | 9.1187 |  |  |  |  | 190 | 298.726 | 311.9687 | 308.4938
2020-01-21 | 12:05 | 1.569 | 2.791 | 2.726 | 298.3457 | 9.3549 | 9.1371 |  |  |  |  | 195 | 298.3457 | 312.1924 | 308.7203
2020-01-21 | 12:06 | 1.567 | 2.793 | 2.729 | 297.9654 | 9.3736 | 9.1588 |  |  |  |  | 200 | 297.9654 | 312.4161 | 309.06
2020-01-21 | 12:06 | 1.565 | 2.794 | 2.732 | 297.5851 | 9.3889 | 9.1806 |  |  |  |  | 205 | 297.5851 | 312.528 | 309.3998
2020-01-21 | 12:06 | 1.567 | 2.795 | 2.734 | 297.9654 | 9.3803 | 9.1756 |  |  |  |  | 210 | 297.9654 | 312.6398 | 309.6263
2020-01-21 | 12:06 | 1.569 | 2.795 | 2.733 | 298.3457 | 9.3683 | 9.1605 |  |  |  |  | 215 | 298.3457 | 312.6398 | 309.513
2020-01-21 | 12:06 | 1.569 | 2.795 | 2.732 | 298.3457 | 9.3683 | 9.1572 |  |  |  |  | 220 | 298.3457 | 312.6398 | 309.3998
2020-01-21 | 12:06 | 1.567 | 2.794 | 2.732 | 297.9654 | 9.3769 | 9.1689 |  |  |  |  | 225 | 297.9654 | 312.528 | 309.3998
2020-01-21 | 12:06 | 1.568 | 2.795 | 2.733 | 298.1555 | 9.3743 | 9.1664 |  |  |  |  | 230 | 298.1555 | 312.6398 | 309.513
2020-01-21 | 12:06 | 1.567 | 2.795 | 2.734 | 297.9654 | 9.3803 | 9.1756 |  |  |  |  | 235 | 297.9654 | 312.6398 | 309.6263
2020-01-21 | 12:06 | 1.567 | 2.797 | 2.736 | 297.9654 | 9.387 | 9.1823 |  |  |  |  | 240 | 297.9654 | 312.8635 | 309.8528
2020-01-21 | 12:06 | 1.566 | 2.8 | 2.74 | 297.7752 | 9.4031 | 9.2016 |  |  |  |  | 245 | 297.7752 | 313.1991 | 310.3058
2020-01-21 | 12:06 | 1.568 | 2.802 | 2.742 | 298.1555 | 9.3978 | 9.1965 |  |  |  |  | 250 | 298.1555 | 313.4228 | 310.5323
2020-01-21 | 12:06 | 1.569 | 2.803 | 2.742 | 298.3457 | 9.3951 | 9.1907 |  |  |  |  | 255 | 298.3457 | 313.5347 | 310.5323
2020-01-21 | 12:06 | 1.571 | 2.803 | 2.741 | 298.726 | 9.3832 | 9.1756 |  |  |  |  | 260 | 298.726 | 313.5347 | 310.419
2020-01-21 | 12:07 | 1.569 | 2.803 | 2.742 | 298.3457 | 9.3951 | 9.1907 |  |  |  |  | 265 | 298.3457 | 313.5347 | 310.5323
2020-01-21 | 12:07 | 1.568 | 2.804 | 2.743 | 298.1555 | 9.4045 | 9.1999 |  |  |  |  | 270 | 298.1555 | 313.6465 | 310.6455
2020-01-21 | 12:07 | 1.569 | 2.804 | 2.744 | 298.3457 | 9.3985 | 9.1974 |  |  |  |  | 275 | 298.3457 | 313.6465 | 310.7588
2020-01-21 | 12:07 | 1.571 | 2.807 | 2.745 | 298.726 | 9.3966 | 9.189 |  |  |  |  | 280 | 298.726 | 313.9821 | 310.872
2020-01-21 | 12:07 | 1.571 | 2.808 | 2.747 | 298.726 | 9.3999 | 9.1957 |  |  |  |  | 285 | 298.726 | 314.094 | 311.0985
2020-01-21 | 12:07 | 1.571 | 2.809 | 2.748 | 298.726 | 9.4033 | 9.1991 |  |  |  |  | 290 | 298.726 | 314.2058 | 311.2118
2020-01-21 | 12:07 | 1.573 | 2.81 | 2.749 | 299.1063 | 9.3947 | 9.1907 |  |  |  |  | 295 | 299.1063 | 314.3177 | 311.325
2020-01-21 | 12:07 | 1.572 | 2.811 | 2.75 | 298.9161 | 9.404 | 9.1999 |  |  |  |  | 300 | 298.9161 | 314.4295 | 311.4383
2020-01-21 | 12:07 | 1.573 | 2.811 | 2.751 | 299.1063 | 9.398 | 9.1974 |  |  |  |  | 305 | 299.1063 | 314.4295 | 311.5515
2020-01-21 | 12:07 | 1.574 | 2.814 | 2.753 | 299.2964 | 9.402 | 9.1982 |  |  |  |  | 310 | 299.2964 | 314.7651 | 311.778
2020-01-21 | 12:07 | 1.573 | 2.812 | 2.752 | 299.1063 | 9.4013 | 9.2007 |  |  |  |  | 315 | 299.1063 | 314.5414 | 311.6648
2020-01-21 | 12:08 | 1.574 | 2.813 | 2.753 | 299.2964 | 9.3987 | 9.1982 |  |  |  |  | 320 | 299.2964 | 314.6532 | 311.778
2020-01-21 | 12:08 | 1.572 | 2.81 | 2.751 | 298.9161 | 9.4006 | 9.2033 |  |  |  |  | 325 | 298.9161 | 314.3177 | 311.5515
2020-01-21 | 12:08 | 1.573 | 2.81 | 2.75 | 299.1063 | 9.3947 | 9.1941 |  |  |  |  | 330 | 299.1063 | 314.3177 | 311.4383
2020-01-21 | 12:08 | 1.572 | 2.808 | 2.748 | 298.9161 | 9.3939 | 9.1932 |  |  |  |  | 335 | 298.9161 | 314.094 | 311.2118
2020-01-21 | 12:08 | 1.57 | 2.806 | 2.746 | 298.5358 | 9.3992 | 9.1982 |  |  |  |  | 340 | 298.5358 | 313.8702 | 310.9853
2020-01-21 | 12:08 | 1.569 | 2.804 | 2.743 | 298.3457 | 9.3985 | 9.194 |  |  |  |  | 345 | 298.3457 | 313.6465 | 310.6455
2020-01-21 | 12:08 | 1.57 | 2.802 | 2.742 | 298.5358 | 9.3858 | 9.1848 |  |  |  |  | 350 | 298.5358 | 313.4228 | 310.5323
2020-01-21 | 12:08 | 1.572 | 2.804 | 2.743 | 298.9161 | 9.3806 | 9.1765 |  |  |  |  | 355 | 298.9161 | 313.6465 | 310.6455
2020-01-21 | 12:08 | 1.572 | 2.805 | 2.743 | 298.9161 | 9.3839 | 9.1765 |  |  |  |  | 360 | 298.9161 | 313.7584 | 310.6455
2020-01-21 | 12:08 | 1.571 | 2.805 | 2.743 | 298.726 | 9.3899 | 9.1823 |  |  |  |  | 365 | 298.726 | 313.7584 | 310.6455
2020-01-21 | 12:08 | 1.574 | 2.806 | 2.742 | 299.2964 | 9.3753 | 9.1615 |  |  |  |  | 370 | 299.2964 | 313.8702 | 310.5323
2020-01-21 | 12:08 | 1.574 | 2.804 | 2.742 | 299.2964 | 9.3686 | 9.1615 |  |  |  |  | 375 | 299.2964 | 313.6465 | 310.5323
2020-01-21 | 12:08 | 1.573 | 2.807 | 2.744 | 299.1063 | 9.3846 | 9.174 |  |  |  |  | 380 | 299.1063 | 313.9821 | 310.7588
2020-01-21 | 12:09 | 1.574 | 2.808 | 2.747 | 299.2964 | 9.382 | 9.1782 |  |  |  |  | 385 | 299.2964 | 314.094 | 311.0985
2020-01-21 | 12:09 | 1.575 | 2.81 | 2.75 | 299.4866 | 9.3827 | 9.1824 |  |  |  |  | 390 | 299.4866 | 314.3177 | 311.4383
2020-01-21 | 12:09 | 1.576 | 2.813 | 2.752 | 299.6767 | 9.3868 | 9.1832 |  |  |  |  | 395 | 299.6767 | 314.6532 | 311.6648
2020-01-21 | 12:09 | 1.579 | 2.814 | 2.754 | 300.2472 | 9.3723 | 9.1724 |  |  |  |  | 400 | 300.2472 | 314.7651 | 311.8913
2020-01-21 | 12:09 | 1.581 | 2.818 | 2.757 | 300.6275 | 9.3737 | 9.1708 |  |  |  |  | 405 | 300.6275 | 315.2125 | 312.231
2020-01-21 | 12:09 | 1.583 | 2.818 | 2.757 | 301.0078 | 9.3619 | 9.1592 |  |  |  |  | 410 | 301.0078 | 315.2125 | 312.231
2020-01-21 | 12:09 | 1.581 | 2.819 | 2.757 | 300.6275 | 9.3771 | 9.1708 |  |  |  |  | 415 | 300.6275 | 315.3244 | 312.231
2020-01-21 | 12:09 | 1.583 | 2.823 | 2.761 | 301.0078 | 9.3785 | 9.1725 |  |  |  |  | 420 | 301.0078 | 315.7718 | 312.684
2020-01-21 | 12:09 | 1.583 | 2.821 | 2.763 | 301.0078 | 9.3719 | 9.1792 |  |  |  |  | 425 | 301.0078 | 315.5481 | 312.9105
2020-01-21 | 12:09 | 1.582 | 2.824 | 2.765 | 300.8176 | 9.3877 | 9.1916 |  |  |  |  | 430 | 300.8176 | 315.8837 | 313.137
2020-01-21 | 12:09 | 1.584 | 2.827 | 2.767 | 301.1979 | 9.3859 | 9.1866 |  |  |  |  | 435 | 301.1979 | 316.2192 | 313.3635
2020-01-21 | 12:10 | 1.585 | 2.828 | 2.769 | 301.3881 | 9.3833 | 9.1875 |  |  |  |  | 440 | 301.3881 | 316.3311 | 313.59
2020-01-21 | 12:10 | 1.583 | 2.826 | 2.766 | 301.0078 | 9.3885 | 9.1891 |  |  |  |  | 445 | 301.0078 | 316.1074 | 313.2503
2020-01-21 | 12:10 | 1.584 | 2.827 | 2.766 | 301.1979 | 9.3859 | 9.1833 |  |  |  |  | 450 | 301.1979 | 316.2192 | 313.2503
2020-01-21 | 12:10 | 1.584 | 2.823 | 2.763 | 301.1979 | 9.3726 | 9.1734 |  |  |  |  | 455 | 301.1979 | 315.7718 | 312.9105
2020-01-21 | 12:10 | 1.581 | 2.823 | 2.764 | 300.6275 | 9.3904 | 9.1941 |  |  |  |  | 460 | 300.6275 | 315.7718 | 313.0238
2020-01-21 | 12:10 | 1.582 | 2.825 | 2.765 | 300.8176 | 9.3911 | 9.1916 |  |  |  |  | 465 | 300.8176 | 315.9955 | 313.137
2020-01-21 | 12:10 | 1.581 | 2.823 | 2.764 | 300.6275 | 9.3904 | 9.1941 |  |  |  |  | 470 | 300.6275 | 315.7718 | 313.0238
2020-01-21 | 12:10 | 1.582 | 2.825 | 2.766 | 300.8176 | 9.3911 | 9.1949 |  |  |  |  | 475 | 300.8176 | 315.9955 | 313.2503
2020-01-21 | 12:10 | 1.583 | 2.826 | 2.767 | 301.0078 | 9.3885 | 9.1925 |  |  |  |  | 480 | 301.0078 | 316.1074 | 313.3635
2020-01-21 | 12:10 | 1.584 | 2.826 | 2.765 | 301.1979 | 9.3825 | 9.18 |  |  |  |  | 485 | 301.1979 | 316.1074 | 313.137
2020-01-21 | 12:10 | 1.586 | 2.826 | 2.765 | 301.5782 | 9.3707 | 9.1684 |  |  |  |  | 490 | 301.5782 | 316.1074 | 313.137
2020-01-21 | 12:10 | 1.586 | 2.826 | 2.765 | 301.5782 | 9.3707 | 9.1684 |  |  |  |  | 495 | 301.5782 | 316.1074 | 313.137
2020-01-21 | 12:10 | 1.585 | 2.825 | 2.764 | 301.3881 | 9.3733 | 9.1709 |  |  |  |  | 500 | 301.3881 | 315.9955 | 313.0238
2020-01-21 | 12:11 | 1.586 | 2.828 | 2.766 | 301.5782 | 9.3773 | 9.1717 |  |  |  |  | 505 | 301.5782 | 316.3311 | 313.2503
2020-01-21 | 12:11 | 1.585 | 2.829 | 2.768 | 301.3881 | 9.3866 | 9.1842 |  |  |  |  | 510 | 301.3881 | 316.443 | 313.4768
2020-01-21 | 12:11 | 1.587 | 2.83 | 2.771 | 301.7684 | 9.3781 | 9.1825 |  |  |  |  | 515 | 301.7684 | 316.5548 | 313.8165
2020-01-21 | 12:11 | 1.586 | 2.831 | 2.772 | 301.5782 | 9.3873 | 9.1916 |  |  |  |  | 520 | 301.5782 | 316.6667 | 313.9298
2020-01-21 | 12:11 | 1.587 | 2.832 | 2.772 | 301.7684 | 9.3847 | 9.1859 |  |  |  |  | 525 | 301.7684 | 316.7785 | 313.9298
2020-01-21 | 12:11 | 1.589 | 2.835 | 2.775 | 302.1487 | 9.3828 | 9.1842 |  |  |  |  | 530 | 302.1487 | 317.1141 | 314.2695
2020-01-21 | 12:11 | 1.589 | 2.837 | 2.777 | 302.1487 | 9.3894 | 9.1908 |  |  |  |  | 535 | 302.1487 | 317.3378 | 314.496
2020-01-21 | 12:11 | 1.59 | 2.838 | 2.779 | 302.3388 | 9.3868 | 9.1917 |  |  |  |  | 540 | 302.3388 | 317.4497 | 314.7225
2020-01-21 | 12:11 | 1.589 | 2.838 | 2.778 | 302.1487 | 9.3927 | 9.1941 |  |  |  |  | 545 | 302.1487 | 317.4497 | 314.6093
2020-01-21 | 12:11 | 1.589 | 2.837 | 2.778 | 302.1487 | 9.3894 | 9.1941 |  |  |  |  | 550 | 302.1487 | 317.3378 | 314.6093
2020-01-21 | 12:11 | 1.59 | 2.838 | 2.778 | 302.3388 | 9.3868 | 9.1884 |  |  |  |  | 555 | 302.3388 | 317.4497 | 314.6093
2020-01-21 | 12:12 | 1.591 | 2.838 | 2.779 | 302.529 | 9.3809 | 9.1859 |  |  |  |  | 560 | 302.529 | 317.4497 | 314.7225
2020-01-21 | 12:12 | 1.591 | 2.838 | 2.779 | 302.529 | 9.3809 | 9.1859 |  |  |  |  | 565 | 302.529 | 317.4497 | 314.7225
2020-01-21 | 12:12 | 1.59 | 2.837 | 2.777 | 302.3388 | 9.3835 | 9.1851 |  |  |  |  | 570 | 302.3388 | 317.3378 | 314.496
2020-01-21 | 12:12 | 1.589 | 2.835 | 2.776 | 302.1487 | 9.3828 | 9.1875 |  |  |  |  | 575 | 302.1487 | 317.1141 | 314.3828
2020-01-21 | 12:12 | 1.589 | 2.835 | 2.776 | 302.1487 | 9.3828 | 9.1875 |  |  |  |  | 580 | 302.1487 | 317.1141 | 314.3828
2020-01-21 | 12:12 | 1.59 | 2.836 | 2.776 | 302.3388 | 9.3802 | 9.1818 |  |  |  |  | 585 | 302.3388 | 317.226 | 314.3828
2020-01-21 | 12:12 | 1.589 | 2.836 | 2.775 | 302.1487 | 9.3861 | 9.1842 |  |  |  |  | 590 | 302.1487 | 317.226 | 314.2695
2020-01-21 | 12:12 | 1.591 | 2.839 | 2.778 | 302.529 | 9.3842 | 9.1826 |  |  |  |  | 595 | 302.529 | 317.5615 | 314.6093
2020-01-21 | 12:12 | 1.591 | 2.839 | 2.779 | 302.529 | 9.3842 | 9.1859 |  |  |  |  | 600 | 302.529 | 317.5615 | 314.7225
2020-01-21 | 12:12 | 1.592 | 2.839 | 2.777 | 302.7191 | 9.3783 | 9.1735 |  |  |  |  | 605 | 302.7191 | 317.5615 | 314.496
2020-01-21 | 12:12 | 1.591 | 2.839 | 2.779 | 302.529 | 9.3842 | 9.1859 |  |  |  |  | 610 | 302.529 | 317.5615 | 314.7225
2020-01-21 | 12:12 | 1.592 | 2.843 | 2.784 | 302.7191 | 9.3915 | 9.1966 |  |  |  |  | 615 | 302.7191 | 318.0089 | 315.2888
2020-01-21 | 12:12 | 1.593 | 2.844 | 2.786 | 302.9093 | 9.3889 | 9.1975 |  |  |  |  | 620 | 302.9093 | 318.1208 | 315.5153
2020-01-21 | 12:13 | 1.593 | 2.845 | 2.785 | 302.9093 | 9.3923 | 9.1942 |  |  |  |  | 625 | 302.9093 | 318.2327 | 315.402
2020-01-21 | 12:13 | 1.594 | 2.844 | 2.785 | 303.0994 | 9.3831 | 9.1884 |  |  |  |  | 630 | 303.0994 | 318.1208 | 315.402
2020-01-21 | 12:13 | 1.594 | 2.845 | 2.785 | 303.0994 | 9.3864 | 9.1884 |  |  |  |  | 635 | 303.0994 | 318.2327 | 315.402
2020-01-21 | 12:13 | 1.595 | 2.845 | 2.785 | 303.2896 | 9.3805 | 9.1826 |  |  |  |  | 640 | 303.2896 | 318.2327 | 315.402
2020-01-21 | 12:13 | 1.594 | 2.844 | 2.785 | 303.0994 | 9.3831 | 9.1884 |  |  |  |  | 645 | 303.0994 | 318.1208 | 315.402
2020-01-21 | 12:13 | 1.594 | 2.843 | 2.784 | 303.0994 | 9.3798 | 9.1851 |  |  |  |  | 650 | 303.0994 | 318.0089 | 315.2888
2020-01-21 | 12:13 | 1.594 | 2.843 | 2.785 | 303.0994 | 9.3798 | 9.1884 |  |  |  |  | 655 | 303.0994 | 318.0089 | 315.402
2020-01-21 | 12:13 | 1.594 | 2.845 | 2.788 | 303.0994 | 9.3864 | 9.1983 |  |  |  |  | 660 | 303.0994 | 318.2327 | 315.7418
2020-01-21 | 12:13 | 1.595 | 2.846 | 2.788 | 303.2896 | 9.3838 | 9.1925 |  |  |  |  | 665 | 303.2896 | 318.3445 | 315.7418
2020-01-21 | 12:13 | 1.597 | 2.846 | 2.788 | 303.6699 | 9.372 | 9.181 |  |  |  |  | 670 | 303.6699 | 318.3445 | 315.7418
2020-01-21 | 12:13 | 1.596 | 2.847 | 2.788 | 303.4797 | 9.3812 | 9.1868 |  |  |  |  | 675 | 303.4797 | 318.4564 | 315.7418
2020-01-21 | 12:14 | 1.595 | 2.848 | 2.788 | 303.2896 | 9.3904 | 9.1925 |  |  |  |  | 680 | 303.2896 | 318.5682 | 315.7418
2020-01-21 | 12:14 | 1.596 | 2.849 | 2.789 | 303.4797 | 9.3878 | 9.1901 |  |  |  |  | 685 | 303.4797 | 318.6801 | 315.855
2020-01-21 | 12:14 | 1.596 | 2.847 | 2.788 | 303.4797 | 9.3812 | 9.1868 |  |  |  |  | 690 | 303.4797 | 318.4564 | 315.7418
2020-01-21 | 12:14 | 1.596 | 2.846 | 2.787 | 303.4797 | 9.3779 | 9.1835 |  |  |  |  | 695 | 303.4797 | 318.3445 | 315.6285
2020-01-21 | 12:14 | 1.597 | 2.848 | 2.789 | 303.6699 | 9.3786 | 9.1843 |  |  |  |  | 700 | 303.6699 | 318.5682 | 315.855
2020-01-21 | 12:14 | 1.596 | 2.847 | 2.787 | 303.4797 | 9.3812 | 9.1835 |  |  |  |  | 705 | 303.4797 | 318.4564 | 315.6285
2020-01-21 | 12:14 | 1.595 | 2.847 | 2.788 | 303.2896 | 9.3871 | 9.1925 |  |  |  |  | 710 | 303.2896 | 318.4564 | 315.7418
2020-01-21 | 12:14 | 1.596 | 2.847 | 2.788 | 303.4797 | 9.3812 | 9.1868 |  |  |  |  | 715 | 303.4797 | 318.4564 | 315.7418
2020-01-21 | 12:14 | 1.597 | 2.846 | 2.785 | 303.6699 | 9.372 | 9.1711 |  |  |  |  | 720 | 303.6699 | 318.3445 | 315.402
2020-01-21 | 12:14 | 1.595 | 2.844 | 2.784 | 303.2896 | 9.3772 | 9.1793 |  |  |  |  | 725 | 303.2896 | 318.1208 | 315.2888
2020-01-21 | 12:14 | 1.594 | 2.844 | 2.785 | 303.0994 | 9.3831 | 9.1884 |  |  |  |  | 730 | 303.0994 | 318.1208 | 315.402
2020-01-21 | 12:14 | 1.593 | 2.844 | 2.784 | 302.9093 | 9.3889 | 9.1909 |  |  |  |  | 735 | 302.9093 | 318.1208 | 315.2888
2020-01-21 | 12:14 | 1.595 | 2.843 | 2.783 | 303.2896 | 9.3739 | 9.176 |  |  |  |  | 740 | 303.2896 | 318.0089 | 315.1755
2020-01-21 | 12:15 | 1.594 | 2.842 | 2.782 | 303.0994 | 9.3765 | 9.1785 |  |  |  |  | 745 | 303.0994 | 317.8971 | 315.0623
2020-01-21 | 12:15 | 1.594 | 2.843 | 2.783 | 303.0994 | 9.3798 | 9.1818 |  |  |  |  | 750 | 303.0994 | 318.0089 | 315.1755
2020-01-21 | 12:15 | 1.595 | 2.848 | 2.788 | 303.2896 | 9.3904 | 9.1925 |  |  |  |  | 755 | 303.2896 | 318.5682 | 315.7418
2020-01-21 | 12:15 | 1.597 | 2.852 | 2.791 | 303.6699 | 9.3918 | 9.1909 |  |  |  |  | 760 | 303.6699 | 319.0157 | 316.0815
2020-01-21 | 12:15 | 1.597 | 2.854 | 2.794 | 303.6699 | 9.3984 | 9.2008 |  |  |  |  | 765 | 303.6699 | 319.2394 | 316.4213
2020-01-21 | 12:15 | 1.598 | 2.855 | 2.795 | 303.86 | 9.3958 | 9.1983 |  |  |  |  | 770 | 303.86 | 319.3512 | 316.5345
2020-01-21 | 12:15 | 1.601 | 2.857 | 2.798 | 304.4305 | 9.3847 | 9.1909 |  |  |  |  | 775 | 304.4305 | 319.5749 | 316.8743
2020-01-21 | 12:15 | 1.602 | 2.858 | 2.798 | 304.6207 | 9.3822 | 9.1852 |  |  |  |  | 780 | 304.6207 | 319.6868 | 316.8743
2020-01-21 | 12:15 | 1.601 | 2.859 | 2.798 | 304.4305 | 9.3913 | 9.1909 |  |  |  |  | 785 | 304.4305 | 319.7987 | 316.8743
2020-01-21 | 12:15 | 1.602 | 2.858 | 2.799 | 304.6207 | 9.3822 | 9.1885 |  |  |  |  | 790 | 304.6207 | 319.6868 | 316.9875
2020-01-21 | 12:15 | 1.602 | 2.86 | 2.8 | 304.6207 | 9.3887 | 9.1918 |  |  |  |  | 795 | 304.6207 | 319.9105 | 317.1008
2020-01-21 | 12:16 | 1.602 | 2.861 | 2.801 | 304.6207 | 9.392 | 9.195 |  |  |  |  | 800 | 304.6207 | 320.0224 | 317.214
2020-01-21 | 12:16 | 1.604 | 2.861 | 2.801 | 305.001 | 9.3803 | 9.1836 |  |  |  |  | 805 | 305.001 | 320.0224 | 317.214
2020-01-21 | 12:16 | 1.606 | 2.861 | 2.8 | 305.3813 | 9.3686 | 9.1689 |  |  |  |  | 810 | 305.3813 | 320.0224 | 317.1008
2020-01-21 | 12:16 | 1.605 | 2.86 | 2.799 | 305.1911 | 9.3712 | 9.1713 |  |  |  |  | 815 | 305.1911 | 319.9105 | 316.9875
2020-01-21 | 12:16 | 1.605 | 2.859 | 2.797 | 305.1911 | 9.3679 | 9.1647 |  |  |  |  | 820 | 305.1911 | 319.7987 | 316.761
2020-01-21 | 12:16 | 1.604 | 2.858 | 2.797 | 305.001 | 9.3705 | 9.1705 |  |  |  |  | 825 | 305.001 | 319.6868 | 316.761
2020-01-21 | 12:16 | 1.605 | 2.858 | 2.797 | 305.1911 | 9.3646 | 9.1647 |  |  |  |  | 830 | 305.1911 | 319.6868 | 316.761
2020-01-21 | 12:16 | 1.605 | 2.856 | 2.795 | 305.1911 | 9.3581 | 9.1582 |  |  |  |  | 835 | 305.1911 | 319.4631 | 316.5345
2020-01-21 | 12:16 | 1.604 | 2.855 | 2.794 | 305.001 | 9.3606 | 9.1606 |  |  |  |  | 840 | 305.001 | 319.3512 | 316.4213
2020-01-21 | 12:16 | 1.605 | 2.857 | 2.795 | 305.1911 | 9.3613 | 9.1582 |  |  |  |  | 845 | 305.1911 | 319.5749 | 316.5345
2020-01-21 | 12:16 | 1.606 | 2.854 | 2.792 | 305.3813 | 9.3457 | 9.1427 |  |  |  |  | 850 | 305.3813 | 319.2394 | 316.1948
2020-01-21 | 12:16 | 1.604 | 2.855 | 2.793 | 305.001 | 9.3606 | 9.1573 |  |  |  |  | 855 | 305.001 | 319.3512 | 316.308
2020-01-21 | 12:16 | 1.604 | 2.853 | 2.792 | 305.001 | 9.3541 | 9.1541 |  |  |  |  | 860 | 305.001 | 319.1275 | 316.1948
2020-01-21 | 12:17 | 1.601 | 2.854 | 2.793 | 304.4305 | 9.3749 | 9.1745 |  |  |  |  | 865 | 304.4305 | 319.2394 | 316.308
2020-01-21 | 12:17 | 1.601 | 2.854 | 2.795 | 304.4305 | 9.3749 | 9.1811 |  |  |  |  | 870 | 304.4305 | 319.2394 | 316.5345
2020-01-21 | 12:17 | 1.602 | 2.858 | 2.799 | 304.6207 | 9.3822 | 9.1885 |  |  |  |  | 875 | 304.6207 | 319.6868 | 316.9875
2020-01-21 | 12:17 | 1.604 | 2.859 | 2.8 | 305.001 | 9.3737 | 9.1803 |  |  |  |  | 880 | 305.001 | 319.7987 | 317.1008
2020-01-21 | 12:17 | 1.605 | 2.86 | 2.801 | 305.1911 | 9.3712 | 9.1779 |  |  |  |  | 885 | 305.1911 | 319.9105 | 317.214
2020-01-21 | 12:17 | 1.605 | 2.862 | 2.802 | 305.1911 | 9.3777 | 9.1811 |  |  |  |  | 890 | 305.1911 | 320.1342 | 317.3273
2020-01-21 | 12:17 | 1.604 | 2.863 | 2.804 | 305.001 | 9.3869 | 9.1934 |  |  |  |  | 895 | 305.001 | 320.2461 | 317.5538
2020-01-21 | 12:17 | 1.604 | 2.865 | 2.805 | 305.001 | 9.3934 | 9.1967 |  |  |  |  | 900 | 305.001 | 320.4698 | 317.667
2020-01-21 | 12:17 | 1.606 | 2.867 | 2.808 | 305.3813 | 9.3883 | 9.1951 |  |  |  |  | 905 | 305.3813 | 320.6935 | 318.0068
2020-01-21 | 12:17 | 1.61 | 2.869 | 2.809 | 306.1419 | 9.3715 | 9.1755 |  |  |  |  | 910 | 306.1419 | 320.9172 | 318.12
2020-01-21 | 12:17 | 1.61 | 2.871 | 2.81 | 306.1419 | 9.378 | 9.1788 |  |  |  |  | 915 | 306.1419 | 321.1409 | 318.2333
2020-01-21 | 12:18 | 1.61 | 2.871 | 2.811 | 306.1419 | 9.378 | 9.182 |  |  |  |  | 920 | 306.1419 | 321.1409 | 318.3465
2020-01-21 | 12:18 | 1.611 | 2.874 | 2.814 | 306.332 | 9.382 | 9.1861 |  |  |  |  | 925 | 306.332 | 321.4765 | 318.6863
2020-01-21 | 12:18 | 1.613 | 2.876 | 2.815 | 306.7123 | 9.3769 | 9.178 |  |  |  |  | 930 | 306.7123 | 321.7002 | 318.7995
2020-01-21 | 12:18 | 1.613 | 2.877 | 2.815 | 306.7123 | 9.3801 | 9.178 |  |  |  |  | 935 | 306.7123 | 321.8121 | 318.7995
2020-01-21 | 12:18 | 1.612 | 2.876 | 2.816 | 306.5222 | 9.3827 | 9.1869 |  |  |  |  | 940 | 306.5222 | 321.7002 | 318.9128
2020-01-21 | 12:18 | 1.612 | 2.876 | 2.817 | 306.5222 | 9.3827 | 9.1902 |  |  |  |  | 945 | 306.5222 | 321.7002 | 319.026
2020-01-21 | 12:18 | 1.613 | 2.876 | 2.817 | 306.7123 | 9.3769 | 9.1845 |  |  |  |  | 950 | 306.7123 | 321.7002 | 319.026
2020-01-21 | 12:18 | 1.613 | 2.873 | 2.814 | 306.7123 | 9.3671 | 9.1747 |  |  |  |  | 955 | 306.7123 | 321.3647 | 318.6863
2020-01-21 | 12:18 | 1.612 | 2.874 | 2.815 | 306.5222 | 9.3762 | 9.1837 |  |  |  |  | 960 | 306.5222 | 321.4765 | 318.7995
2020-01-21 | 12:18 | 1.611 | 2.873 | 2.814 | 306.332 | 9.3787 | 9.1861 |  |  |  |  | 965 | 306.332 | 321.3647 | 318.6863
2020-01-21 | 12:18 | 1.611 | 2.875 | 2.815 | 306.332 | 9.3852 | 9.1894 |  |  |  |  | 970 | 306.332 | 321.5884 | 318.7995
2020-01-21 | 12:18 | 1.612 | 2.877 | 2.817 | 306.5222 | 9.3859 | 9.1902 |  |  |  |  | 975 | 306.5222 | 321.8121 | 319.026
2020-01-21 | 12:18 | 1.612 | 2.877 | 2.82 | 306.5222 | 9.3859 | 9.2 |  |  |  |  | 980 | 306.5222 | 321.8121 | 319.3658
2020-01-21 | 12:19 | 1.612 | 2.879 | 2.821 | 306.5222 | 9.3925 | 9.2032 |  |  |  |  | 985 | 306.5222 | 322.0358 | 319.479
2020-01-21 | 12:19 | 1.613 | 2.881 | 2.824 | 306.7123 | 9.3932 | 9.2073 |  |  |  |  | 990 | 306.7123 | 322.2595 | 319.8188
2020-01-21 | 12:19 | 1.615 | 2.884 | 2.825 | 307.0926 | 9.3913 | 9.1992 |  |  |  |  | 995 | 307.0926 | 322.5951 | 319.932
2020-01-21 | 12:19 | 1.616 | 2.887 | 2.828 | 307.2828 | 9.3953 | 9.2033 |  |  |  |  | 1000 | 307.2828 | 322.9306 | 320.2718
2020-01-21 | 12:19 | 1.616 | 2.887 | 2.829 | 307.2828 | 9.3953 | 9.2065 |  |  |  |  | 1005 | 307.2828 | 322.9306 | 320.3851
2020-01-21 | 12:19 | 1.617 | 2.888 | 2.831 | 307.4729 | 9.3927 | 9.2073 |  |  |  |  | 1010 | 307.4729 | 323.0425 | 320.6116
2020-01-21 | 12:19 | 1.617 | 2.889 | 2.831 | 307.4729 | 9.3959 | 9.2073 |  |  |  |  | 1015 | 307.4729 | 323.1544 | 320.6116
2020-01-21 | 12:19 | 1.619 | 2.889 | 2.831 | 307.8532 | 9.3843 | 9.1959 |  |  |  |  | 1020 | 307.8532 | 323.1544 | 320.6116
2020-01-21 | 12:19 | 1.619 | 2.888 | 2.829 | 307.8532 | 9.3811 | 9.1894 |  |  |  |  | 1025 | 307.8532 | 323.0425 | 320.3851
2020-01-21 | 12:19 | 1.618 | 2.889 | 2.832 | 307.6631 | 9.3901 | 9.2049 |  |  |  |  | 1030 | 307.6631 | 323.1544 | 320.7248
2020-01-21 | 12:19 | 1.619 | 2.89 | 2.832 | 307.8532 | 9.3876 | 9.1992 |  |  |  |  | 1035 | 307.8532 | 323.2662 | 320.7248
2020-01-21 | 12:20 | 1.618 | 2.887 | 2.829 | 307.6631 | 9.3836 | 9.1951 |  |  |  |  | 1040 | 307.6631 | 322.9306 | 320.3851
2020-01-21 | 12:20 | 1.618 | 2.885 | 2.827 | 307.6631 | 9.3771 | 9.1886 |  |  |  |  | 1045 | 307.6631 | 322.7069 | 320.1586
2020-01-21 | 12:20 | 1.617 | 2.886 | 2.828 | 307.4729 | 9.3862 | 9.1976 |  |  |  |  | 1050 | 307.4729 | 322.8188 | 320.2718
2020-01-21 | 12:20 | 1.616 | 2.883 | 2.826 | 307.2828 | 9.3822 | 9.1967 |  |  |  |  | 1055 | 307.2828 | 322.4832 | 320.0453
2020-01-21 | 12:20 | 1.616 | 2.885 | 2.826 | 307.2828 | 9.3887 | 9.1967 |  |  |  |  | 1060 | 307.2828 | 322.7069 | 320.0453
2020-01-21 | 12:20 | 1.618 | 2.886 | 2.827 | 307.6631 | 9.3804 | 9.1886 |  |  |  |  | 1065 | 307.6631 | 322.8188 | 320.1586
2020-01-21 | 12:20 | 1.618 | 2.886 | 2.826 | 307.6631 | 9.3804 | 9.1854 |  |  |  |  | 1070 | 307.6631 | 322.8188 | 320.0453
2020-01-21 | 12:20 | 1.618 | 2.886 | 2.827 | 307.6631 | 9.3804 | 9.1886 |  |  |  |  | 1075 | 307.6631 | 322.8188 | 320.1586
2020-01-21 | 12:20 | 1.619 | 2.886 | 2.827 | 307.8532 | 9.3746 | 9.1829 |  |  |  |  | 1080 | 307.8532 | 322.8188 | 320.1586
2020-01-21 | 12:20 | 1.618 | 2.887 | 2.828 | 307.6631 | 9.3836 | 9.1919 |  |  |  |  | 1085 | 307.6631 | 322.9306 | 320.2718
2020-01-21 | 12:20 | 1.618 | 2.886 | 2.828 | 307.6631 | 9.3804 | 9.1919 |  |  |  |  | 1090 | 307.6631 | 322.8188 | 320.2718
2020-01-21 | 12:20 | 1.618 | 2.885 | 2.825 | 307.6631 | 9.3771 | 9.1821 |  |  |  |  | 1095 | 307.6631 | 322.7069 | 319.932
2020-01-21 | 12:20 | 1.618 | 2.885 | 2.825 | 307.6631 | 9.3771 | 9.1821 |  |  |  |  | 1100 | 307.6631 | 322.7069 | 319.932
2020-01-21 | 12:21 | 1.618 | 2.885 | 2.825 | 307.6631 | 9.3771 | 9.1821 |  |  |  |  | 1105 | 307.6631 | 322.7069 | 319.932
2020-01-21 | 12:21 | 1.618 | 2.886 | 2.826 | 307.6631 | 9.3804 | 9.1854 |  |  |  |  | 1110 | 307.6631 | 322.8188 | 320.0453
2020-01-21 | 12:21 | 1.619 | 2.886 | 2.826 | 307.8532 | 9.3746 | 9.1797 |  |  |  |  | 1115 | 307.8532 | 322.8188 | 320.0453
2020-01-21 | 12:21 | 1.62 | 2.885 | 2.824 | 308.0434 | 9.3656 | 9.1675 |  |  |  |  | 1120 | 308.0434 | 322.7069 | 319.8188
2020-01-21 | 12:21 | 1.618 | 2.883 | 2.821 | 307.6631 | 9.3706 | 9.1691 |  |  |  |  | 1125 | 307.6631 | 322.4832 | 319.479
2020-01-21 | 12:21 | 1.617 | 2.882 | 2.822 | 307.4729 | 9.3732 | 9.178 |  |  |  |  | 1130 | 307.4729 | 322.3714 | 319.5923
2020-01-21 | 12:21 | 1.616 | 2.879 | 2.818 | 307.2828 | 9.3692 | 9.1707 |  |  |  |  | 1135 | 307.2828 | 322.0358 | 319.1393
2020-01-21 | 12:21 | 1.616 | 2.879 | 2.818 | 307.2828 | 9.3692 | 9.1707 |  |  |  |  | 1140 | 307.2828 | 322.0358 | 319.1393
2020-01-21 | 12:21 | 1.615 | 2.877 | 2.818 | 307.0926 | 9.3685 | 9.1764 |  |  |  |  | 1145 | 307.0926 | 321.8121 | 319.1393
2020-01-21 | 12:21 | 1.614 | 2.877 | 2.817 | 306.9025 | 9.3743 | 9.1788 |  |  |  |  | 1150 | 306.9025 | 321.8121 | 319.026
2020-01-21 | 12:21 | 1.614 | 2.878 | 2.819 | 306.9025 | 9.3776 | 9.1853 |  |  |  |  | 1155 | 306.9025 | 321.9239 | 319.2525
2020-01-21 | 12:22 | 1.616 | 2.878 | 2.819 | 307.2828 | 9.366 | 9.174 |  |  |  |  | 1160 | 307.2828 | 321.9239 | 319.2525
2020-01-21 | 12:22 | 1.617 | 2.88 | 2.82 | 307.4729 | 9.3667 | 9.1715 |  |  |  |  | 1165 | 307.4729 | 322.1477 | 319.3658
2020-01-21 | 12:22 | 1.617 | 2.88 | 2.82 | 307.4729 | 9.3667 | 9.1715 |  |  |  |  | 1170 | 307.4729 | 322.1477 | 319.3658
2020-01-21 | 12:22 | 1.617 | 2.882 | 2.823 | 307.4729 | 9.3732 | 9.1813 |  |  |  |  | 1175 | 307.4729 | 322.3714 | 319.7055
2020-01-21 | 12:22 | 1.617 | 2.882 | 2.822 | 307.4729 | 9.3732 | 9.178 |  |  |  |  | 1180 | 307.4729 | 322.3714 | 319.5923
2020-01-21 | 12:22 | 1.62 | 2.889 | 2.829 | 308.0434 | 9.3785 | 9.1838 |  |  |  |  | 1185 | 308.0434 | 323.1544 | 320.3851
2020-01-21 | 12:22 | 1.621 | 2.888 | 2.829 | 308.2335 | 9.3695 | 9.1781 |  |  |  |  | 1190 | 308.2335 | 323.0425 | 320.3851
2020-01-21 | 12:22 | 1.621 | 2.89 | 2.832 | 308.2335 | 9.376 | 9.1878 |  |  |  |  | 1195 | 308.2335 | 323.2662 | 320.7248
2020-01-21 | 12:22 | 1.62 | 2.892 | 2.833 | 308.0434 | 9.3883 | 9.1968 |  |  |  |  | 1200 | 308.0434 | 323.4899 | 320.8381
2020-01-21 | 12:22 | 1.621 | 2.89 | 2.831 | 308.2335 | 9.376 | 9.1846 |  |  |  |  | 1205 | 308.2335 | 323.2662 | 320.6116
2020-01-21 | 12:22 | 1.621 | 2.891 | 2.833 | 308.2335 | 9.3793 | 9.1911 |  |  |  |  | 1210 | 308.2335 | 323.3781 | 320.8381
2020-01-21 | 12:22 | 1.622 | 2.89 | 2.831 | 308.4237 | 9.3702 | 9.1789 |  |  |  |  | 1215 | 308.4237 | 323.2662 | 320.6116
2020-01-21 | 12:22 | 1.622 | 2.89 | 2.832 | 308.4237 | 9.3702 | 9.1822 |  |  |  |  | 1220 | 308.4237 | 323.2662 | 320.7248
2020-01-21 | 12:23 | 1.623 | 2.89 | 2.831 | 308.6138 | 9.3645 | 9.1733 |  |  |  |  | 1225 | 308.6138 | 323.2662 | 320.6116
2020-01-21 | 12:23 | 1.622 | 2.891 | 2.832 | 308.4237 | 9.3735 | 9.1822 |  |  |  |  | 1230 | 308.4237 | 323.3781 | 320.7248
2020-01-21 | 12:23 | 1.621 | 2.891 | 2.832 | 308.2335 | 9.3793 | 9.1878 |  |  |  |  | 1235 | 308.2335 | 323.3781 | 320.7248
2020-01-21 | 12:23 | 1.622 | 2.893 | 2.835 | 308.4237 | 9.38 | 9.1919 |  |  |  |  | 1240 | 308.4237 | 323.6018 | 321.0646
2020-01-21 | 12:23 | 1.622 | 2.893 | 2.835 | 308.4237 | 9.38 | 9.1919 |  |  |  |  | 1245 | 308.4237 | 323.6018 | 321.0646
2020-01-21 | 12:23 | 1.622 | 2.896 | 2.838 | 308.4237 | 9.3897 | 9.2016 |  |  |  |  | 1250 | 308.4237 | 323.9374 | 321.4043
2020-01-21 | 12:23 | 1.624 | 2.896 | 2.838 | 308.804 | 9.3781 | 9.1903 |  |  |  |  | 1255 | 308.804 | 323.9374 | 321.4043
2020-01-21 | 12:23 | 1.624 | 2.895 | 2.837 | 308.804 | 9.3749 | 9.1871 |  |  |  |  | 1260 | 308.804 | 323.8255 | 321.2911
2020-01-21 | 12:23 | 1.624 | 2.895 | 2.838 | 308.804 | 9.3749 | 9.1903 |  |  |  |  | 1265 | 308.804 | 323.8255 | 321.4043
2020-01-21 | 12:23 | 1.624 | 2.898 | 2.841 | 308.804 | 9.3846 | 9.2 |  |  |  |  | 1270 | 308.804 | 324.1611 | 321.7441
2020-01-21 | 12:23 | 1.625 | 2.897 | 2.838 | 308.9941 | 9.3756 | 9.1846 |  |  |  |  | 1275 | 308.9941 | 324.0492 | 321.4043
2020-01-21 | 12:24 | 1.625 | 2.896 | 2.838 | 308.9941 | 9.3723 | 9.1846 |  |  |  |  | 1280 | 308.9941 | 323.9374 | 321.4043
2020-01-21 | 12:24 | 1.625 | 2.896 | 2.838 | 308.9941 | 9.3723 | 9.1846 |  |  |  |  | 1285 | 308.9941 | 323.9374 | 321.4043
2020-01-21 | 12:24 | 1.624 | 2.894 | 2.837 | 308.804 | 9.3716 | 9.1871 |  |  |  |  | 1290 | 308.804 | 323.7136 | 321.2911
2020-01-21 | 12:24 | 1.623 | 2.892 | 2.833 | 308.6138 | 9.3709 | 9.1798 |  |  |  |  | 1295 | 308.6138 | 323.4899 | 320.8381
2020-01-21 | 12:24 | 1.622 | 2.89 | 2.833 | 308.4237 | 9.3702 | 9.1854 |  |  |  |  | 1300 | 308.4237 | 323.2662 | 320.8381
2020-01-21 | 12:24 | 1.623 | 2.891 | 2.832 | 308.6138 | 9.3677 | 9.1765 |  |  |  |  | 1305 | 308.6138 | 323.3781 | 320.7248
2020-01-21 | 12:24 | 1.622 | 2.891 | 2.832 | 308.4237 | 9.3735 | 9.1822 |  |  |  |  | 1310 | 308.4237 | 323.3781 | 320.7248
2020-01-21 | 12:24 | 1.622 | 2.888 | 2.828 | 308.4237 | 9.3637 | 9.1692 |  |  |  |  | 1315 | 308.4237 | 323.0425 | 320.2718
2020-01-21 | 12:24 | 1.62 | 2.889 | 2.83 | 308.0434 | 9.3785 | 9.187 |  |  |  |  | 1320 | 308.0434 | 323.1544 | 320.4983
2020-01-21 | 12:24 | 1.621 | 2.891 | 2.833 | 308.2335 | 9.3793 | 9.1911 |  |  |  |  | 1325 | 308.2335 | 323.3781 | 320.8381
2020-01-21 | 12:24 | 1.621 | 2.89 | 2.831 | 308.2335 | 9.376 | 9.1846 |  |  |  |  | 1330 | 308.2335 | 323.2662 | 320.6116
2020-01-21 | 12:24 | 1.621 | 2.89 | 2.832 | 308.2335 | 9.376 | 9.1878 |  |  |  |  | 1335 | 308.2335 | 323.2662 | 320.7248
2020-01-21 | 12:24 | 1.622 | 2.891 | 2.832 | 308.4237 | 9.3735 | 9.1822 |  |  |  |  | 1340 | 308.4237 | 323.3781 | 320.7248
2020-01-21 | 12:25 | 1.622 | 2.893 | 2.834 | 308.4237 | 9.38 | 9.1887 |  |  |  |  | 1345 | 308.4237 | 323.6018 | 320.9513
2020-01-21 | 12:25 | 1.623 | 2.894 | 2.836 | 308.6138 | 9.3774 | 9.1895 |  |  |  |  | 1350 | 308.6138 | 323.7136 | 321.1778
2020-01-21 | 12:25 | 1.622 | 2.893 | 2.837 | 308.4237 | 9.38 | 9.1984 |  |  |  |  | 1355 | 308.4237 | 323.6018 | 321.2911
2020-01-21 | 12:25 | 1.624 | 2.898 | 2.84 | 308.804 | 9.3846 | 9.1968 |  |  |  |  | 1360 | 308.804 | 324.1611 | 321.6308
2020-01-21 | 12:25 | 1.625 | 2.899 | 2.842 | 308.9941 | 9.3821 | 9.1976 |  |  |  |  | 1365 | 308.9941 | 324.2729 | 321.8573
2020-01-21 | 12:25 | 1.626 | 2.901 | 2.842 | 309.1843 | 9.3828 | 9.1919 |  |  |  |  | 1370 | 309.1843 | 324.4966 | 321.8573
2020-01-21 | 12:25 | 1.627 | 2.9 | 2.843 | 309.3744 | 9.3738 | 9.1895 |  |  |  |  | 1375 | 309.3744 | 324.3848 | 321.9706
2020-01-21 | 12:25 | 1.628 | 2.904 | 2.845 | 309.5646 | 9.3809 | 9.1903 |  |  |  |  | 1380 | 309.5646 | 324.8322 | 322.1971
2020-01-21 | 12:25 | 1.628 | 2.903 | 2.845 | 309.5646 | 9.3777 | 9.1903 |  |  |  |  | 1385 | 309.5646 | 324.7204 | 322.1971
2020-01-21 | 12:25 | 1.629 | 2.904 | 2.845 | 309.7547 | 9.3752 | 9.1847 |  |  |  |  | 1390 | 309.7547 | 324.8322 | 322.1971
2020-01-21 | 12:25 | 1.628 | 2.904 | 2.846 | 309.5646 | 9.3809 | 9.1936 |  |  |  |  | 1395 | 309.5646 | 324.8322 | 322.3103
2020-01-21 | 12:26 | 1.628 | 2.902 | 2.845 | 309.5646 | 9.3745 | 9.1903 |  |  |  |  | 1400 | 309.5646 | 324.6085 | 322.1971
2020-01-21 | 12:26 | 1.629 | 2.903 | 2.845 | 309.7547 | 9.3719 | 9.1847 |  |  |  |  | 1405 | 309.7547 | 324.7204 | 322.1971
2020-01-21 | 12:26 | 1.628 | 2.9 | 2.843 | 309.5646 | 9.368 | 9.1839 |  |  |  |  | 1410 | 309.5646 | 324.3848 | 321.9706
2020-01-21 | 12:26 | 1.628 | 2.9 | 2.842 | 309.5646 | 9.368 | 9.1806 |  |  |  |  | 1415 | 309.5646 | 324.3848 | 321.8573
2020-01-21 | 12:26 | 1.628 | 2.903 | 2.844 | 309.5646 | 9.3777 | 9.1871 |  |  |  |  | 1420 | 309.5646 | 324.7204 | 322.0838
2020-01-21 | 12:26 | 1.628 | 2.901 | 2.844 | 309.5646 | 9.3712 | 9.1871 |  |  |  |  | 1425 | 309.5646 | 324.4966 | 322.0838
2020-01-21 | 12:26 | 1.627 | 2.899 | 2.841 | 309.3744 | 9.3705 | 9.183 |  |  |  |  | 1430 | 309.3744 | 324.2729 | 321.7441
2020-01-21 | 12:26 | 1.625 | 2.898 | 2.842 | 308.9941 | 9.3788 | 9.1976 |  |  |  |  | 1435 | 308.9941 | 324.1611 | 321.8573
2020-01-21 | 12:26 | 1.626 | 2.899 | 2.84 | 309.1843 | 9.3763 | 9.1855 |  |  |  |  | 1440 | 309.1843 | 324.2729 | 321.6308
2020-01-21 | 12:26 | 1.625 | 2.899 | 2.84 | 308.9941 | 9.3821 | 9.1911 |  |  |  |  | 1445 | 308.9941 | 324.2729 | 321.6308
2020-01-21 | 12:26 | 1.626 | 2.899 | 2.841 | 309.1843 | 9.3763 | 9.1887 |  |  |  |  | 1450 | 309.1843 | 324.2729 | 321.7441
2020-01-21 | 12:26 | 1.626 | 2.896 | 2.839 | 309.1843 | 9.3666 | 9.1822 |  |  |  |  | 1455 | 309.1843 | 323.9374 | 321.5176
2020-01-21 | 12:26 | 1.625 | 2.894 | 2.837 | 308.9941 | 9.3659 | 9.1814 |  |  |  |  | 1460 | 308.9941 | 323.7136 | 321.2911
2020-01-21 | 12:27 | 1.624 | 2.892 | 2.834 | 308.804 | 9.3652 | 9.1773 |  |  |  |  | 1465 | 308.804 | 323.4899 | 320.9513
2020-01-21 | 12:27 | 1.624 | 2.894 | 2.837 | 308.804 | 9.3716 | 9.1871 |  |  |  |  | 1470 | 308.804 | 323.7136 | 321.2911
2020-01-21 | 12:27 | 1.623 | 2.892 | 2.833 | 308.6138 | 9.3709 | 9.1798 |  |  |  |  | 1475 | 308.6138 | 323.4899 | 320.8381
2020-01-21 | 12:27 | 1.622 | 2.892 | 2.833 | 308.4237 | 9.3767 | 9.1854 |  |  |  |  | 1480 | 308.4237 | 323.4899 | 320.8381
2020-01-21 | 12:27 | 1.623 | 2.892 | 2.835 | 308.6138 | 9.3709 | 9.1862 |  |  |  |  | 1485 | 308.6138 | 323.4899 | 321.0646
2020-01-21 | 12:27 | 1.624 | 2.892 | 2.835 | 308.804 | 9.3652 | 9.1806 |  |  |  |  | 1490 | 308.804 | 323.4899 | 321.0646
2020-01-21 | 12:27 | 1.623 | 2.893 | 2.837 | 308.6138 | 9.3742 | 9.1927 |  |  |  |  | 1495 | 308.6138 | 323.6018 | 321.2911
2020-01-21 | 12:27 | 1.623 | 2.895 | 2.837 | 308.6138 | 9.3807 | 9.1927 |  |  |  |  | 1500 | 308.6138 | 323.8255 | 321.2911
2020-01-21 | 12:27 | 1.625 | 2.898 | 2.841 | 308.9941 | 9.3788 | 9.1944 |  |  |  |  | 1505 | 308.9941 | 324.1611 | 321.7441
2020-01-21 | 12:27 | 1.625 | 2.898 | 2.841 | 308.9941 | 9.3788 | 9.1944 |  |  |  |  | 1510 | 308.9941 | 324.1611 | 321.7441
2020-01-21 | 12:27 | 1.626 | 2.896 | 2.839 | 309.1843 | 9.3666 | 9.1822 |  |  |  |  | 1515 | 309.1843 | 323.9374 | 321.5176
2020-01-21 | 12:28 | 1.625 | 2.896 | 2.839 | 308.9941 | 9.3723 | 9.1879 |  |  |  |  | 1520 | 308.9941 | 323.9374 | 321.5176
2020-01-21 | 12:28 | 1.626 | 2.898 | 2.839 | 309.1843 | 9.3731 | 9.1822 |  |  |  |  | 1525 | 309.1843 | 324.1611 | 321.5176
2020-01-21 | 12:28 | 1.626 | 2.898 | 2.84 | 309.1843 | 9.3731 | 9.1855 |  |  |  |  | 1530 | 309.1843 | 324.1611 | 321.6308
2020-01-21 | 12:28 | 1.626 | 2.898 | 2.84 | 309.1843 | 9.3731 | 9.1855 |  |  |  |  | 1535 | 309.1843 | 324.1611 | 321.6308
2020-01-21 | 12:28 | 1.628 | 2.9 | 2.842 | 309.5646 | 9.368 | 9.1806 |  |  |  |  | 1540 | 309.5646 | 324.3848 | 321.8573
2020-01-21 | 12:28 | 1.627 | 2.901 | 2.842 | 309.3744 | 9.377 | 9.1863 |  |  |  |  | 1545 | 309.3744 | 324.4966 | 321.8573
2020-01-21 | 12:28 | 1.628 | 2.903 | 2.845 | 309.5646 | 9.3777 | 9.1903 |  |  |  |  | 1550 | 309.5646 | 324.7204 | 322.1971
2020-01-21 | 12:28 | 1.627 | 2.904 | 2.846 | 309.3744 | 9.3867 | 9.1992 |  |  |  |  | 1555 | 309.3744 | 324.8322 | 322.3103
2020-01-21 | 12:28 | 1.629 | 2.905 | 2.847 | 309.7547 | 9.3784 | 9.1911 |  |  |  |  | 1560 | 309.7547 | 324.9441 | 322.4236
2020-01-21 | 12:28 | 1.629 | 2.905 | 2.848 | 309.7547 | 9.3784 | 9.1944 |  |  |  |  | 1565 | 309.7547 | 324.9441 | 322.5368
2020-01-21 | 12:28 | 1.629 | 2.903 | 2.847 | 309.7547 | 9.3719 | 9.1911 |  |  |  |  | 1570 | 309.7547 | 324.7204 | 322.4236
2020-01-21 | 12:28 | 1.629 | 2.902 | 2.845 | 309.7547 | 9.3687 | 9.1847 |  |  |  |  | 1575 | 309.7547 | 324.6085 | 322.1971
2020-01-21 | 12:28 | 1.628 | 2.902 | 2.844 | 309.5646 | 9.3745 | 9.1871 |  |  |  |  | 1580 | 309.5646 | 324.6085 | 322.0838
2020-01-21 | 12:29 | 1.627 | 2.9 | 2.841 | 309.3744 | 9.3738 | 9.183 |  |  |  |  | 1585 | 309.3744 | 324.3848 | 321.7441
2020-01-21 | 12:29 | 1.627 | 2.902 | 2.843 | 309.3744 | 9.3802 | 9.1895 |  |  |  |  | 1590 | 309.3744 | 324.6085 | 321.9706
2020-01-21 | 12:29 | 1.627 | 2.899 | 2.84 | 309.3744 | 9.3705 | 9.1798 |  |  |  |  | 1595 | 309.3744 | 324.2729 | 321.6308
2020-01-21 | 12:29 | 1.627 | 2.898 | 2.839 | 309.3744 | 9.3673 | 9.1766 |  |  |  |  | 1600 | 309.3744 | 324.1611 | 321.5176
2020-01-21 | 12:29 | 1.625 | 2.896 | 2.837 | 308.9941 | 9.3723 | 9.1814 |  |  |  |  | 1605 | 308.9941 | 323.9374 | 321.2911
2020-01-21 | 12:29 | 1.624 | 2.895 | 2.835 | 308.804 | 9.3749 | 9.1806 |  |  |  |  | 1610 | 308.804 | 323.8255 | 321.0646
2020-01-21 | 12:29 | 1.623 | 2.894 | 2.836 | 308.6138 | 9.3774 | 9.1895 |  |  |  |  | 1615 | 308.6138 | 323.7136 | 321.1778
2020-01-21 | 12:29 | 1.624 | 2.896 | 2.837 | 308.804 | 9.3781 | 9.1871 |  |  |  |  | 1620 | 308.804 | 323.9374 | 321.2911
2020-01-21 | 12:29 | 1.626 | 2.899 | 2.839 | 309.1843 | 9.3763 | 9.1822 |  |  |  |  | 1625 | 309.1843 | 324.2729 | 321.5176
2020-01-21 | 12:29 | 1.625 | 2.899 | 2.84 | 308.9941 | 9.3821 | 9.1911 |  |  |  |  | 1630 | 308.9941 | 324.2729 | 321.6308
2020-01-21 | 12:29 | 1.626 | 2.9 | 2.843 | 309.1843 | 9.3795 | 9.1952 |  |  |  |  | 1635 | 309.1843 | 324.3848 | 321.9706
2020-01-21 | 12:30 | 1.628 | 2.9 | 2.842 | 309.5646 | 9.368 | 9.1806 |  |  |  |  | 1640 | 309.5646 | 324.3848 | 321.8573
2020-01-21 | 12:30 | 1.627 | 2.9 | 2.841 | 309.3744 | 9.3738 | 9.183 |  |  |  |  | 1645 | 309.3744 | 324.3848 | 321.7441
2020-01-21 | 12:30 | 1.627 | 2.901 | 2.843 | 309.3744 | 9.377 | 9.1895 |  |  |  |  | 1650 | 309.3744 | 324.4966 | 321.9706
2020-01-21 | 12:30 | 1.627 | 2.901 | 2.845 | 309.3744 | 9.377 | 9.196 |  |  |  |  | 1655 | 309.3744 | 324.4966 | 322.1971
2020-01-21 | 12:30 | 1.628 | 2.902 | 2.845 | 309.5646 | 9.3745 | 9.1903 |  |  |  |  | 1660 | 309.5646 | 324.6085 | 322.1971
2020-01-21 | 12:30 | 1.628 | 2.905 | 2.848 | 309.5646 | 9.3841 | 9.2 |  |  |  |  | 1665 | 309.5646 | 324.9441 | 322.5368
2020-01-21 | 12:30 | 1.628 | 2.904 | 2.847 | 309.5646 | 9.3809 | 9.1968 |  |  |  |  | 1670 | 309.5646 | 324.8322 | 322.4236
2020-01-21 | 12:30 | 1.628 | 2.901 | 2.843 | 309.5646 | 9.3712 | 9.1839 |  |  |  |  | 1675 | 309.5646 | 324.4966 | 321.9706
2020-01-21 | 12:30 | 1.626 | 2.901 | 2.843 | 309.1843 | 9.3828 | 9.1952 |  |  |  |  | 1680 | 309.1843 | 324.4966 | 321.9706
2020-01-21 | 12:30 | 1.625 | 2.9 | 2.842 | 308.9941 | 9.3853 | 9.1976 |  |  |  |  | 1685 | 308.9941 | 324.3848 | 321.8573
2020-01-21 | 12:30 | 1.625 | 2.901 | 2.844 | 308.9941 | 9.3885 | 9.2041 |  |  |  |  | 1690 | 308.9941 | 324.4966 | 322.0838
2020-01-21 | 12:30 | 1.626 | 2.902 | 2.844 | 309.1843 | 9.386 | 9.1984 |  |  |  |  | 1695 | 309.1843 | 324.6085 | 322.0838
2020-01-21 | 12:30 | 1.628 | 2.904 | 2.847 | 309.5646 | 9.3809 | 9.1968 |  |  |  |  | 1700 | 309.5646 | 324.8322 | 322.4236
2020-01-21 | 12:31 | 1.627 | 2.904 | 2.846 | 309.3744 | 9.3867 | 9.1992 |  |  |  |  | 1705 | 309.3744 | 324.8322 | 322.3103
2020-01-21 | 12:31 | 1.629 | 2.906 | 2.849 | 309.7547 | 9.3816 | 9.1976 |  |  |  |  | 1710 | 309.7547 | 325.0559 | 322.6501
2020-01-21 | 12:31 | 1.63 | 2.909 | 2.851 | 309.9449 | 9.3855 | 9.1984 |  |  |  |  | 1715 | 309.9449 | 325.3915 | 322.8766
2020-01-21 | 12:31 | 1.631 | 2.908 | 2.85 | 310.135 | 9.3766 | 9.1895 |  |  |  |  | 1720 | 310.135 | 325.2796 | 322.7633
2020-01-21 | 12:31 | 1.629 | 2.906 | 2.846 | 309.7547 | 9.3816 | 9.1879 |  |  |  |  | 1725 | 309.7547 | 325.0559 | 322.3103
2020-01-21 | 12:31 | 1.629 | 2.907 | 2.847 | 309.7547 | 9.3848 | 9.1911 |  |  |  |  | 1730 | 309.7547 | 325.1678 | 322.4236
2020-01-21 | 12:31 | 1.629 | 2.907 | 2.848 | 309.7547 | 9.3848 | 9.1944 |  |  |  |  | 1735 | 309.7547 | 325.1678 | 322.5368
2020-01-21 | 12:31 | 1.629 | 2.904 | 2.847 | 309.7547 | 9.3752 | 9.1911 |  |  |  |  | 1740 | 309.7547 | 324.8322 | 322.4236
2020-01-21 | 12:31 | 1.631 | 2.905 | 2.847 | 310.135 | 9.3669 | 9.1799 |  |  |  |  | 1745 | 310.135 | 324.9441 | 322.4236
2020-01-21 | 12:31 | 1.63 | 2.903 | 2.845 | 309.9449 | 9.3662 | 9.1791 |  |  |  |  | 1750 | 309.9449 | 324.7204 | 322.1971
2020-01-21 | 12:31 | 1.629 | 2.903 | 2.844 | 309.7547 | 9.3719 | 9.1815 |  |  |  |  | 1755 | 309.7547 | 324.7204 | 322.0838
2020-01-21 | 12:32 | 1.627 | 2.9 | 2.841 | 309.3744 | 9.3738 | 9.183 |  |  |  |  | 1760 | 309.3744 | 324.3848 | 321.7441
2020-01-21 | 12:32 | 1.628 | 2.901 | 2.843 | 309.5646 | 9.3712 | 9.1839 |  |  |  |  | 1765 | 309.5646 | 324.4966 | 321.9706
2020-01-21 | 12:32 | 1.628 | 2.899 | 2.84 | 309.5646 | 9.3648 | 9.1742 |  |  |  |  | 1770 | 309.5646 | 324.2729 | 321.6308
2020-01-21 | 12:32 | 1.627 | 2.899 | 2.84 | 309.3744 | 9.3705 | 9.1798 |  |  |  |  | 1775 | 309.3744 | 324.2729 | 321.6308
2020-01-21 | 12:32 | 1.627 | 2.899 | 2.841 | 309.3744 | 9.3705 | 9.183 |  |  |  |  | 1780 | 309.3744 | 324.2729 | 321.7441
2020-01-21 | 12:32 | 1.625 | 2.897 | 2.839 | 308.9941 | 9.3756 | 9.1879 |  |  |  |  | 1785 | 308.9941 | 324.0492 | 321.5176
2020-01-21 | 12:32 | 1.624 | 2.895 | 2.838 | 308.804 | 9.3749 | 9.1903 |  |  |  |  | 1790 | 308.804 | 323.8255 | 321.4043
2020-01-21 | 12:32 | 1.624 | 2.893 | 2.835 | 308.804 | 9.3684 | 9.1806 |  |  |  |  | 1795 | 308.804 | 323.6018 | 321.0646
2020-01-21 | 12:32 | 1.622 | 2.89 | 2.833 | 308.4237 | 9.3702 | 9.1854 |  |  |  |  | 1800 | 308.4237 | 323.2662 | 320.8381
2020-01-21 | 12:32 | 1.62 | 2.885 | 2.828 | 308.0434 | 9.3656 | 9.1805 |  |  |  |  | 1805 | 308.0434 | 322.7069 | 320.2718
2020-01-21 | 12:32 | 1.618 | 2.883 | 2.828 | 307.6631 | 9.3706 | 9.1919 |  |  |  |  | 1810 | 307.6631 | 322.4832 | 320.2718
2020-01-21 | 12:32 | 1.617 | 2.884 | 2.828 | 307.4729 | 9.3797 | 9.1976 |  |  |  |  | 1815 | 307.4729 | 322.5951 | 320.2718
2020-01-21 | 12:32 | 1.618 | 2.883 | 2.829 | 307.6631 | 9.3706 | 9.1951 |  |  |  |  | 1820 | 307.6631 | 322.4832 | 320.3851
2020-01-21 | 12:33 | 1.616 | 2.883 | 2.828 | 307.2828 | 9.3822 | 9.2033 |  |  |  |  | 1825 | 307.2828 | 322.4832 | 320.2718
2020-01-21 | 12:33 | 1.616 | 2.882 | 2.826 | 307.2828 | 9.379 | 9.1967 |  |  |  |  | 1830 | 307.2828 | 322.3714 | 320.0453
2020-01-21 | 12:33 | 1.615 | 2.88 | 2.824 | 307.0926 | 9.3783 | 9.1959 |  |  |  |  | 1835 | 307.0926 | 322.1477 | 319.8188
2020-01-21 | 12:33 | 1.615 | 2.877 | 2.821 | 307.0926 | 9.3685 | 9.1862 |  |  |  |  | 1840 | 307.0926 | 321.8121 | 319.479
2020-01-21 | 12:33 | 1.614 | 2.878 | 2.824 | 306.9025 | 9.3776 | 9.2016 |  |  |  |  | 1845 | 306.9025 | 321.9239 | 319.8188
2020-01-21 | 12:33 | 1.613 | 2.875 | 2.819 | 306.7123 | 9.3736 | 9.191 |  |  |  |  | 1850 | 306.7123 | 321.5884 | 319.2525
2020-01-21 | 12:33 | 1.611 | 2.872 | 2.816 | 306.332 | 9.3754 | 9.1926 |  |  |  |  | 1855 | 306.332 | 321.2528 | 318.9128
2020-01-21 | 12:33 | 1.613 | 2.872 | 2.817 | 306.7123 | 9.3638 | 9.1845 |  |  |  |  | 1860 | 306.7123 | 321.2528 | 319.026
2020-01-21 | 12:33 | 1.613 | 2.874 | 2.817 | 306.7123 | 9.3703 | 9.1845 |  |  |  |  | 1865 | 306.7123 | 321.4765 | 319.026
2020-01-21 | 12:33 | 1.613 | 2.875 | 2.816 | 306.7123 | 9.3736 | 9.1812 |  |  |  |  | 1870 | 306.7123 | 321.5884 | 318.9128
2020-01-21 | 12:33 | 1.613 | 2.875 | 2.818 | 306.7123 | 9.3736 | 9.1878 |  |  |  |  | 1875 | 306.7123 | 321.5884 | 319.1393
2020-01-21 | 12:34 | 1.614 | 2.877 | 2.82 | 306.9025 | 9.3743 | 9.1886 |  |  |  |  | 1880 | 306.9025 | 321.8121 | 319.3658
2020-01-21 | 12:34 | 1.615 | 2.879 | 2.821 | 307.0926 | 9.375 | 9.1862 |  |  |  |  | 1885 | 307.0926 | 322.0358 | 319.479
2020-01-21 | 12:34 | 1.616 | 2.88 | 2.823 | 307.2828 | 9.3725 | 9.187 |  |  |  |  | 1890 | 307.2828 | 322.1477 | 319.7055
2020-01-21 | 12:34 | 1.618 | 2.884 | 2.827 | 307.6631 | 9.3739 | 9.1886 |  |  |  |  | 1895 | 307.6631 | 322.5951 | 320.1586
2020-01-21 | 12:34 | 1.621 | 2.888 | 2.832 | 308.2335 | 9.3695 | 9.1878 |  |  |  |  | 1900 | 308.2335 | 323.0425 | 320.7248
2020-01-21 | 12:34 | 1.623 | 2.895 | 2.837 | 308.6138 | 9.3807 | 9.1927 |  |  |  |  | 1905 | 308.6138 | 323.8255 | 321.2911
2020-01-21 | 12:34 | 1.626 | 2.9 | 2.843 | 309.1843 | 9.3795 | 9.1952 |  |  |  |  | 1910 | 309.1843 | 324.3848 | 321.9706
2020-01-21 | 12:34 | 1.629 | 2.907 | 2.849 | 309.7547 | 9.3848 | 9.1976 |  |  |  |  | 1915 | 309.7547 | 325.1678 | 322.6501
2020-01-21 | 12:34 | 1.632 | 2.911 | 2.853 | 310.3252 | 9.3805 | 9.1936 |  |  |  |  | 1920 | 310.3252 | 325.6152 | 323.1031
2020-01-21 | 12:34 | 1.634 | 2.914 | 2.858 | 310.7055 | 9.3787 | 9.1984 |  |  |  |  | 1925 | 310.7055 | 325.9508 | 323.6693
2020-01-21 | 12:34 | 1.636 | 2.919 | 2.862 | 311.0858 | 9.3833 | 9.2 |  |  |  |  | 1930 | 311.0858 | 326.5101 | 324.1223
2020-01-21 | 12:34 | 1.637 | 2.917 | 2.861 | 311.2759 | 9.3711 | 9.1912 |  |  |  |  | 1935 | 311.2759 | 326.2864 | 324.0091
2020-01-21 | 12:34 | 1.637 | 2.919 | 2.862 | 311.2759 | 9.3775 | 9.1944 |  |  |  |  | 1940 | 311.2759 | 326.5101 | 324.1223
2020-01-21 | 12:35 | 1.638 | 2.92 | 2.864 | 311.4661 | 9.375 | 9.1952 |  |  |  |  | 1945 | 311.4661 | 326.6219 | 324.3488
2020-01-21 | 12:35 | 1.638 | 2.919 | 2.863 | 311.4661 | 9.3718 | 9.192 |  |  |  |  | 1950 | 311.4661 | 326.5101 | 324.2356
2020-01-21 | 12:35 | 1.638 | 2.92 | 2.864 | 311.4661 | 9.375 | 9.1952 |  |  |  |  | 1955 | 311.4661 | 326.6219 | 324.3488
2020-01-21 | 12:35 | 1.638 | 2.921 | 2.864 | 311.4661 | 9.3782 | 9.1952 |  |  |  |  | 1960 | 311.4661 | 326.7338 | 324.3488
2020-01-21 | 12:35 | 1.639 | 2.921 | 2.863 | 311.6562 | 9.3725 | 9.1864 |  |  |  |  | 1965 | 311.6562 | 326.7338 | 324.2356
2020-01-21 | 12:35 | 1.639 | 2.921 | 2.861 | 311.6562 | 9.3725 | 9.18 |  |  |  |  | 1970 | 311.6562 | 326.7338 | 324.0091
Zeitspanne: | 0:33 |  |  |  | 304.9186 | 9.3783 | 9.1833 |  |  |  |  | 1970

## Measurements - Akwatia
Albedo Measurements
[7 empty] | Cloudless time periods |  | Average Albedo
Site: | Akwatia St. Dominics Hospital | [5 empty] | 12:06 | 12:16 | 0.20497
Location: | Latitude: 6.047049°; Longitude: -0.8005275° | [5 empty] | 12:55 | 13:02 | 0.20552 | Inclination
Date: | 06.02.2020; 11:16pm-1:16pm | [7 empty] | 0.20524 | 2.577%
|  | Top Pyranometer |  |  |  | Bottom Pyranometer
|  | Min. Value [V] | Max. Value [V] | Average Value [V] | Irradiance [W/m^2] | Min. Value [V] | Max. Value [V] | Average Value [V] | Irradiance [W/m^2] | Albedo | Average Albedo: | 0.19995
2020-02-06 | 11:16 | 8.173 | 8.181 | 8.177 | 871.9072 | 1.68 | 1.682 | 1.681 | 183.05 | 0.20994
2020-02-06 | 11:16 | 8.189 | 8.21 | 8.1995 | 874.3063 | 1.684 | 1.687 | 1.6855 | 183.54 | 0.20993
2020-02-06 | 11:16 | 8.213 | 8.221 | 8.217 | 876.1723 | 1.687 | 1.689 | 1.688 | 183.8122 | 0.20979
2020-02-06 | 11:16 | 8.224 | 8.229 | 8.2265 | 877.1853 | 1.689 | 1.691 | 1.69 | 184.03 | 0.2098
2020-02-06 | 11:16 | 8.232 | 8.233 | 8.2325 | 877.8251 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.20977
2020-02-06 | 11:16 | 8.222 | 8.231 | 8.2265 | 877.1853 | 1.689 | 1.69 | 1.6895 | 183.9756 | 0.20973
2020-02-06 | 11:16 | 8.208 | 8.218 | 8.213 | 875.7458 | 1.686 | 1.689 | 1.6875 | 183.7578 | 0.20983
2020-02-06 | 11:17 | 8.19 | 8.204 | 8.197 | 874.0397 | 1.682 | 1.686 | 1.684 | 183.3767 | 0.2098
2020-02-06 | 11:17 | 8.19 | 8.192 | 8.191 | 873.4 | 1.682 | 1.682 | 1.682 | 183.1589 | 0.20971
2020-02-06 | 11:17 | 8.181 | 8.191 | 8.186 | 872.8668 | 1.679 | 1.682 | 1.6805 | 182.9955 | 0.20965
2020-02-06 | 11:17 | 8.17 | 8.178 | 8.174 | 871.5873 | 1.677 | 1.68 | 1.6785 | 182.7777 | 0.20971
2020-02-06 | 11:17 | 8.159 | 8.166 | 8.1625 | 870.361 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.20963
2020-02-06 | 11:17 | 8.154 | 8.158 | 8.156 | 869.6679 | 1.672 | 1.674 | 1.673 | 182.1788 | 0.20948
2020-02-06 | 11:17 | 8.144 | 8.151 | 8.1475 | 868.7616 | 1.67 | 1.673 | 1.6715 | 182.0155 | 0.20951
2020-02-06 | 11:17 | 8.104 | 8.14 | 8.122 | 866.0425 | 1.658 | 1.668 | 1.663 | 181.0899 | 0.2091
2020-02-06 | 11:17 | 7.238 | 8.092 | 7.665 | 817.313 | 1.449 | 1.656 | 1.5525 | 169.0572 | 0.20685
2020-02-06 | 11:17 | 5.768 | 6.986 | 6.377 | 679.9746 | 1.095 | 1.384 | 1.2395 | 134.9735 | 0.1985
2020-02-06 | 11:17 | 4.126 | 5.316 | 4.721 | 503.3966 | 0.741 | 0.977 | 0.859 | 93.5395 | 0.18582
2020-02-06 | 11:18 | 4.13 | 4.253 | 4.1915 | 446.9364 | 0.744 | 0.775 | 0.7595 | 82.7046 | 0.18505
2020-02-06 | 11:18 | 4.317 | 5.345 | 4.831 | 515.1258 | 0.789 | 1.031 | 0.91 | 99.0931 | 0.19237
2020-02-06 | 11:18 | 4.296 | 5.391 | 4.8435 | 516.4586 | 0.773 | 1.022 | 0.8975 | 97.7319 | 0.18923
2020-02-06 | 11:18 | 3.855 | 4.168 | 4.0115 | 427.7431 | 0.688 | 0.747 | 0.7175 | 78.1311 | 0.18266
2020-02-06 | 11:18 | 3.793 | 3.831 | 3.812 | 406.4706 | 0.68 | 0.685 | 0.6825 | 74.3198 | 0.18284
2020-02-06 | 11:18 | 3.751 | 3.786 | 3.7685 | 401.8322 | 0.673 | 0.679 | 0.676 | 73.612 | 0.18319
2020-02-06 | 11:18 | 3.72 | 3.743 | 3.7315 | 397.8869 | 0.668 | 0.672 | 0.67 | 72.9586 | 0.18337
2020-02-06 | 11:18 | 3.702 | 3.714 | 3.708 | 395.3812 | 0.664 | 0.668 | 0.666 | 72.5231 | 0.18343
2020-02-06 | 11:18 | 3.696 | 3.7 | 3.698 | 394.3149 | 0.664 | 0.666 | 0.665 | 72.4142 | 0.18365
2020-02-06 | 11:18 | 3.698 | 3.709 | 3.7035 | 394.9013 | 0.666 | 0.668 | 0.667 | 72.632 | 0.18392
2020-02-06 | 11:18 | 3.712 | 3.734 | 3.723 | 396.9806 | 0.669 | 0.674 | 0.6715 | 73.122 | 0.1842
2020-02-06 | 11:18 | 3.742 | 3.781 | 3.7615 | 401.0858 | 0.674 | 0.683 | 0.6785 | 73.8842 | 0.18421
2020-02-06 | 11:18 | 3.792 | 3.885 | 3.8385 | 409.2963 | 0.686 | 0.708 | 0.697 | 75.8988 | 0.18544
2020-02-06 | 11:19 | 3.945 | 4.67 | 4.3075 | 459.3054 | 0.723 | 0.892 | 0.8075 | 87.9315 | 0.19144
2020-02-06 | 11:19 | 4.962 | 7.046 | 6.004 | 640.2019 | 0.961 | 1.422 | 1.1915 | 129.7466 | 0.20267
2020-02-06 | 11:19 | 5.638 | 7.163 | 6.4005 | 682.4803 | 1.099 | 1.427 | 1.263 | 137.5325 | 0.20152
2020-02-06 | 11:19 | 5.961 | 7.803 | 6.882 | 733.8223 | 1.193 | 1.613 | 1.403 | 152.7776 | 0.20819
2020-02-06 | 11:19 | 7.953 | 8.1 | 8.0265 | 855.8595 | 1.643 | 1.67 | 1.6565 | 180.3821 | 0.21076
2020-02-06 | 11:19 | 8.108 | 8.129 | 8.1185 | 865.6693 | 1.672 | 1.674 | 1.673 | 182.1788 | 0.21045
2020-02-06 | 11:19 | 8.134 | 8.148 | 8.141 | 868.0685 | 1.674 | 1.677 | 1.6755 | 182.4511 | 0.21018
2020-02-06 | 11:19 | 8.151 | 8.169 | 8.16 | 870.0945 | 1.676 | 1.679 | 1.6775 | 182.6688 | 0.20994
2020-02-06 | 11:19 | 8.164 | 8.169 | 8.1665 | 870.7875 | 1.677 | 1.68 | 1.6785 | 182.7777 | 0.2099
2020-02-06 | 11:19 | 8.169 | 8.177 | 8.173 | 871.4806 | 1.678 | 1.681 | 1.6795 | 182.8866 | 0.20986
2020-02-06 | 11:19 | 8.175 | 8.178 | 8.1765 | 871.8538 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.20977
2020-02-06 | 11:20 | 8.179 | 8.182 | 8.1805 | 872.2804 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.20966
2020-02-06 | 11:20 | 8.178 | 8.18 | 8.179 | 872.1204 | 1.677 | 1.679 | 1.678 | 182.7233 | 0.20952
2020-02-06 | 11:20 | 8.173 | 8.181 | 8.177 | 871.9072 | 1.675 | 1.679 | 1.677 | 182.6144 | 0.20944
2020-02-06 | 11:20 | 8.149 | 8.17 | 8.1595 | 870.0411 | 1.668 | 1.674 | 1.671 | 181.961 | 0.20914
2020-02-06 | 11:20 | 6.598 | 8.139 | 7.3685 | 785.6974 | 1.238 | 1.666 | 1.452 | 158.1134 | 0.20124
2020-02-06 | 11:20 | 3.868 | 5.31 | 4.589 | 489.3215 | 0.687 | 0.957 | 0.822 | 89.5105 | 0.18293
2020-02-06 | 11:20 | 3.73 | 3.82 | 3.775 | 402.5253 | 0.667 | 0.681 | 0.674 | 73.3942 | 0.18233
2020-02-06 | 11:20 | 3.662 | 3.716 | 3.689 | 393.3552 | 0.656 | 0.664 | 0.66 | 71.8697 | 0.18271
2020-02-06 | 11:20 | 3.611 | 3.652 | 3.6315 | 387.224 | 0.646 | 0.655 | 0.6505 | 70.8352 | 0.18293
2020-02-06 | 11:20 | 3.566 | 3.6 | 3.583 | 382.0525 | 0.639 | 0.647 | 0.643 | 70.0185 | 0.18327
2020-02-06 | 11:20 | 3.535 | 3.56 | 3.5475 | 378.2672 | 0.633 | 0.639 | 0.636 | 69.2563 | 0.18309
2020-02-06 | 11:20 | 3.509 | 3.53 | 3.5195 | 375.2816 | 0.628 | 0.632 | 0.63 | 68.6029 | 0.1828
2020-02-06 | 11:20 | 3.485 | 3.504 | 3.4945 | 372.6158 | 0.624 | 0.628 | 0.626 | 68.1673 | 0.18294
2020-02-06 | 11:21 | 3.463 | 3.481 | 3.472 | 370.2167 | 0.62 | 0.624 | 0.622 | 67.7318 | 0.18295
2020-02-06 | 11:21 | 3.439 | 3.457 | 3.448 | 367.6576 | 0.616 | 0.619 | 0.6175 | 67.2417 | 0.18289
2020-02-06 | 11:21 | 3.418 | 3.436 | 3.427 | 365.4183 | 0.612 | 0.616 | 0.614 | 66.8606 | 0.18297
2020-02-06 | 11:21 | 3.396 | 3.413 | 3.4045 | 363.0192 | 0.609 | 0.612 | 0.6105 | 66.4795 | 0.18313
2020-02-06 | 11:21 | 3.375 | 3.391 | 3.383 | 360.7267 | 0.605 | 0.608 | 0.6065 | 66.0439 | 0.18309
2020-02-06 | 11:21 | 3.353 | 3.371 | 3.362 | 358.4874 | 0.601 | 0.605 | 0.603 | 65.6628 | 0.18317
2020-02-06 | 11:21 | 3.328 | 3.348 | 3.338 | 355.9283 | 0.597 | 0.6 | 0.5985 | 65.1728 | 0.18311
2020-02-06 | 11:21 | 3.307 | 3.325 | 3.316 | 353.5825 | 0.594 | 0.596 | 0.595 | 64.7916 | 0.18324
2020-02-06 | 11:21 | 3.291 | 3.304 | 3.2975 | 351.6099 | 0.59 | 0.592 | 0.591 | 64.3561 | 0.18303
2020-02-06 | 11:21 | 3.272 | 3.288 | 3.28 | 349.7439 | 0.586 | 0.59 | 0.588 | 64.0294 | 0.18308
2020-02-06 | 11:21 | 3.256 | 3.269 | 3.2625 | 347.8778 | 0.583 | 0.586 | 0.5845 | 63.6483 | 0.18296
2020-02-06 | 11:22 | 3.238 | 3.252 | 3.245 | 346.0118 | 0.58 | 0.584 | 0.582 | 63.376 | 0.18316
2020-02-06 | 11:22 | 3.225 | 3.235 | 3.23 | 344.4124 | 0.58 | 0.582 | 0.581 | 63.2671 | 0.1837
2020-02-06 | 11:22 | 3.212 | 3.221 | 3.2165 | 342.9729 | 0.578 | 0.58 | 0.579 | 63.0493 | 0.18383
2020-02-06 | 11:22 | 3.201 | 3.209 | 3.205 | 341.7467 | 0.576 | 0.578 | 0.577 | 62.8316 | 0.18385
2020-02-06 | 11:22 | 3.191 | 3.199 | 3.195 | 340.6804 | 0.574 | 0.576 | 0.575 | 62.6138 | 0.18379
2020-02-06 | 11:22 | 3.179 | 3.189 | 3.184 | 339.5074 | 0.573 | 0.574 | 0.5735 | 62.4504 | 0.18394
2020-02-06 | 11:22 | 3.165 | 3.177 | 3.171 | 338.1213 | 0.571 | 0.573 | 0.572 | 62.2871 | 0.18422
2020-02-06 | 11:22 | 3.152 | 3.163 | 3.1575 | 336.6818 | 0.569 | 0.57 | 0.5695 | 62.0148 | 0.18419
2020-02-06 | 11:22 | 3.139 | 3.15 | 3.1445 | 335.2956 | 0.568 | 0.569 | 0.5685 | 61.906 | 0.18463
2020-02-06 | 11:22 | 3.123 | 3.137 | 3.13 | 333.7495 | 0.565 | 0.567 | 0.566 | 61.6337 | 0.18467
2020-02-06 | 11:22 | 3.103 | 3.119 | 3.111 | 331.7235 | 0.561 | 0.564 | 0.5625 | 61.2526 | 0.18465
2020-02-06 | 11:22 | 3.082 | 3.1 | 3.091 | 329.5909 | 0.556 | 0.56 | 0.558 | 60.7626 | 0.18436
2020-02-06 | 11:22 | 3.066 | 3.079 | 3.0725 | 327.6183 | 0.554 | 0.556 | 0.555 | 60.4359 | 0.18447
2020-02-06 | 11:23 | 3.05 | 3.064 | 3.057 | 325.9655 | 0.552 | 0.554 | 0.553 | 60.2181 | 0.18474
2020-02-06 | 11:23 | 3.034 | 3.047 | 3.0405 | 324.2062 | 0.549 | 0.553 | 0.551 | 60.0003 | 0.18507
2020-02-06 | 11:23 | 3.022 | 3.032 | 3.027 | 322.7667 | 0.546 | 0.548 | 0.547 | 59.5647 | 0.18454
2020-02-06 | 11:23 | 3.01 | 3.018 | 3.014 | 321.3805 | 0.544 | 0.546 | 0.545 | 59.347 | 0.18466
2020-02-06 | 11:23 | 3.003 | 3.008 | 3.0055 | 320.4741 | 0.542 | 0.543 | 0.5425 | 59.0747 | 0.18434
2020-02-06 | 11:23 | 3.001 | 3.002 | 3.0015 | 320.0476 | 0.541 | 0.543 | 0.542 | 59.0203 | 0.18441
2020-02-06 | 11:23 | 3.003 | 3.006 | 3.0045 | 320.3675 | 0.542 | 0.543 | 0.5425 | 59.0747 | 0.1844
2020-02-06 | 11:23 | 3.006 | 3.011 | 3.0085 | 320.794 | 0.543 | 0.544 | 0.5435 | 59.1836 | 0.18449
2020-02-06 | 11:23 | 3.012 | 3.017 | 3.0145 | 321.4338 | 0.544 | 0.545 | 0.5445 | 59.2925 | 0.18446
2020-02-06 | 11:23 | 3.018 | 3.02 | 3.019 | 321.9136 | 0.545 | 0.546 | 0.5455 | 59.4014 | 0.18453
2020-02-06 | 11:23 | 3.02 | 3.021 | 3.0205 | 322.0736 | 0.544 | 0.546 | 0.545 | 59.347 | 0.18427
2020-02-06 | 11:24 | 3.018 | 3.02 | 3.019 | 321.9136 | 0.544 | 0.546 | 0.545 | 59.347 | 0.18436
2020-02-06 | 11:24 | 3.013 | 3.017 | 3.015 | 321.4871 | 0.543 | 0.545 | 0.544 | 59.2381 | 0.18426
2020-02-06 | 11:24 | 3.009 | 3.012 | 3.0105 | 321.0073 | 0.542 | 0.543 | 0.5425 | 59.0747 | 0.18403
2020-02-06 | 11:24 | 3.009 | 3.01 | 3.0095 | 320.9006 | 0.542 | 0.543 | 0.5425 | 59.0747 | 0.18409
2020-02-06 | 11:24 | 3.008 | 3.01 | 3.009 | 320.8473 | 0.541 | 0.543 | 0.542 | 59.0203 | 0.18395
2020-02-06 | 11:24 | 3.003 | 3.007 | 3.005 | 320.4208 | 0.539 | 0.543 | 0.541 | 58.9114 | 0.18386
2020-02-06 | 11:24 | 2.997 | 3.001 | 2.999 | 319.781 | 0.538 | 0.54 | 0.539 | 58.6936 | 0.18354
2020-02-06 | 11:24 | 2.991 | 2.997 | 2.994 | 319.2479 | 0.537 | 0.538 | 0.5375 | 58.5303 | 0.18334
2020-02-06 | 11:24 | 2.982 | 2.99 | 2.986 | 318.3949 | 0.534 | 0.536 | 0.535 | 58.258 | 0.18297
2020-02-06 | 11:24 | 2.974 | 2.981 | 2.9775 | 317.4885 | 0.532 | 0.534 | 0.533 | 58.0402 | 0.18281
2020-02-06 | 11:24 | 2.97 | 2.973 | 2.9715 | 316.8487 | 0.53 | 0.532 | 0.531 | 57.8224 | 0.18249
2020-02-06 | 11:24 | 2.966 | 2.968 | 2.967 | 316.3689 | 0.529 | 0.53 | 0.5295 | 57.6591 | 0.18225
2020-02-06 | 11:24 | 2.966 | 2.968 | 2.967 | 316.3689 | 0.528 | 0.53 | 0.529 | 57.6047 | 0.18208
2020-02-06 | 11:25 | 2.969 | 2.97 | 2.9695 | 316.6355 | 0.529 | 0.53 | 0.5295 | 57.6591 | 0.1821
2020-02-06 | 11:25 | 2.971 | 2.974 | 2.9725 | 316.9554 | 0.529 | 0.53 | 0.5295 | 57.6591 | 0.18192
2020-02-06 | 11:25 | 2.976 | 2.981 | 2.9785 | 317.5951 | 0.531 | 0.532 | 0.5315 | 57.8769 | 0.18223
2020-02-06 | 11:25 | 2.981 | 2.988 | 2.9845 | 318.2349 | 0.532 | 0.533 | 0.5325 | 57.9858 | 0.18221
2020-02-06 | 11:25 | 2.989 | 2.999 | 2.994 | 319.2479 | 0.533 | 0.534 | 0.5335 | 58.0947 | 0.18197
2020-02-06 | 11:25 | 3.001 | 3.014 | 3.0075 | 320.6874 | 0.536 | 0.538 | 0.537 | 58.4758 | 0.18235
2020-02-06 | 11:25 | 3.017 | 3.033 | 3.025 | 322.5534 | 0.539 | 0.541 | 0.54 | 58.8025 | 0.1823
2020-02-06 | 11:25 | 3.036 | 3.052 | 3.044 | 324.5794 | 0.543 | 0.545 | 0.544 | 59.2381 | 0.18251
2020-02-06 | 11:25 | 3.055 | 3.074 | 3.0645 | 326.7653 | 0.546 | 0.551 | 0.5485 | 59.7281 | 0.18279
2020-02-06 | 11:25 | 3.075 | 3.093 | 3.084 | 328.8445 | 0.549 | 0.554 | 0.5515 | 60.0548 | 0.18262
2020-02-06 | 11:25 | 3.097 | 3.114 | 3.1055 | 331.1371 | 0.554 | 0.558 | 0.556 | 60.5448 | 0.18284
2020-02-06 | 11:26 | 3.118 | 3.135 | 3.1265 | 333.3763 | 0.557 | 0.562 | 0.5595 | 60.9259 | 0.18275
2020-02-06 | 11:26 | 3.14 | 3.159 | 3.1495 | 335.8287 | 0.561 | 0.566 | 0.5635 | 61.3615 | 0.18272
2020-02-06 | 11:26 | 3.164 | 3.185 | 3.1745 | 338.4945 | 0.566 | 0.57 | 0.568 | 61.8515 | 0.18273
2020-02-06 | 11:26 | 3.191 | 3.217 | 3.204 | 341.64 | 0.572 | 0.577 | 0.5745 | 62.5593 | 0.18311
2020-02-06 | 11:26 | 3.224 | 3.252 | 3.238 | 345.2654 | 0.578 | 0.584 | 0.581 | 63.2671 | 0.18324
2020-02-06 | 11:26 | 3.259 | 3.29 | 3.2745 | 349.1574 | 0.585 | 0.591 | 0.588 | 64.0294 | 0.18338
2020-02-06 | 11:26 | 3.298 | 3.333 | 3.3155 | 353.5292 | 0.592 | 0.598 | 0.595 | 64.7916 | 0.18327
2020-02-06 | 11:26 | 3.342 | 3.382 | 3.362 | 358.4874 | 0.6 | 0.608 | 0.604 | 65.7717 | 0.18347
2020-02-06 | 11:26 | 3.392 | 3.481 | 3.4365 | 366.4313 | 0.61 | 0.629 | 0.6195 | 67.4595 | 0.1841
2020-02-06 | 11:26 | 3.504 | 3.593 | 3.5485 | 378.3738 | 0.633 | 0.657 | 0.645 | 70.2363 | 0.18563
2020-02-06 | 11:26 | 3.744 | 5.197 | 4.4705 | 476.6859 | 0.701 | 1.036 | 0.8685 | 94.574 | 0.1984
2020-02-06 | 11:26 | 5.563 | 7.064 | 6.3135 | 673.2036 | 1.124 | 1.469 | 1.2965 | 141.1804 | 0.20971
2020-02-06 | 11:26 | 6.916 | 7.366 | 7.141 | 761.4393 | 1.401 | 1.521 | 1.461 | 159.0934 | 0.20894
2020-02-06 | 11:27 | 4.797 | 6.718 | 5.7575 | 613.9178 | 0.896 | 1.351 | 1.1235 | 122.3418 | 0.19928
2020-02-06 | 11:27 | 3.564 | 4.318 | 3.941 | 420.2258 | 0.636 | 0.791 | 0.7135 | 77.6955 | 0.18489
2020-02-06 | 11:27 | 3.502 | 3.533 | 3.5175 | 375.0683 | 0.629 | 0.633 | 0.631 | 68.7118 | 0.1832
2020-02-06 | 11:27 | 3.47 | 3.491 | 3.4805 | 371.123 | 0.623 | 0.626 | 0.6245 | 68.004 | 0.18324
2020-02-06 | 11:27 | 3.455 | 3.467 | 3.461 | 369.0437 | 0.62 | 0.622 | 0.621 | 67.6229 | 0.18324
2020-02-06 | 11:27 | 3.443 | 3.452 | 3.4475 | 367.6042 | 0.618 | 0.62 | 0.619 | 67.4051 | 0.18336
2020-02-06 | 11:27 | 3.433 | 3.441 | 3.437 | 366.4846 | 0.617 | 0.618 | 0.6175 | 67.2417 | 0.18348
2020-02-06 | 11:27 | 3.425 | 3.432 | 3.4285 | 365.5783 | 0.613 | 0.616 | 0.6145 | 66.9151 | 0.18304
2020-02-06 | 11:27 | 3.417 | 3.423 | 3.42 | 364.6719 | 0.611 | 0.614 | 0.6125 | 66.6973 | 0.1829
2020-02-06 | 11:27 | 3.408 | 3.415 | 3.4115 | 363.7656 | 0.609 | 0.611 | 0.61 | 66.425 | 0.1826
2020-02-06 | 11:27 | 3.402 | 3.406 | 3.404 | 362.9659 | 0.608 | 0.609 | 0.6085 | 66.2617 | 0.18256
2020-02-06 | 11:28 | 3.398 | 3.401 | 3.3995 | 362.486 | 0.606 | 0.608 | 0.607 | 66.0984 | 0.18235
2020-02-06 | 11:28 | 3.397 | 3.398 | 3.3975 | 362.2728 | 0.605 | 0.607 | 0.606 | 65.9895 | 0.18215
2020-02-06 | 11:28 | 3.396 | 3.397 | 3.3965 | 362.1662 | 0.605 | 0.606 | 0.6055 | 65.935 | 0.18206
2020-02-06 | 11:28 | 3.396 | 3.397 | 3.3965 | 362.1662 | 0.605 | 0.606 | 0.6055 | 65.935 | 0.18206
2020-02-06 | 11:28 | 3.398 | 3.402 | 3.4 | 362.5394 | 0.605 | 0.607 | 0.606 | 65.9895 | 0.18202
2020-02-06 | 11:28 | 3.403 | 3.408 | 3.4055 | 363.1258 | 0.606 | 0.608 | 0.607 | 66.0984 | 0.18203
2020-02-06 | 11:28 | 3.411 | 3.422 | 3.4165 | 364.2987 | 0.609 | 0.611 | 0.61 | 66.425 | 0.18234
2020-02-06 | 11:28 | 3.425 | 3.442 | 3.4335 | 366.1114 | 0.613 | 0.616 | 0.6145 | 66.9151 | 0.18277
2020-02-06 | 11:28 | 3.448 | 3.472 | 3.46 | 368.9371 | 0.617 | 0.623 | 0.62 | 67.514 | 0.183
2020-02-06 | 11:28 | 3.48 | 3.506 | 3.493 | 372.4559 | 0.623 | 0.629 | 0.626 | 68.1673 | 0.18302
2020-02-06 | 11:28 | 3.512 | 3.533 | 3.5225 | 375.6014 | 0.63 | 0.633 | 0.6315 | 68.7662 | 0.18308
2020-02-06 | 11:28 | 3.534 | 3.54 | 3.537 | 377.1476 | 0.633 | 0.635 | 0.634 | 69.0385 | 0.18305
2020-02-06 | 11:28 | 3.529 | 3.539 | 3.534 | 376.8277 | 0.632 | 0.634 | 0.633 | 68.9296 | 0.18292
2020-02-06 | 11:29 | 3.498 | 3.523 | 3.5105 | 374.3219 | 0.625 | 0.631 | 0.628 | 68.3851 | 0.18269
2020-02-06 | 11:29 | 3.469 | 3.491 | 3.48 | 371.0697 | 0.619 | 0.625 | 0.622 | 67.7318 | 0.18253
2020-02-06 | 11:29 | 3.449 | 3.463 | 3.456 | 368.5106 | 0.617 | 0.619 | 0.618 | 67.2962 | 0.18262
2020-02-06 | 11:29 | 3.435 | 3.446 | 3.4405 | 366.8578 | 0.614 | 0.616 | 0.615 | 66.9695 | 0.18255
2020-02-06 | 11:29 | 3.427 | 3.435 | 3.431 | 365.8449 | 0.611 | 0.614 | 0.6125 | 66.6973 | 0.18231
2020-02-06 | 11:29 | 3.415 | 3.424 | 3.4195 | 364.6186 | 0.609 | 0.612 | 0.6105 | 66.4795 | 0.18233
2020-02-06 | 11:29 | 3.402 | 3.413 | 3.4075 | 363.3391 | 0.606 | 0.609 | 0.6075 | 66.1528 | 0.18207
2020-02-06 | 11:29 | 3.39 | 3.4 | 3.395 | 362.0062 | 0.604 | 0.607 | 0.6055 | 65.935 | 0.18214
2020-02-06 | 11:29 | 3.378 | 3.387 | 3.3825 | 360.6733 | 0.601 | 0.605 | 0.603 | 65.6628 | 0.18206
2020-02-06 | 11:29 | 3.367 | 3.376 | 3.3715 | 359.5004 | 0.599 | 0.6 | 0.5995 | 65.2817 | 0.18159
2020-02-06 | 11:29 | 3.355 | 3.363 | 3.359 | 358.1676 | 0.598 | 0.6 | 0.599 | 65.2272 | 0.18211
2020-02-06 | 11:30 | 3.344 | 3.352 | 3.348 | 356.9946 | 0.595 | 0.597 | 0.596 | 64.9005 | 0.1818
2020-02-06 | 11:30 | 3.331 | 3.341 | 3.336 | 355.7151 | 0.591 | 0.595 | 0.593 | 64.5738 | 0.18153
2020-02-06 | 11:30 | 3.32 | 3.329 | 3.3245 | 354.4889 | 0.59 | 0.592 | 0.591 | 64.3561 | 0.18155
2020-02-06 | 11:30 | 3.311 | 3.318 | 3.3145 | 353.4226 | 0.589 | 0.591 | 0.59 | 64.2472 | 0.18179
2020-02-06 | 11:30 | 3.299 | 3.308 | 3.3035 | 352.2496 | 0.587 | 0.589 | 0.588 | 64.0294 | 0.18177
2020-02-06 | 11:30 | 3.289 | 3.297 | 3.293 | 351.13 | 0.585 | 0.587 | 0.586 | 63.8116 | 0.18173
2020-02-06 | 11:30 | 3.279 | 3.288 | 3.2835 | 350.1171 | 0.583 | 0.585 | 0.584 | 63.5938 | 0.18164
2020-02-06 | 11:30 | 3.269 | 3.276 | 3.2725 | 348.9441 | 0.581 | 0.583 | 0.582 | 63.376 | 0.18162
2020-02-06 | 11:30 | 3.26 | 3.267 | 3.2635 | 347.9845 | 0.579 | 0.581 | 0.58 | 63.1582 | 0.1815
2020-02-06 | 11:30 | 3.252 | 3.258 | 3.255 | 347.0781 | 0.578 | 0.58 | 0.579 | 63.0493 | 0.18166
2020-02-06 | 11:30 | 3.249 | 3.251 | 3.25 | 346.545 | 0.577 | 0.579 | 0.578 | 62.9404 | 0.18162
2020-02-06 | 11:30 | 3.244 | 3.248 | 3.246 | 346.1185 | 0.577 | 0.578 | 0.5775 | 62.886 | 0.18169
2020-02-06 | 11:30 | 3.241 | 3.243 | 3.242 | 345.6919 | 0.576 | 0.577 | 0.5765 | 62.7771 | 0.1816
2020-02-06 | 11:31 | 3.237 | 3.24 | 3.2385 | 345.3187 | 0.576 | 0.577 | 0.5765 | 62.7771 | 0.18179
2020-02-06 | 11:31 | 3.234 | 3.237 | 3.2355 | 344.9989 | 0.575 | 0.576 | 0.5755 | 62.6682 | 0.18165
2020-02-06 | 11:31 | 3.235 | 3.236 | 3.2355 | 344.9989 | 0.575 | 0.576 | 0.5755 | 62.6682 | 0.18165
2020-02-06 | 11:31 | 3.235 | 3.238 | 3.2365 | 345.1055 | 0.575 | 0.577 | 0.576 | 62.7227 | 0.18175
2020-02-06 | 11:31 | 3.238 | 3.24 | 3.239 | 345.3721 | 0.576 | 0.578 | 0.577 | 62.8316 | 0.18192
2020-02-06 | 11:31 | 3.242 | 3.246 | 3.244 | 345.9052 | 0.577 | 0.578 | 0.5775 | 62.886 | 0.1818
2020-02-06 | 11:31 | 3.245 | 3.251 | 3.248 | 346.3317 | 0.578 | 0.58 | 0.579 | 63.0493 | 0.18205
2020-02-06 | 11:31 | 3.251 | 3.255 | 3.253 | 346.8649 | 0.579 | 0.581 | 0.58 | 63.1582 | 0.18208
2020-02-06 | 11:31 | 3.257 | 3.261 | 3.259 | 347.5046 | 0.581 | 0.581 | 0.581 | 63.2671 | 0.18206
2020-02-06 | 11:31 | 3.264 | 3.271 | 3.2675 | 348.411 | 0.581 | 0.583 | 0.582 | 63.376 | 0.1819
2020-02-06 | 11:31 | 3.274 | 3.284 | 3.279 | 349.6372 | 0.584 | 0.585 | 0.5845 | 63.6483 | 0.18204
2020-02-06 | 11:32 | 3.286 | 3.3 | 3.293 | 351.13 | 0.585 | 0.589 | 0.587 | 63.9205 | 0.18204
2020-02-06 | 11:32 | 3.304 | 3.32 | 3.312 | 353.156 | 0.589 | 0.592 | 0.5905 | 64.3016 | 0.18208
2020-02-06 | 11:32 | 3.325 | 3.345 | 3.335 | 355.6085 | 0.594 | 0.598 | 0.596 | 64.9005 | 0.18251
2020-02-06 | 11:32 | 3.351 | 3.376 | 3.3635 | 358.6474 | 0.598 | 0.603 | 0.6005 | 65.3905 | 0.18233
2020-02-06 | 11:32 | 3.382 | 3.409 | 3.3955 | 362.0595 | 0.605 | 0.611 | 0.608 | 66.2072 | 0.18286
2020-02-06 | 11:32 | 3.416 | 3.453 | 3.4345 | 366.2181 | 0.611 | 0.619 | 0.615 | 66.9695 | 0.18287
2020-02-06 | 11:32 | 3.462 | 3.508 | 3.485 | 371.6028 | 0.622 | 0.631 | 0.6265 | 68.2218 | 0.18359
2020-02-06 | 11:32 | 3.522 | 3.593 | 3.5575 | 379.3335 | 0.634 | 0.648 | 0.641 | 69.8007 | 0.18401
2020-02-06 | 11:32 | 3.619 | 3.723 | 3.671 | 391.4359 | 0.654 | 0.672 | 0.663 | 72.1964 | 0.18444
2020-02-06 | 11:32 | 3.738 | 4.116 | 3.927 | 418.733 | 0.675 | 0.76 | 0.7175 | 78.1311 | 0.18659
2020-02-06 | 11:32 | 4.013 | 4.169 | 4.091 | 436.2202 | 0.73 | 0.765 | 0.7475 | 81.3979 | 0.1866
2020-02-06 | 11:32 | 4.025 | 4.18 | 4.1025 | 437.4464 | 0.734 | 0.771 | 0.7525 | 81.9424 | 0.18732
2020-02-06 | 11:32 | 4.321 | 5.351 | 4.836 | 515.6589 | 0.813 | 1.027 | 0.92 | 100.182 | 0.19428
2020-02-06 | 11:33 | 4.42 | 5.217 | 4.8185 | 513.7929 | 0.815 | 0.981 | 0.898 | 97.7864 | 0.19032
2020-02-06 | 11:33 | 4.537 | 5.135 | 4.836 | 515.6589 | 0.85 | 0.968 | 0.909 | 98.9842 | 0.19196
2020-02-06 | 11:33 | 5.18 | 5.372 | 5.276 | 562.5758 | 0.982 | 1.022 | 1.002 | 109.1113 | 0.19395
2020-02-06 | 11:33 | 5.001 | 5.301 | 5.151 | 549.2471 | 0.949 | 1.01 | 0.9795 | 106.6612 | 0.1942
2020-02-06 | 11:33 | 5.083 | 5.752 | 5.4175 | 577.6638 | 0.974 | 1.114 | 1.044 | 113.6848 | 0.1968
2020-02-06 | 11:33 | 5.138 | 5.609 | 5.3735 | 572.9721 | 0.972 | 1.068 | 1.02 | 111.0714 | 0.19385
2020-02-06 | 11:33 | 5.255 | 6.042 | 5.6485 | 602.2952 | 1.018 | 1.184 | 1.101 | 119.8917 | 0.19906
2020-02-06 | 11:33 | 4.719 | 5.908 | 5.3135 | 566.5744 | 0.859 | 1.131 | 0.995 | 108.349 | 0.19124
2020-02-06 | 11:33 | 4.349 | 4.522 | 4.4355 | 472.9539 | 0.794 | 0.821 | 0.8075 | 87.9315 | 0.18592
2020-02-06 | 11:33 | 4.352 | 4.404 | 4.378 | 466.8227 | 0.797 | 0.808 | 0.8025 | 87.387 | 0.1872
2020-02-06 | 11:33 | 4.428 | 4.646 | 4.537 | 483.7768 | 0.813 | 0.87 | 0.8415 | 91.6339 | 0.18941
2020-02-06 | 11:34 | 4.832 | 5.957 | 5.3945 | 575.2113 | 0.919 | 1.182 | 1.0505 | 114.3926 | 0.19887
2020-02-06 | 11:34 | 6.224 | 7.109 | 6.6665 | 710.8437 | 1.242 | 1.458 | 1.35 | 147.0062 | 0.20681
2020-02-06 | 11:34 | 7.57 | 8.422 | 7.996 | 852.6073 | 1.565 | 1.737 | 1.651 | 179.7832 | 0.21086
2020-02-06 | 11:34 | 8.449 | 8.472 | 8.4605 | 902.1365 | 1.74 | 1.745 | 1.7425 | 189.7469 | 0.21033
2020-02-06 | 11:34 | 8.445 | 8.469 | 8.457 | 901.7633 | 1.74 | 1.744 | 1.742 | 189.6925 | 0.21036
2020-02-06 | 11:34 | 8.444 | 8.446 | 8.445 | 900.4838 | 1.739 | 1.74 | 1.7395 | 189.4202 | 0.21035
2020-02-06 | 11:34 | 8.447 | 8.469 | 8.458 | 901.87 | 1.74 | 1.743 | 1.7415 | 189.638 | 0.21027
2020-02-06 | 11:34 | 8.472 | 8.515 | 8.4935 | 905.6553 | 1.743 | 1.749 | 1.746 | 190.1281 | 0.20993
2020-02-06 | 11:34 | 8.523 | 8.537 | 8.53 | 909.5473 | 1.751 | 1.752 | 1.7515 | 190.727 | 0.20969
2020-02-06 | 11:34 | 8.48 | 8.532 | 8.506 | 906.9882 | 1.728 | 1.75 | 1.739 | 189.3658 | 0.20879
2020-02-06 | 11:34 | 7.713 | 8.303 | 8.008 | 853.8868 | 1.556 | 1.678 | 1.617 | 176.0808 | 0.20621
2020-02-06 | 11:34 | 8.049 | 8.397 | 8.223 | 876.8121 | 1.651 | 1.724 | 1.6875 | 183.7578 | 0.20957
2020-02-06 | 11:34 | 8.426 | 8.456 | 8.441 | 900.0573 | 1.73 | 1.735 | 1.7325 | 188.658 | 0.20961
2020-02-06 | 11:35 | 8.364 | 8.446 | 8.405 | 896.2186 | 1.715 | 1.731 | 1.723 | 187.6235 | 0.20935
2020-02-06 | 11:35 | 8.352 | 8.434 | 8.393 | 894.9391 | 1.715 | 1.731 | 1.723 | 187.6235 | 0.20965
2020-02-06 | 11:35 | 8.444 | 8.453 | 8.4485 | 900.857 | 1.732 | 1.732 | 1.732 | 188.6035 | 0.20936
2020-02-06 | 11:35 | 8.455 | 8.471 | 8.463 | 902.4031 | 1.731 | 1.735 | 1.733 | 188.7124 | 0.20912
2020-02-06 | 11:35 | 8.477 | 8.483 | 8.48 | 904.2158 | 1.734 | 1.736 | 1.735 | 188.9302 | 0.20894
2020-02-06 | 11:35 | 8.485 | 8.506 | 8.4955 | 905.8686 | 1.737 | 1.741 | 1.739 | 189.3658 | 0.20904
2020-02-06 | 11:35 | 8.514 | 8.519 | 8.5165 | 908.1078 | 1.74 | 1.741 | 1.7405 | 189.5291 | 0.20871
2020-02-06 | 11:35 | 8.518 | 8.528 | 8.523 | 908.8009 | 1.742 | 1.743 | 1.7425 | 189.7469 | 0.20879
2020-02-06 | 11:35 | 8.529 | 8.54 | 8.5345 | 910.0271 | 1.742 | 1.745 | 1.7435 | 189.8558 | 0.20863
2020-02-06 | 11:35 | 8.537 | 8.543 | 8.54 | 910.6136 | 1.745 | 1.746 | 1.7455 | 190.0736 | 0.20873
2020-02-06 | 11:35 | 8.541 | 8.56 | 8.5505 | 911.7332 | 1.746 | 1.749 | 1.7475 | 190.2914 | 0.20871
2020-02-06 | 11:36 | 8.565 | 8.592 | 8.5785 | 914.7188 | 1.75 | 1.756 | 1.753 | 190.8903 | 0.20869
2020-02-06 | 11:36 | 8.591 | 8.598 | 8.5945 | 916.4249 | 1.754 | 1.755 | 1.7545 | 191.0536 | 0.20848
2020-02-06 | 11:36 | 8.557 | 8.587 | 8.572 | 914.0257 | 1.748 | 1.752 | 1.75 | 190.5636 | 0.20849
2020-02-06 | 11:36 | 8.498 | 8.544 | 8.521 | 908.5876 | 1.738 | 1.745 | 1.7415 | 189.638 | 0.20872
2020-02-06 | 11:36 | 8.44 | 8.491 | 8.4655 | 902.6697 | 1.727 | 1.738 | 1.7325 | 188.658 | 0.209
2020-02-06 | 11:36 | 8.405 | 8.427 | 8.416 | 897.3915 | 1.722 | 1.724 | 1.723 | 187.6235 | 0.20908
2020-02-06 | 11:36 | 8.408 | 8.444 | 8.426 | 898.4578 | 1.723 | 1.729 | 1.726 | 187.9502 | 0.20919
2020-02-06 | 11:36 | 8.435 | 8.446 | 8.4405 | 900.004 | 1.727 | 1.73 | 1.7285 | 188.2224 | 0.20914
2020-02-06 | 11:36 | 8.45 | 8.461 | 8.4555 | 901.6034 | 1.731 | 1.733 | 1.732 | 188.6035 | 0.20919
2020-02-06 | 11:36 | 8.472 | 8.489 | 8.4805 | 904.2691 | 1.734 | 1.737 | 1.7355 | 188.9847 | 0.20899
2020-02-06 | 11:36 | 8.49 | 8.499 | 8.4945 | 905.7619 | 1.737 | 1.74 | 1.7385 | 189.3114 | 0.20901
2020-02-06 | 11:36 | 8.483 | 8.499 | 8.491 | 905.3887 | 1.737 | 1.741 | 1.739 | 189.3658 | 0.20915
2020-02-06 | 11:36 | 8.447 | 8.469 | 8.458 | 901.87 | 1.731 | 1.734 | 1.7325 | 188.658 | 0.20919
2020-02-06 | 11:37 | 8.458 | 8.491 | 8.4745 | 903.6294 | 1.734 | 1.74 | 1.737 | 189.148 | 0.20932
2020-02-06 | 11:37 | 8.5 | 8.554 | 8.527 | 909.2274 | 1.742 | 1.748 | 1.745 | 190.0192 | 0.20899
2020-02-06 | 11:37 | 8.556 | 8.558 | 8.557 | 912.4263 | 1.748 | 1.749 | 1.7485 | 190.4003 | 0.20867
2020-02-06 | 11:37 | 8.555 | 8.568 | 8.5615 | 912.9061 | 1.748 | 1.752 | 1.75 | 190.5636 | 0.20874
2020-02-06 | 11:37 | 8.571 | 8.578 | 8.5745 | 914.2923 | 1.751 | 1.753 | 1.752 | 190.7814 | 0.20867
2020-02-06 | 11:37 | 8.582 | 8.588 | 8.585 | 915.4119 | 1.753 | 1.756 | 1.7545 | 191.0536 | 0.20871
2020-02-06 | 11:37 | 8.589 | 8.605 | 8.597 | 916.6914 | 1.756 | 1.759 | 1.7575 | 191.3803 | 0.20877
2020-02-06 | 11:37 | 8.612 | 8.634 | 8.623 | 919.4638 | 1.76 | 1.766 | 1.763 | 191.9792 | 0.20879
2020-02-06 | 11:37 | 8.636 | 8.638 | 8.637 | 920.9566 | 1.765 | 1.766 | 1.7655 | 192.2515 | 0.20875
2020-02-06 | 11:37 | 8.636 | 8.64 | 8.638 | 921.0632 | 1.764 | 1.768 | 1.766 | 192.3059 | 0.20879
2020-02-06 | 11:37 | 8.642 | 8.649 | 8.6455 | 921.8629 | 1.768 | 1.77 | 1.769 | 192.6326 | 0.20896
2020-02-06 | 11:38 | 8.635 | 8.647 | 8.641 | 921.3831 | 1.768 | 1.769 | 1.7685 | 192.5782 | 0.20901
2020-02-06 | 11:38 | 8.615 | 8.632 | 8.6235 | 919.5171 | 1.764 | 1.766 | 1.765 | 192.197 | 0.20902
2020-02-06 | 11:38 | 8.615 | 8.632 | 8.6235 | 919.5171 | 1.766 | 1.769 | 1.7675 | 192.4693 | 0.20932
2020-02-06 | 11:38 | 8.635 | 8.655 | 8.645 | 921.8096 | 1.768 | 1.774 | 1.771 | 192.8504 | 0.20921
2020-02-06 | 11:38 | 8.661 | 8.676 | 8.6685 | 924.3154 | 1.774 | 1.778 | 1.776 | 193.3949 | 0.20923
2020-02-06 | 11:38 | 8.678 | 8.687 | 8.6825 | 925.8082 | 1.779 | 1.78 | 1.7795 | 193.776 | 0.2093
2020-02-06 | 11:38 | 8.687 | 8.689 | 8.688 | 926.3947 | 1.779 | 1.781 | 1.78 | 193.8304 | 0.20923
2020-02-06 | 11:38 | 8.683 | 8.689 | 8.686 | 926.1814 | 1.779 | 1.781 | 1.78 | 193.8304 | 0.20928
2020-02-06 | 11:38 | 8.664 | 8.68 | 8.672 | 924.6886 | 1.777 | 1.779 | 1.778 | 193.6126 | 0.20938
2020-02-06 | 11:38 | 8.64 | 8.659 | 8.6495 | 922.2895 | 1.774 | 1.776 | 1.775 | 193.286 | 0.20957
2020-02-06 | 11:38 | 8.632 | 8.639 | 8.6355 | 920.7967 | 1.773 | 1.775 | 1.774 | 193.1771 | 0.20979
2020-02-06 | 11:38 | 8.626 | 8.641 | 8.6335 | 920.5834 | 1.773 | 1.775 | 1.774 | 193.1771 | 0.20984
2020-02-06 | 11:39 | 8.613 | 8.623 | 8.618 | 918.9306 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.20998
2020-02-06 | 11:39 | 8.615 | 8.629 | 8.622 | 919.3572 | 1.772 | 1.776 | 1.774 | 193.1771 | 0.21012
2020-02-06 | 11:39 | 8.61 | 8.635 | 8.6225 | 919.4105 | 1.768 | 1.776 | 1.772 | 192.9593 | 0.20987
2020-02-06 | 11:39 | 8.499 | 8.593 | 8.546 | 911.2533 | 1.748 | 1.765 | 1.7565 | 191.2714 | 0.2099
2020-02-06 | 11:39 | 8.553 | 8.574 | 8.5635 | 913.1194 | 1.758 | 1.764 | 1.761 | 191.7615 | 0.21001
2020-02-06 | 11:39 | 8.564 | 8.57 | 8.567 | 913.4926 | 1.763 | 1.767 | 1.765 | 192.197 | 0.2104
2020-02-06 | 11:39 | 8.552 | 8.563 | 8.5575 | 912.4796 | 1.764 | 1.765 | 1.7645 | 192.1426 | 0.21057
2020-02-06 | 11:39 | 8.543 | 8.552 | 8.5475 | 911.4133 | 1.762 | 1.765 | 1.7635 | 192.0337 | 0.2107
2020-02-06 | 11:39 | 8.535 | 8.546 | 8.5405 | 910.6669 | 1.761 | 1.762 | 1.7615 | 191.8159 | 0.21063
2020-02-06 | 11:39 | 8.509 | 8.53 | 8.5195 | 908.4277 | 1.758 | 1.759 | 1.7585 | 191.4892 | 0.21079
2020-02-06 | 11:39 | 8.357 | 8.494 | 8.4255 | 898.4045 | 1.729 | 1.756 | 1.7425 | 189.7469 | 0.2112
2020-02-06 | 11:39 | 8.325 | 8.348 | 8.3365 | 888.9145 | 1.724 | 1.732 | 1.728 | 188.168 | 0.21168
2020-02-06 | 11:40 | 8.361 | 8.387 | 8.374 | 892.9131 | 1.73 | 1.735 | 1.7325 | 188.658 | 0.21128
2020-02-06 | 11:40 | 8.346 | 8.376 | 8.361 | 891.5269 | 1.727 | 1.732 | 1.7295 | 188.3313 | 0.21125
2020-02-06 | 11:40 | 8.329 | 8.369 | 8.349 | 890.2474 | 1.717 | 1.727 | 1.722 | 187.5146 | 0.21063
2020-02-06 | 11:40 | 8.275 | 8.321 | 8.298 | 884.8093 | 1.707 | 1.717 | 1.712 | 186.4257 | 0.2107
2020-02-06 | 11:40 | 8.207 | 8.263 | 8.235 | 878.0917 | 1.697 | 1.705 | 1.701 | 185.2278 | 0.21094
2020-02-06 | 11:40 | 8.118 | 8.191 | 8.1545 | 869.508 | 1.681 | 1.693 | 1.687 | 183.7033 | 0.21127
2020-02-06 | 11:40 | 8.068 | 8.1 | 8.084 | 861.9906 | 1.679 | 1.688 | 1.6835 | 183.3222 | 0.21267
2020-02-06 | 11:40 | 7.941 | 8.041 | 7.991 | 852.0741 | 1.65 | 1.67 | 1.66 | 180.7632 | 0.21214
2020-02-06 | 11:40 | 7.711 | 7.915 | 7.813 | 833.0941 | 1.598 | 1.644 | 1.621 | 176.5164 | 0.21188
2020-02-06 | 11:40 | 6.818 | 7.662 | 7.24 | 771.9956 | 1.392 | 1.589 | 1.4905 | 162.3058 | 0.21024
2020-02-06 | 11:40 | 4.885 | 6.593 | 5.739 | 611.9451 | 0.921 | 1.332 | 1.1265 | 122.6685 | 0.20046
2020-02-06 | 11:40 | 4.05 | 4.525 | 4.2875 | 457.1728 | 0.755 | 0.846 | 0.8005 | 87.1692 | 0.19067
2020-02-06 | 11:41 | 3.824 | 4.003 | 3.9135 | 417.2935 | 0.712 | 0.747 | 0.7295 | 79.4378 | 0.19036
2020-02-06 | 11:41 | 3.703 | 3.793 | 3.748 | 399.6463 | 0.688 | 0.704 | 0.696 | 75.7899 | 0.18964
2020-02-06 | 11:41 | 3.633 | 3.689 | 3.661 | 390.3696 | 0.677 | 0.686 | 0.6815 | 74.2109 | 0.1901
2020-02-06 | 11:41 | 3.599 | 3.624 | 3.6115 | 385.0914 | 0.67 | 0.674 | 0.672 | 73.1764 | 0.19002
2020-02-06 | 11:41 | 3.573 | 3.592 | 3.5825 | 381.9992 | 0.665 | 0.669 | 0.667 | 72.632 | 0.19014
2020-02-06 | 11:41 | 3.562 | 3.571 | 3.5665 | 380.2931 | 0.663 | 0.664 | 0.6635 | 72.2508 | 0.18999
2020-02-06 | 11:41 | 3.552 | 3.56 | 3.556 | 379.1735 | 0.661 | 0.663 | 0.662 | 72.0875 | 0.19012
2020-02-06 | 11:41 | 3.543 | 3.55 | 3.5465 | 378.1605 | 0.659 | 0.662 | 0.6605 | 71.9242 | 0.19019
2020-02-06 | 11:41 | 3.537 | 3.541 | 3.539 | 377.3608 | 0.658 | 0.66 | 0.659 | 71.7608 | 0.19016
2020-02-06 | 11:41 | 3.534 | 3.536 | 3.535 | 376.9343 | 0.658 | 0.659 | 0.6585 | 71.7064 | 0.19024
2020-02-06 | 11:41 | 3.535 | 3.565 | 3.55 | 378.5337 | 0.659 | 0.671 | 0.665 | 72.4142 | 0.1913
2020-02-06 | 11:41 | 3.646 | 3.861 | 3.7535 | 400.2328 | 0.692 | 0.741 | 0.7165 | 78.0222 | 0.19494
2020-02-06 | 11:42 | 3.916 | 4.246 | 4.081 | 435.1539 | 0.756 | 0.823 | 0.7895 | 85.9714 | 0.19757
2020-02-06 | 11:42 | 4.025 | 4.258 | 4.1415 | 441.6049 | 0.761 | 0.82 | 0.7905 | 86.0803 | 0.19493
2020-02-06 | 11:42 | 3.644 | 3.928 | 3.786 | 403.6982 | 0.68 | 0.739 | 0.7095 | 77.2599 | 0.19138
2020-02-06 | 11:42 | 3.47 | 3.589 | 3.5295 | 376.3478 | 0.646 | 0.67 | 0.658 | 71.6519 | 0.19039
2020-02-06 | 11:42 | 3.454 | 3.464 | 3.459 | 368.8305 | 0.645 | 0.647 | 0.646 | 70.3452 | 0.19073
2020-02-06 | 11:42 | 3.44 | 3.465 | 3.4525 | 368.1374 | 0.643 | 0.648 | 0.6455 | 70.2908 | 0.19094
2020-02-06 | 11:42 | 3.393 | 3.432 | 3.4125 | 363.8722 | 0.633 | 0.64 | 0.6365 | 69.3107 | 0.19048
2020-02-06 | 11:42 | 3.332 | 3.38 | 3.356 | 357.8477 | 0.621 | 0.632 | 0.6265 | 68.2218 | 0.19064
2020-02-06 | 11:42 | 3.288 | 3.321 | 3.3045 | 352.3563 | 0.613 | 0.618 | 0.6155 | 67.024 | 0.19022
2020-02-06 | 11:42 | 3.259 | 3.28 | 3.2695 | 348.6242 | 0.607 | 0.613 | 0.61 | 66.425 | 0.19053
2020-02-06 | 11:42 | 3.238 | 3.255 | 3.2465 | 346.1718 | 0.603 | 0.606 | 0.6045 | 65.8261 | 0.19015
2020-02-06 | 11:42 | 3.219 | 3.233 | 3.226 | 343.9859 | 0.6 | 0.602 | 0.601 | 65.445 | 0.19025
2020-02-06 | 11:43 | 3.206 | 3.217 | 3.2115 | 342.4397 | 0.597 | 0.6 | 0.5985 | 65.1728 | 0.19032
2020-02-06 | 11:43 | 3.192 | 3.202 | 3.197 | 340.8936 | 0.595 | 0.597 | 0.596 | 64.9005 | 0.19038
2020-02-06 | 11:43 | 3.181 | 3.19 | 3.1855 | 339.6674 | 0.593 | 0.594 | 0.5935 | 64.6283 | 0.19027
2020-02-06 | 11:43 | 3.173 | 3.179 | 3.176 | 338.6544 | 0.589 | 0.592 | 0.5905 | 64.3016 | 0.18987
2020-02-06 | 11:43 | 3.17 | 3.171 | 3.1705 | 338.068 | 0.589 | 0.591 | 0.59 | 64.2472 | 0.19004
2020-02-06 | 11:43 | 3.168 | 3.17 | 3.169 | 337.908 | 0.589 | 0.591 | 0.59 | 64.2472 | 0.19013
2020-02-06 | 11:43 | 3.17 | 3.171 | 3.1705 | 338.068 | 0.589 | 0.591 | 0.59 | 64.2472 | 0.19004
2020-02-06 | 11:43 | 3.172 | 3.176 | 3.174 | 338.4412 | 0.589 | 0.592 | 0.5905 | 64.3016 | 0.18999
2020-02-06 | 11:43 | 3.176 | 3.184 | 3.18 | 339.0809 | 0.591 | 0.593 | 0.592 | 64.465 | 0.19012
2020-02-06 | 11:43 | 3.185 | 3.191 | 3.188 | 339.934 | 0.593 | 0.596 | 0.5945 | 64.7372 | 0.19044
2020-02-06 | 11:43 | 3.192 | 3.202 | 3.197 | 340.8936 | 0.595 | 0.597 | 0.596 | 64.9005 | 0.19038
2020-02-06 | 11:43 | 3.204 | 3.212 | 3.208 | 342.0665 | 0.597 | 0.6 | 0.5985 | 65.1728 | 0.19053
2020-02-06 | 11:44 | 3.216 | 3.224 | 3.22 | 343.3461 | 0.6 | 0.602 | 0.601 | 65.445 | 0.19061
2020-02-06 | 11:44 | 3.226 | 3.234 | 3.23 | 344.4124 | 0.602 | 0.605 | 0.6035 | 65.7172 | 0.19081
2020-02-06 | 11:44 | 3.235 | 3.243 | 3.239 | 345.3721 | 0.603 | 0.605 | 0.604 | 65.7717 | 0.19044
2020-02-06 | 11:44 | 3.244 | 3.254 | 3.249 | 346.4383 | 0.604 | 0.607 | 0.6055 | 65.935 | 0.19032
2020-02-06 | 11:44 | 3.256 | 3.264 | 3.26 | 347.6113 | 0.608 | 0.609 | 0.6085 | 66.2617 | 0.19062
2020-02-06 | 11:44 | 3.266 | 3.274 | 3.27 | 348.6776 | 0.61 | 0.612 | 0.611 | 66.5339 | 0.19082
2020-02-06 | 11:44 | 3.276 | 3.285 | 3.2805 | 349.7972 | 0.611 | 0.612 | 0.6115 | 66.5884 | 0.19036
2020-02-06 | 11:44 | 3.287 | 3.296 | 3.2915 | 350.9701 | 0.613 | 0.614 | 0.6135 | 66.8062 | 0.19035
2020-02-06 | 11:44 | 3.299 | 3.307 | 3.303 | 352.1963 | 0.614 | 0.617 | 0.6155 | 67.024 | 0.1903
2020-02-06 | 11:44 | 3.309 | 3.317 | 3.313 | 353.2626 | 0.617 | 0.618 | 0.6175 | 67.2417 | 0.19034
2020-02-06 | 11:44 | 3.321 | 3.329 | 3.325 | 354.5422 | 0.618 | 0.62 | 0.619 | 67.4051 | 0.19012
2020-02-06 | 11:44 | 3.331 | 3.342 | 3.3365 | 355.7684 | 0.62 | 0.622 | 0.621 | 67.6229 | 0.19008
2020-02-06 | 11:44 | 3.344 | 3.353 | 3.3485 | 357.048 | 0.622 | 0.623 | 0.6225 | 67.7862 | 0.18985
2020-02-06 | 11:45 | 3.358 | 3.37 | 3.364 | 358.7007 | 0.625 | 0.627 | 0.626 | 68.1673 | 0.19004
2020-02-06 | 11:45 | 3.374 | 3.388 | 3.381 | 360.5134 | 0.627 | 0.631 | 0.629 | 68.494 | 0.18999
2020-02-06 | 11:45 | 3.393 | 3.408 | 3.4005 | 362.5927 | 0.63 | 0.634 | 0.632 | 68.8207 | 0.1898
2020-02-06 | 11:45 | 3.414 | 3.43 | 3.422 | 364.8852 | 0.635 | 0.639 | 0.637 | 69.3652 | 0.1901
2020-02-06 | 11:45 | 3.434 | 3.451 | 3.4425 | 367.0711 | 0.64 | 0.642 | 0.641 | 69.8007 | 0.19016
2020-02-06 | 11:45 | 3.455 | 3.473 | 3.464 | 369.3636 | 0.644 | 0.647 | 0.6455 | 70.2908 | 0.1903
2020-02-06 | 11:45 | 3.479 | 3.495 | 3.487 | 371.8161 | 0.647 | 0.651 | 0.649 | 70.6719 | 0.19007
2020-02-06 | 11:45 | 3.498 | 3.514 | 3.506 | 373.8421 | 0.652 | 0.653 | 0.6525 | 71.053 | 0.19006
2020-02-06 | 11:45 | 3.519 | 3.535 | 3.527 | 376.0813 | 0.654 | 0.657 | 0.6555 | 71.3797 | 0.1898
2020-02-06 | 11:45 | 3.538 | 3.551 | 3.5445 | 377.9473 | 0.657 | 0.66 | 0.6585 | 71.7064 | 0.18973
2020-02-06 | 11:45 | 3.555 | 3.571 | 3.563 | 379.9199 | 0.662 | 0.663 | 0.6625 | 72.1419 | 0.18989
2020-02-06 | 11:46 | 3.576 | 3.59 | 3.583 | 382.0525 | 0.664 | 0.666 | 0.665 | 72.4142 | 0.18954
2020-02-06 | 11:46 | 3.595 | 3.612 | 3.6035 | 384.2384 | 0.668 | 0.671 | 0.6695 | 72.9042 | 0.18974
2020-02-06 | 11:46 | 3.617 | 3.635 | 3.626 | 386.6376 | 0.671 | 0.674 | 0.6725 | 73.2309 | 0.1894
2020-02-06 | 11:46 | 3.641 | 3.661 | 3.651 | 389.3033 | 0.674 | 0.678 | 0.676 | 73.612 | 0.18909
2020-02-06 | 11:46 | 3.665 | 3.684 | 3.6745 | 391.8091 | 0.679 | 0.682 | 0.6805 | 74.102 | 0.18913
2020-02-06 | 11:46 | 3.689 | 3.706 | 3.6975 | 394.2616 | 0.682 | 0.687 | 0.6845 | 74.5376 | 0.18906
2020-02-06 | 11:46 | 3.713 | 3.731 | 3.722 | 396.874 | 0.686 | 0.691 | 0.6885 | 74.9732 | 0.18891
2020-02-06 | 11:46 | 3.734 | 3.753 | 3.7435 | 399.1665 | 0.691 | 0.695 | 0.693 | 75.4632 | 0.18905
2020-02-06 | 11:46 | 3.757 | 3.776 | 3.7665 | 401.619 | 0.695 | 0.698 | 0.6965 | 75.8443 | 0.18885
2020-02-06 | 11:46 | 3.78 | 3.798 | 3.789 | 404.0181 | 0.699 | 0.701 | 0.7 | 76.2255 | 0.18867
2020-02-06 | 11:46 | 3.803 | 3.82 | 3.8115 | 406.4173 | 0.702 | 0.705 | 0.7035 | 76.6066 | 0.18849
2020-02-06 | 11:46 | 3.825 | 3.84 | 3.8325 | 408.6565 | 0.704 | 0.707 | 0.7055 | 76.8244 | 0.18799
2020-02-06 | 11:46 | 3.843 | 3.859 | 3.851 | 410.6291 | 0.707 | 0.709 | 0.708 | 77.0966 | 0.18775
2020-02-06 | 11:47 | 3.864 | 3.881 | 3.8725 | 412.9217 | 0.71 | 0.713 | 0.7115 | 77.4777 | 0.18763
2020-02-06 | 11:47 | 3.886 | 3.916 | 3.901 | 415.9606 | 0.714 | 0.719 | 0.7165 | 78.0222 | 0.18757
2020-02-06 | 11:47 | 3.926 | 3.952 | 3.939 | 420.0125 | 0.719 | 0.725 | 0.722 | 78.6211 | 0.18719
2020-02-06 | 11:47 | 3.958 | 4.003 | 3.9805 | 424.4376 | 0.726 | 0.734 | 0.73 | 79.4923 | 0.18729
2020-02-06 | 11:47 | 4.015 | 4.088 | 4.0515 | 432.0083 | 0.736 | 0.75 | 0.743 | 80.9079 | 0.18728
2020-02-06 | 11:47 | 4.111 | 4.232 | 4.1715 | 444.8038 | 0.756 | 0.782 | 0.769 | 83.7391 | 0.18826
2020-02-06 | 11:47 | 4.261 | 4.338 | 4.2995 | 458.4523 | 0.787 | 0.804 | 0.7955 | 86.6248 | 0.18895
2020-02-06 | 11:47 | 4.36 | 4.528 | 4.444 | 473.8603 | 0.808 | 0.845 | 0.8265 | 90.0005 | 0.18993
2020-02-06 | 11:47 | 4.584 | 4.824 | 4.704 | 501.5839 | 0.858 | 0.909 | 0.8835 | 96.2074 | 0.19181
2020-02-06 | 11:47 | 4.886 | 5.192 | 5.039 | 537.3047 | 0.924 | 0.99 | 0.957 | 104.2111 | 0.19395
2020-02-06 | 11:47 | 5.255 | 5.335 | 5.295 | 564.6017 | 1.003 | 1.019 | 1.011 | 110.0913 | 0.19499
2020-02-06 | 11:48 | 5.315 | 5.405 | 5.36 | 571.5326 | 1.017 | 1.05 | 1.0335 | 112.5414 | 0.19691
2020-02-06 | 11:48 | 5.464 | 5.875 | 5.6695 | 604.5344 | 1.065 | 1.16 | 1.1125 | 121.144 | 0.20039
2020-02-06 | 11:48 | 6.023 | 6.563 | 6.293 | 671.0177 | 1.193 | 1.307 | 1.25 | 136.1169 | 0.20285
2020-02-06 | 11:48 | 6.692 | 7.253 | 6.9725 | 743.4723 | 1.337 | 1.458 | 1.3975 | 152.1787 | 0.20469
2020-02-06 | 11:48 | 5.219 | 7.342 | 6.2805 | 669.6848 | 0.931 | 1.46 | 1.1955 | 130.1822 | 0.19439
2020-02-06 | 11:48 | 4.516 | 4.788 | 4.652 | 496.0391 | 0.816 | 0.853 | 0.8345 | 90.8716 | 0.18319
2020-02-06 | 11:48 | 4.733 | 5.111 | 4.922 | 524.829 | 0.887 | 0.96 | 0.9235 | 100.5631 | 0.19161
2020-02-06 | 11:48 | 5.072 | 6.363 | 5.7175 | 609.6526 | 0.96 | 1.227 | 1.0935 | 119.075 | 0.19532
2020-02-06 | 11:48 | 6.432 | 7.486 | 6.959 | 742.0328 | 1.25 | 1.513 | 1.3815 | 150.4364 | 0.20274
2020-02-06 | 11:48 | 7.843 | 8.595 | 8.219 | 876.3856 | 1.6 | 1.749 | 1.6745 | 182.3422 | 0.20806
2020-02-06 | 11:48 | 8.565 | 8.72 | 8.6425 | 921.5431 | 1.716 | 1.77 | 1.743 | 189.8014 | 0.20596
2020-02-06 | 11:48 | 8.388 | 8.464 | 8.426 | 898.4578 | 1.677 | 1.694 | 1.6855 | 183.54 | 0.20428
2020-02-06 | 11:48 | 8.417 | 8.511 | 8.464 | 902.5097 | 1.687 | 1.712 | 1.6995 | 185.0645 | 0.20506
2020-02-06 | 11:49 | 8.612 | 8.807 | 8.7095 | 928.6872 | 1.74 | 1.783 | 1.7615 | 191.8159 | 0.20655
2020-02-06 | 11:49 | 8.817 | 8.83 | 8.8235 | 940.8429 | 1.785 | 1.787 | 1.786 | 194.4838 | 0.20671
2020-02-06 | 11:49 | 8.836 | 8.858 | 8.847 | 943.3487 | 1.788 | 1.791 | 1.7895 | 194.8649 | 0.20657
2020-02-06 | 11:49 | 8.862 | 8.865 | 8.8635 | 945.1081 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20647
2020-02-06 | 11:49 | 8.866 | 8.875 | 8.8705 | 945.8545 | 1.792 | 1.793 | 1.7925 | 195.1916 | 0.20637
2020-02-06 | 11:49 | 8.878 | 8.886 | 8.882 | 947.0808 | 1.793 | 1.795 | 1.794 | 195.3549 | 0.20627
2020-02-06 | 11:49 | 8.701 | 8.887 | 8.794 | 937.6974 | 1.75 | 1.794 | 1.772 | 192.9593 | 0.20578
2020-02-06 | 11:49 | 7.717 | 8.667 | 8.192 | 873.5066 | 1.52 | 1.734 | 1.627 | 177.1697 | 0.20283
2020-02-06 | 11:49 | 7.698 | 8.207 | 7.9525 | 847.9689 | 1.527 | 1.644 | 1.5855 | 172.6506 | 0.2036
2020-02-06 | 11:49 | 8.192 | 8.699 | 8.4455 | 900.5371 | 1.644 | 1.76 | 1.702 | 185.3367 | 0.20581
2020-02-06 | 11:49 | 8.784 | 8.918 | 8.851 | 943.7753 | 1.776 | 1.797 | 1.7865 | 194.5382 | 0.20613
2020-02-06 | 11:50 | 8.924 | 8.927 | 8.9255 | 951.7191 | 1.791 | 1.796 | 1.7935 | 195.3005 | 0.20521
2020-02-06 | 11:50 | 8.919 | 8.922 | 8.9205 | 951.186 | 1.791 | 1.792 | 1.7915 | 195.0827 | 0.20509
2020-02-06 | 11:50 | 8.92 | 8.922 | 8.921 | 951.2393 | 1.788 | 1.79 | 1.789 | 194.8105 | 0.2048
2020-02-06 | 11:50 | 8.919 | 8.922 | 8.9205 | 951.186 | 1.788 | 1.79 | 1.789 | 194.8105 | 0.20481
2020-02-06 | 11:50 | 8.921 | 8.928 | 8.9245 | 951.6125 | 1.787 | 1.789 | 1.788 | 194.7016 | 0.2046
2020-02-06 | 11:50 | 8.93 | 8.932 | 8.931 | 952.3056 | 1.789 | 1.793 | 1.791 | 195.0283 | 0.2048
2020-02-06 | 11:50 | 8.927 | 8.929 | 8.928 | 951.9857 | 1.792 | 1.793 | 1.7925 | 195.1916 | 0.20504
2020-02-06 | 11:50 | 8.918 | 8.927 | 8.9225 | 951.3992 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20499
2020-02-06 | 11:50 | 8.906 | 8.918 | 8.912 | 950.2796 | 1.788 | 1.79 | 1.789 | 194.8105 | 0.205
2020-02-06 | 11:50 | 8.926 | 8.94 | 8.933 | 952.5188 | 1.792 | 1.794 | 1.793 | 195.246 | 0.20498
2020-02-06 | 11:50 | 8.943 | 8.955 | 8.949 | 954.2249 | 1.793 | 1.797 | 1.795 | 195.4638 | 0.20484
2020-02-06 | 11:50 | 8.953 | 8.957 | 8.955 | 954.8647 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20487
2020-02-06 | 11:50 | 8.954 | 8.955 | 8.9545 | 954.8114 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20489
2020-02-06 | 11:51 | 8.957 | 8.96 | 8.9585 | 955.2379 | 1.799 | 1.802 | 1.8005 | 196.0628 | 0.20525
2020-02-06 | 11:51 | 8.949 | 8.953 | 8.951 | 954.4382 | 1.802 | 1.802 | 1.802 | 196.2261 | 0.20559
2020-02-06 | 11:51 | 8.958 | 8.978 | 8.968 | 956.2509 | 1.803 | 1.807 | 1.805 | 196.5528 | 0.20555
2020-02-06 | 11:51 | 8.98 | 8.983 | 8.9815 | 957.6904 | 1.807 | 1.807 | 1.807 | 196.7706 | 0.20546
2020-02-06 | 11:51 | 8.978 | 8.984 | 8.981 | 957.6371 | 1.806 | 1.808 | 1.807 | 196.7706 | 0.20548
2020-02-06 | 11:51 | 8.977 | 8.979 | 8.978 | 957.3172 | 1.805 | 1.807 | 1.806 | 196.6617 | 0.20543
2020-02-06 | 11:51 | 8.977 | 8.983 | 8.98 | 957.5304 | 1.806 | 1.808 | 1.807 | 196.7706 | 0.2055
2020-02-06 | 11:51 | 8.983 | 8.986 | 8.9845 | 958.0103 | 1.807 | 1.81 | 1.8085 | 196.9339 | 0.20557
2020-02-06 | 11:51 | 8.98 | 8.984 | 8.982 | 957.7437 | 1.807 | 1.808 | 1.8075 | 196.825 | 0.20551
2020-02-06 | 11:51 | 8.982 | 8.985 | 8.9835 | 957.9036 | 1.809 | 1.81 | 1.8095 | 197.0428 | 0.2057
2020-02-06 | 11:51 | 8.987 | 8.99 | 8.9885 | 958.4368 | 1.809 | 1.81 | 1.8095 | 197.0428 | 0.20559
2020-02-06 | 11:52 | 8.991 | 8.994 | 8.9925 | 958.8633 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20544
2020-02-06 | 11:52 | 8.985 | 8.992 | 8.9885 | 958.4368 | 1.81 | 1.811 | 1.8105 | 197.1517 | 0.2057
2020-02-06 | 11:52 | 8.983 | 8.987 | 8.985 | 958.0636 | 1.81 | 1.811 | 1.8105 | 197.1517 | 0.20578
2020-02-06 | 11:52 | 8.978 | 8.984 | 8.981 | 957.6371 | 1.8 | 1.81 | 1.805 | 196.5528 | 0.20525
2020-02-06 | 11:52 | 8.977 | 8.978 | 8.9775 | 957.2638 | 1.798 | 1.801 | 1.7995 | 195.9539 | 0.2047
2020-02-06 | 11:52 | 8.985 | 8.996 | 8.9905 | 958.65 | 1.792 | 1.801 | 1.7965 | 195.6272 | 0.20407
2020-02-06 | 11:52 | 8.999 | 9.022 | 9.0105 | 960.7826 | 1.791 | 1.795 | 1.793 | 195.246 | 0.20322
2020-02-06 | 11:52 | 9.025 | 9.053 | 9.039 | 963.8215 | 1.798 | 1.81 | 1.804 | 196.4439 | 0.20382
2020-02-06 | 11:52 | 9.064 | 9.082 | 9.073 | 967.4469 | 1.806 | 1.809 | 1.8075 | 196.825 | 0.20345
2020-02-06 | 11:52 | 9.088 | 9.107 | 9.0975 | 970.0594 | 1.812 | 1.834 | 1.823 | 198.5129 | 0.20464
2020-02-06 | 11:52 | 9.057 | 9.131 | 9.094 | 969.6862 | 1.821 | 1.837 | 1.829 | 199.1662 | 0.20539
2020-02-06 | 11:52 | 8.987 | 9.14 | 9.0635 | 966.434 | 1.794 | 1.838 | 1.816 | 197.7506 | 0.20462
2020-02-06 | 11:52 | 8.638 | 8.789 | 8.7135 | 929.1137 | 1.72 | 1.755 | 1.7375 | 189.2025 | 0.20364
2020-02-06 | 11:53 | 7.866 | 8.663 | 8.2645 | 881.2372 | 1.554 | 1.708 | 1.631 | 177.6053 | 0.20154
2020-02-06 | 11:53 | 7.997 | 8.332 | 8.1645 | 870.5743 | 1.564 | 1.644 | 1.604 | 174.6652 | 0.20063
2020-02-06 | 11:53 | 5.454 | 7.67 | 6.562 | 699.701 | 1 | 1.473 | 1.2365 | 134.6468 | 0.19243
2020-02-06 | 11:53 | 5.302 | 7.599 | 6.4505 | 687.8118 | 0.985 | 1.562 | 1.2735 | 138.6759 | 0.20162
2020-02-06 | 11:53 | 8.235 | 8.907 | 8.571 | 913.9191 | 1.688 | 1.801 | 1.7445 | 189.9647 | 0.20786
2020-02-06 | 11:53 | 5.084 | 8.791 | 6.9375 | 739.7402 | 0.904 | 1.739 | 1.3215 | 143.9028 | 0.19453
2020-02-06 | 11:53 | 4.592 | 4.846 | 4.719 | 503.1833 | 0.82 | 0.912 | 0.866 | 94.3018 | 0.18741
2020-02-06 | 11:53 | 5.357 | 8.094 | 6.7255 | 717.1348 | 1.056 | 1.643 | 1.3495 | 146.9518 | 0.20492
2020-02-06 | 11:53 | 8.351 | 8.771 | 8.561 | 912.8528 | 1.681 | 1.773 | 1.727 | 188.0591 | 0.20601
2020-02-06 | 11:53 | 8.758 | 8.894 | 8.826 | 941.1095 | 1.756 | 1.793 | 1.7745 | 193.2315 | 0.20532
2020-02-06 | 11:53 | 8.545 | 8.652 | 8.5985 | 916.8514 | 1.717 | 1.745 | 1.731 | 188.4947 | 0.20559
2020-02-06 | 11:54 | 8.479 | 8.836 | 8.6575 | 923.1425 | 1.714 | 1.789 | 1.7515 | 190.727 | 0.20661
2020-02-06 | 11:54 | 8.859 | 8.915 | 8.887 | 947.6139 | 1.793 | 1.803 | 1.798 | 195.7905 | 0.20661
2020-02-06 | 11:54 | 8.785 | 8.914 | 8.8495 | 943.6153 | 1.774 | 1.802 | 1.788 | 194.7016 | 0.20634
2020-02-06 | 11:54 | 8.83 | 8.963 | 8.8965 | 948.6269 | 1.786 | 1.81 | 1.798 | 195.7905 | 0.20639
2020-02-06 | 11:54 | 8.963 | 8.968 | 8.9655 | 955.9843 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20606
2020-02-06 | 11:54 | 8.959 | 8.963 | 8.961 | 955.5045 | 1.808 | 1.808 | 1.808 | 196.8795 | 0.20605
2020-02-06 | 11:54 | 8.967 | 8.98 | 8.9735 | 956.8373 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20587
2020-02-06 | 11:54 | 8.982 | 8.988 | 8.985 | 958.0636 | 1.81 | 1.812 | 1.811 | 197.2061 | 0.20584
2020-02-06 | 11:54 | 8.99 | 8.994 | 8.992 | 958.81 | 1.812 | 1.814 | 1.813 | 197.4239 | 0.20591
2020-02-06 | 11:54 | 8.988 | 8.994 | 8.991 | 958.7033 | 1.81 | 1.813 | 1.8115 | 197.2606 | 0.20576
2020-02-06 | 11:54 | 8.977 | 8.986 | 8.9815 | 957.6904 | 1.81 | 1.812 | 1.811 | 197.2061 | 0.20592
2020-02-06 | 11:54 | 8.975 | 8.978 | 8.9765 | 957.1572 | 1.809 | 1.811 | 1.81 | 197.0972 | 0.20592
2020-02-06 | 11:54 | 8.976 | 8.98 | 8.978 | 957.3172 | 1.81 | 1.811 | 1.8105 | 197.1517 | 0.20594
2020-02-06 | 11:55 | 8.976 | 8.978 | 8.977 | 957.2105 | 1.81 | 1.811 | 1.8105 | 197.1517 | 0.20596
2020-02-06 | 11:55 | 8.974 | 8.976 | 8.975 | 956.9973 | 1.81 | 1.812 | 1.811 | 197.2061 | 0.20607
2020-02-06 | 11:55 | 8.974 | 8.975 | 8.9745 | 956.944 | 1.811 | 1.812 | 1.8115 | 197.2606 | 0.20614
2020-02-06 | 11:55 | 8.972 | 8.976 | 8.974 | 956.8906 | 1.811 | 1.813 | 1.812 | 197.315 | 0.2062
2020-02-06 | 11:55 | 8.97 | 8.976 | 8.973 | 956.784 | 1.811 | 1.812 | 1.8115 | 197.2606 | 0.20617
2020-02-06 | 11:55 | 8.955 | 8.969 | 8.962 | 955.6111 | 1.81 | 1.813 | 1.8115 | 197.2606 | 0.20642
2020-02-06 | 11:55 | 8.937 | 8.952 | 8.9445 | 953.7451 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20654
2020-02-06 | 11:55 | 8.921 | 8.934 | 8.9275 | 951.9324 | 1.805 | 1.809 | 1.807 | 196.7706 | 0.20671
2020-02-06 | 11:55 | 8.712 | 8.915 | 8.8135 | 939.7767 | 1.756 | 1.804 | 1.78 | 193.8304 | 0.20625
2020-02-06 | 11:55 | 7.912 | 8.679 | 8.2955 | 884.5427 | 1.536 | 1.748 | 1.642 | 178.8031 | 0.20214
2020-02-06 | 11:55 | 4.672 | 7.459 | 6.0655 | 646.7596 | 0.825 | 1.441 | 1.133 | 123.3763 | 0.19076
2020-02-06 | 11:56 | 4.029 | 4.358 | 4.1935 | 447.1496 | 0.718 | 0.77 | 0.744 | 81.0168 | 0.18118
2020-02-06 | 11:56 | 3.952 | 4.099 | 4.0255 | 429.2359 | 0.707 | 0.74 | 0.7235 | 78.7844 | 0.18355
2020-02-06 | 11:56 | 3.821 | 3.903 | 3.862 | 411.8021 | 0.685 | 0.698 | 0.6915 | 75.2999 | 0.18285
2020-02-06 | 11:56 | 3.768 | 3.808 | 3.788 | 403.9115 | 0.68 | 0.685 | 0.6825 | 74.3198 | 0.184
2020-02-06 | 11:56 | 3.734 | 3.759 | 3.7465 | 399.4864 | 0.675 | 0.679 | 0.677 | 73.7209 | 0.18454
2020-02-06 | 11:56 | 3.726 | 3.733 | 3.7295 | 397.6737 | 0.673 | 0.674 | 0.6735 | 73.3398 | 0.18442
2020-02-06 | 11:56 | 3.741 | 3.85 | 3.7955 | 404.7112 | 0.677 | 0.703 | 0.69 | 75.1365 | 0.18565
2020-02-06 | 11:56 | 3.91 | 4.189 | 4.0495 | 431.795 | 0.719 | 0.775 | 0.747 | 81.3434 | 0.18838
2020-02-06 | 11:56 | 3.92 | 4.368 | 4.144 | 441.8715 | 0.711 | 0.815 | 0.763 | 83.0857 | 0.18803
2020-02-06 | 11:56 | 4.361 | 4.755 | 4.558 | 486.016 | 0.805 | 0.904 | 0.8545 | 93.0495 | 0.19145
2020-02-06 | 11:56 | 4.52 | 5.208 | 4.864 | 518.6445 | 0.821 | 0.997 | 0.909 | 98.9842 | 0.19085
2020-02-06 | 11:56 | 3.943 | 4.212 | 4.0775 | 434.7807 | 0.714 | 0.761 | 0.7375 | 80.309 | 0.18471
2020-02-06 | 11:56 | 3.929 | 4.242 | 4.0855 | 435.6337 | 0.713 | 0.799 | 0.756 | 82.3235 | 0.18897
2020-02-06 | 11:57 | 4.641 | 6.386 | 5.5135 | 587.9002 | 0.909 | 1.269 | 1.089 | 118.585 | 0.20171
2020-02-06 | 11:57 | 6.064 | 6.542 | 6.303 | 672.084 | 1.18 | 1.329 | 1.2545 | 136.6069 | 0.20326
2020-02-06 | 11:57 | 5.693 | 6.949 | 6.321 | 674.0033 | 1.072 | 1.401 | 1.2365 | 134.6468 | 0.19977
2020-02-06 | 11:57 | 5.2 | 5.334 | 5.267 | 561.6161 | 0.995 | 1.008 | 1.0015 | 109.0568 | 0.19418
2020-02-06 | 11:57 | 3.649 | 4.913 | 4.281 | 456.4797 | 0.639 | 0.916 | 0.7775 | 84.6647 | 0.18547
2020-02-06 | 11:57 | 3.259 | 3.488 | 3.3735 | 359.7137 | 0.572 | 0.611 | 0.5915 | 64.4105 | 0.17906
2020-02-06 | 11:57 | 3.142 | 3.231 | 3.1865 | 339.774 | 0.551 | 0.569 | 0.56 | 60.9804 | 0.17947
2020-02-06 | 11:57 | 3.075 | 3.126 | 3.1005 | 330.6039 | 0.539 | 0.546 | 0.5425 | 59.0747 | 0.17869
2020-02-06 | 11:57 | 3.036 | 3.066 | 3.051 | 325.3258 | 0.532 | 0.536 | 0.534 | 58.1491 | 0.17874
2020-02-06 | 11:57 | 3.023 | 3.032 | 3.0275 | 322.82 | 0.531 | 0.532 | 0.5315 | 57.8769 | 0.17929
2020-02-06 | 11:57 | 3.022 | 3.024 | 3.023 | 322.3401 | 0.53 | 0.532 | 0.531 | 57.8224 | 0.17938
2020-02-06 | 11:58 | 3.024 | 3.03 | 3.027 | 322.7667 | 0.531 | 0.534 | 0.5325 | 57.9858 | 0.17965
2020-02-06 | 11:58 | 3.032 | 3.048 | 3.04 | 324.1528 | 0.535 | 0.537 | 0.536 | 58.3669 | 0.18006
2020-02-06 | 11:58 | 3.05 | 3.062 | 3.056 | 325.8589 | 0.538 | 0.541 | 0.5395 | 58.748 | 0.18029
2020-02-06 | 11:58 | 3.064 | 3.077 | 3.0705 | 327.405 | 0.541 | 0.546 | 0.5435 | 59.1836 | 0.18077
2020-02-06 | 11:58 | 3.084 | 3.129 | 3.1065 | 331.2437 | 0.549 | 0.558 | 0.5535 | 60.2726 | 0.18196
2020-02-06 | 11:58 | 3.143 | 3.224 | 3.1835 | 339.4541 | 0.561 | 0.577 | 0.569 | 61.9604 | 0.18253
2020-02-06 | 11:58 | 3.25 | 3.614 | 3.432 | 365.9515 | 0.581 | 0.68 | 0.6305 | 68.6574 | 0.18761
2020-02-06 | 11:58 | 3.985 | 5.049 | 4.517 | 481.6442 | 0.779 | 0.977 | 0.878 | 95.6085 | 0.1985
2020-02-06 | 11:58 | 4.598 | 5.342 | 4.97 | 529.9472 | 0.84 | 1.038 | 0.939 | 102.251 | 0.19295
2020-02-06 | 11:58 | 3.483 | 4.183 | 3.833 | 408.7098 | 0.624 | 0.748 | 0.686 | 74.7009 | 0.18277
2020-02-06 | 11:58 | 3.448 | 3.465 | 3.4565 | 368.5639 | 0.621 | 0.624 | 0.6225 | 67.7862 | 0.18392
2020-02-06 | 11:58 | 3.451 | 3.47 | 3.4605 | 368.9904 | 0.622 | 0.627 | 0.6245 | 68.004 | 0.1843
2020-02-06 | 11:58 | 3.437 | 3.451 | 3.444 | 367.231 | 0.619 | 0.623 | 0.621 | 67.6229 | 0.18414
2020-02-06 | 11:59 | 3.41 | 3.433 | 3.4215 | 364.8319 | 0.614 | 0.618 | 0.616 | 67.0784 | 0.18386
2020-02-06 | 11:59 | 3.361 | 3.4 | 3.3805 | 360.4601 | 0.605 | 0.613 | 0.609 | 66.3161 | 0.18398
2020-02-06 | 11:59 | 3.305 | 3.349 | 3.327 | 354.7554 | 0.594 | 0.601 | 0.5975 | 65.0639 | 0.1834
2020-02-06 | 11:59 | 3.266 | 3.296 | 3.281 | 349.8505 | 0.587 | 0.593 | 0.59 | 64.2472 | 0.18364
2020-02-06 | 11:59 | 3.237 | 3.26 | 3.2485 | 346.385 | 0.582 | 0.586 | 0.584 | 63.5938 | 0.18359
2020-02-06 | 11:59 | 3.23 | 3.272 | 3.251 | 346.6516 | 0.579 | 0.592 | 0.5855 | 63.7571 | 0.18392
2020-02-06 | 11:59 | 3.382 | 3.849 | 3.6155 | 385.518 | 0.625 | 0.73 | 0.6775 | 73.7753 | 0.19137
2020-02-06 | 11:59 | 3.318 | 3.905 | 3.6115 | 385.0914 | 0.593 | 0.739 | 0.666 | 72.5231 | 0.18833
2020-02-06 | 11:59 | 3.239 | 3.269 | 3.254 | 346.9715 | 0.579 | 0.584 | 0.5815 | 63.3216 | 0.1825
2020-02-06 | 11:59 | 3.247 | 3.343 | 3.295 | 351.3433 | 0.58 | 0.604 | 0.592 | 64.465 | 0.18348
2020-02-06 | 11:59 | 3.406 | 3.853 | 3.6295 | 387.0108 | 0.618 | 0.71 | 0.664 | 72.3053 | 0.18683
2020-02-06 | 12:00 | 3.862 | 4.235 | 4.0485 | 431.6884 | 0.713 | 0.826 | 0.7695 | 83.7935 | 0.19411
2020-02-06 | 12:00 | 4.721 | 5.657 | 5.189 | 553.299 | 0.951 | 1.122 | 1.0365 | 112.8681 | 0.20399
2020-02-06 | 12:00 | 3.687 | 5.343 | 4.515 | 481.4309 | 0.664 | 1.008 | 0.836 | 91.035 | 0.18909
2020-02-06 | 12:00 | 3.59 | 3.797 | 3.6935 | 393.835 | 0.643 | 0.693 | 0.668 | 72.7409 | 0.1847
2020-02-06 | 12:00 | 3.412 | 3.506 | 3.459 | 368.8305 | 0.614 | 0.628 | 0.621 | 67.6229 | 0.18334
2020-02-06 | 12:00 | 3.394 | 3.404 | 3.399 | 362.4327 | 0.612 | 0.614 | 0.613 | 66.7517 | 0.18418
2020-02-06 | 12:00 | 3.394 | 3.411 | 3.4025 | 362.8059 | 0.614 | 0.618 | 0.616 | 67.0784 | 0.18489
2020-02-06 | 12:00 | 3.418 | 3.45 | 3.434 | 366.1648 | 0.62 | 0.626 | 0.623 | 67.8407 | 0.18527
2020-02-06 | 12:00 | 3.459 | 3.515 | 3.487 | 371.8161 | 0.628 | 0.639 | 0.6335 | 68.984 | 0.18553
2020-02-06 | 12:00 | 3.531 | 3.61 | 3.5705 | 380.7196 | 0.642 | 0.656 | 0.649 | 70.6719 | 0.18563
2020-02-06 | 12:00 | 3.634 | 4.083 | 3.8585 | 411.4289 | 0.661 | 0.775 | 0.718 | 78.1855 | 0.19003
2020-02-06 | 12:00 | 4.423 | 5.489 | 4.956 | 528.4544 | 0.862 | 1.081 | 0.9715 | 105.79 | 0.20019
2020-02-06 | 12:00 | 5.532 | 6.311 | 5.9215 | 631.4049 | 1.074 | 1.25 | 1.162 | 126.5342 | 0.2004
2020-02-06 | 12:01 | 5.845 | 6.437 | 6.141 | 654.8101 | 1.141 | 1.29 | 1.2155 | 132.3601 | 0.20214
2020-02-06 | 12:01 | 3.924 | 5.93 | 4.927 | 525.3622 | 0.7 | 1.123 | 0.9115 | 99.2564 | 0.18893
2020-02-06 | 12:01 | 3.734 | 3.836 | 3.785 | 403.5916 | 0.674 | 0.688 | 0.681 | 74.1565 | 0.18374
2020-02-06 | 12:01 | 3.709 | 3.726 | 3.7175 | 396.3941 | 0.671 | 0.673 | 0.672 | 73.1764 | 0.18461
2020-02-06 | 12:01 | 3.711 | 3.726 | 3.7185 | 396.5008 | 0.671 | 0.675 | 0.673 | 73.2853 | 0.18483
2020-02-06 | 12:01 | 3.736 | 3.865 | 3.8005 | 405.2444 | 0.677 | 0.71 | 0.6935 | 75.5176 | 0.18635
2020-02-06 | 12:01 | 3.938 | 4.02 | 3.979 | 424.2777 | 0.724 | 0.739 | 0.7315 | 79.6556 | 0.18774
2020-02-06 | 12:01 | 3.987 | 4.078 | 4.0325 | 429.9823 | 0.725 | 0.746 | 0.7355 | 80.0912 | 0.18627
2020-02-06 | 12:01 | 4.169 | 5.216 | 4.6925 | 500.3576 | 0.769 | 1.005 | 0.887 | 96.5885 | 0.19304
2020-02-06 | 12:01 | 5.405 | 6.775 | 6.09 | 649.372 | 1.048 | 1.349 | 1.1985 | 130.5089 | 0.20098
2020-02-06 | 12:01 | 5.252 | 6.635 | 5.9435 | 633.7508 | 0.963 | 1.288 | 1.1255 | 122.5596 | 0.19339
2020-02-06 | 12:02 | 4.84 | 5.288 | 5.064 | 539.9704 | 0.899 | 0.993 | 0.946 | 103.0133 | 0.19078
2020-02-06 | 12:02 | 4.25 | 4.597 | 4.4235 | 471.6744 | 0.766 | 0.825 | 0.7955 | 86.6248 | 0.18365
2020-02-06 | 12:02 | 4.376 | 4.612 | 4.494 | 479.1917 | 0.805 | 0.854 | 0.8295 | 90.3272 | 0.1885
2020-02-06 | 12:02 | 4.426 | 5.354 | 4.89 | 521.4169 | 0.808 | 1.093 | 0.9505 | 103.5033 | 0.1985
2020-02-06 | 12:02 | 6.612 | 8.639 | 7.6255 | 813.1011 | 1.38 | 1.763 | 1.5715 | 171.1261 | 0.21046
2020-02-06 | 12:02 | 8.766 | 8.875 | 8.8205 | 940.5231 | 1.784 | 1.795 | 1.7895 | 194.8649 | 0.20719
2020-02-06 | 12:02 | 8.873 | 8.89 | 8.8815 | 947.0274 | 1.789 | 1.795 | 1.792 | 195.1372 | 0.20605
2020-02-06 | 12:02 | 8.777 | 8.863 | 8.82 | 940.4697 | 1.771 | 1.79 | 1.7805 | 193.8849 | 0.20616
2020-02-06 | 12:02 | 8.874 | 8.878 | 8.876 | 946.441 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20606
2020-02-06 | 12:02 | 8.869 | 8.873 | 8.871 | 945.9078 | 1.787 | 1.789 | 1.788 | 194.7016 | 0.20584
2020-02-06 | 12:02 | 8.861 | 8.871 | 8.866 | 945.3747 | 1.785 | 1.787 | 1.786 | 194.4838 | 0.20572
2020-02-06 | 12:02 | 8.855 | 8.858 | 8.8565 | 944.3617 | 1.782 | 1.783 | 1.7825 | 194.1027 | 0.20554
2020-02-06 | 12:02 | 8.86 | 8.874 | 8.867 | 945.4813 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.20535
2020-02-06 | 12:03 | 8.877 | 8.891 | 8.884 | 947.294 | 1.784 | 1.785 | 1.7845 | 194.3205 | 0.20513
2020-02-06 | 12:03 | 8.894 | 8.912 | 8.903 | 949.32 | 1.785 | 1.79 | 1.7875 | 194.6471 | 0.20504
2020-02-06 | 12:03 | 8.919 | 8.943 | 8.931 | 952.3056 | 1.791 | 1.794 | 1.7925 | 195.1916 | 0.20497
2020-02-06 | 12:03 | 8.947 | 8.961 | 8.954 | 954.7581 | 1.794 | 1.798 | 1.796 | 195.5727 | 0.20484
2020-02-06 | 12:03 | 8.962 | 8.973 | 8.9675 | 956.1976 | 1.796 | 1.799 | 1.7975 | 195.7361 | 0.2047
2020-02-06 | 12:03 | 8.975 | 8.986 | 8.9805 | 957.5837 | 1.798 | 1.801 | 1.7995 | 195.9539 | 0.20463
2020-02-06 | 12:03 | 8.986 | 8.991 | 8.9885 | 958.4368 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20451
2020-02-06 | 12:03 | 8.99 | 8.995 | 8.9925 | 958.8633 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.20464
2020-02-06 | 12:03 | 8.995 | 9.001 | 8.998 | 959.4497 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20463
2020-02-06 | 12:03 | 9.001 | 9.006 | 9.0035 | 960.0362 | 1.804 | 1.805 | 1.8045 | 196.4983 | 0.20468
2020-02-06 | 12:03 | 9.008 | 9.011 | 9.0095 | 960.676 | 1.804 | 1.805 | 1.8045 | 196.4983 | 0.20454
2020-02-06 | 12:04 | 9.002 | 9.01 | 9.006 | 960.3028 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20445
2020-02-06 | 12:04 | 8.999 | 9.002 | 9.0005 | 959.7163 | 1.802 | 1.803 | 1.8025 | 196.2805 | 0.20452
2020-02-06 | 12:04 | 8.99 | 9 | 8.995 | 959.1299 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.20459
2020-02-06 | 12:04 | 8.977 | 8.988 | 8.9825 | 957.797 | 1.798 | 1.8 | 1.799 | 195.8994 | 0.20453
2020-02-06 | 12:04 | 8.975 | 8.981 | 8.978 | 957.3172 | 1.798 | 1.799 | 1.7985 | 195.845 | 0.20458
2020-02-06 | 12:04 | 8.982 | 8.989 | 8.9855 | 958.1169 | 1.799 | 1.8 | 1.7995 | 195.9539 | 0.20452
2020-02-06 | 12:04 | 8.991 | 8.999 | 8.995 | 959.1299 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20436
2020-02-06 | 12:04 | 8.998 | 9.008 | 9.003 | 959.9829 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.20441
2020-02-06 | 12:04 | 9.01 | 9.011 | 9.0105 | 960.7826 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.20424
2020-02-06 | 12:04 | 9.005 | 9.014 | 9.0095 | 960.676 | 1.801 | 1.804 | 1.8025 | 196.2805 | 0.20432
2020-02-06 | 12:04 | 9.011 | 9.016 | 9.0135 | 961.1025 | 1.803 | 1.805 | 1.804 | 196.4439 | 0.20439
2020-02-06 | 12:04 | 9.004 | 9.01 | 9.007 | 960.4094 | 1.8 | 1.802 | 1.801 | 196.1172 | 0.2042
2020-02-06 | 12:04 | 8.997 | 9.002 | 8.9995 | 959.6097 | 1.8 | 1.8 | 1.8 | 196.0083 | 0.20426
2020-02-06 | 12:05 | 8.995 | 8.998 | 8.9965 | 959.2898 | 1.799 | 1.8 | 1.7995 | 195.9539 | 0.20427
2020-02-06 | 12:05 | 8.992 | 8.995 | 8.9935 | 958.9699 | 1.798 | 1.799 | 1.7985 | 195.845 | 0.20422
2020-02-06 | 12:05 | 8.985 | 8.99 | 8.9875 | 958.3301 | 1.797 | 1.798 | 1.7975 | 195.7361 | 0.20425
2020-02-06 | 12:05 | 8.982 | 8.983 | 8.9825 | 957.797 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.2043
2020-02-06 | 12:05 | 8.982 | 8.983 | 8.9825 | 957.797 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20419
2020-02-06 | 12:05 | 8.98 | 8.985 | 8.9825 | 957.797 | 1.795 | 1.796 | 1.7955 | 195.5183 | 0.20413
2020-02-06 | 12:05 | 8.983 | 8.987 | 8.985 | 958.0636 | 1.795 | 1.796 | 1.7955 | 195.5183 | 0.20408
2020-02-06 | 12:05 | 8.98 | 8.984 | 8.982 | 957.7437 | 1.794 | 1.796 | 1.795 | 195.4638 | 0.20409
2020-02-06 | 12:05 | 8.981 | 8.983 | 8.982 | 957.7437 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.2042
2020-02-06 | 12:05 | 8.982 | 8.984 | 8.983 | 957.8503 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20418
2020-02-06 | 12:05 | 8.983 | 8.988 | 8.9855 | 958.1169 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20418
2020-02-06 | 12:06 | 8.987 | 8.989 | 8.988 | 958.3835 | 1.797 | 1.798 | 1.7975 | 195.7361 | 0.20424
2020-02-06 | 12:06 | 8.98 | 8.987 | 8.9835 | 957.9036 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20417
2020-02-06 | 12:06 | 8.98 | 8.982 | 8.981 | 957.6371 | 1.795 | 1.796 | 1.7955 | 195.5183 | 0.20417
2020-02-06 | 12:06 | 8.98 | 8.983 | 8.9815 | 957.6904 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20421
2020-02-06 | 12:06 | 8.98 | 8.986 | 8.983 | 957.8503 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20418
2020-02-06 | 12:06 | 8.985 | 8.987 | 8.986 | 958.1702 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20417
2020-02-06 | 12:06 | 8.985 | 8.988 | 8.9865 | 958.2235 | 1.795 | 1.798 | 1.7965 | 195.6272 | 0.20416
2020-02-06 | 12:06 | 8.987 | 8.988 | 8.9875 | 958.3301 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20419
2020-02-06 | 12:06 | 8.988 | 8.992 | 8.99 | 958.5967 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20425
2020-02-06 | 12:06 | 8.993 | 8.996 | 8.9945 | 959.0765 | 1.798 | 1.799 | 1.7985 | 195.845 | 0.2042
2020-02-06 | 12:06 | 8.995 | 8.999 | 8.997 | 959.3431 | 1.798 | 1.8 | 1.799 | 195.8994 | 0.2042
2020-02-06 | 12:06 | 8.997 | 8.999 | 8.998 | 959.4497 | 1.798 | 1.8 | 1.799 | 195.8994 | 0.20418
2020-02-06 | 12:06 | 9.001 | 9.004 | 9.0025 | 959.9296 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20419
2020-02-06 | 12:07 | 9.004 | 9.006 | 9.005 | 960.1962 | 1.8 | 1.802 | 1.801 | 196.1172 | 0.20425
2020-02-06 | 12:07 | 9.004 | 9.008 | 9.006 | 960.3028 | 1.8 | 1.802 | 1.801 | 196.1172 | 0.20422
2020-02-06 | 12:07 | 9.006 | 9.009 | 9.0075 | 960.4627 | 1.801 | 1.802 | 1.8015 | 196.1716 | 0.20425
2020-02-06 | 12:07 | 9.011 | 9.013 | 9.012 | 960.9426 | 1.803 | 1.804 | 1.8035 | 196.3894 | 0.20437
2020-02-06 | 12:07 | 9.01 | 9.014 | 9.012 | 960.9426 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20432
2020-02-06 | 12:07 | 9.011 | 9.012 | 9.0115 | 960.8892 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20433
2020-02-06 | 12:07 | 9.01 | 9.012 | 9.011 | 960.8359 | 1.802 | 1.803 | 1.8025 | 196.2805 | 0.20428
2020-02-06 | 12:07 | 9.005 | 9.009 | 9.007 | 960.4094 | 1.801 | 1.802 | 1.8015 | 196.1716 | 0.20426
2020-02-06 | 12:07 | 9.002 | 9.005 | 9.0035 | 960.0362 | 1.8 | 1.802 | 1.801 | 196.1172 | 0.20428
2020-02-06 | 12:07 | 9.001 | 9.004 | 9.0025 | 959.9296 | 1.8 | 1.801 | 1.8005 | 196.0628 | 0.20425
2020-02-06 | 12:07 | 9 | 9.002 | 9.001 | 959.7696 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20422
2020-02-06 | 12:08 | 8.994 | 9.002 | 8.998 | 959.4497 | 1.797 | 1.8 | 1.7985 | 195.845 | 0.20412
2020-02-06 | 12:08 | 8.989 | 8.994 | 8.9915 | 958.7567 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.2041
2020-02-06 | 12:08 | 8.985 | 8.989 | 8.987 | 958.2768 | 1.794 | 1.797 | 1.7955 | 195.5183 | 0.20403
2020-02-06 | 12:08 | 8.979 | 8.983 | 8.981 | 957.6371 | 1.794 | 1.796 | 1.795 | 195.4638 | 0.20411
2020-02-06 | 12:08 | 8.975 | 8.98 | 8.9775 | 957.2638 | 1.794 | 1.795 | 1.7945 | 195.4094 | 0.20413
2020-02-06 | 12:08 | 8.969 | 8.975 | 8.972 | 956.6774 | 1.793 | 1.794 | 1.7935 | 195.3005 | 0.20414
2020-02-06 | 12:08 | 8.968 | 8.97 | 8.969 | 956.3575 | 1.793 | 1.794 | 1.7935 | 195.3005 | 0.20421
2020-02-06 | 12:08 | 8.957 | 8.968 | 8.9625 | 955.6644 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20419
2020-02-06 | 12:08 | 8.949 | 8.953 | 8.951 | 954.4382 | 1.789 | 1.791 | 1.79 | 194.9194 | 0.20422
2020-02-06 | 12:08 | 8.951 | 8.955 | 8.953 | 954.6514 | 1.781 | 1.79 | 1.7855 | 194.4293 | 0.20367
2020-02-06 | 12:08 | 8.945 | 8.949 | 8.947 | 954.0117 | 1.764 | 1.772 | 1.768 | 192.5237 | 0.2018
2020-02-06 | 12:08 | 8.937 | 8.943 | 8.94 | 953.2653 | 1.771 | 1.779 | 1.775 | 193.286 | 0.20276
2020-02-06 | 12:08 | 8.937 | 8.941 | 8.939 | 953.1586 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.2037
2020-02-06 | 12:09 | 8.932 | 8.937 | 8.9345 | 952.6788 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.2038
2020-02-06 | 12:09 | 8.927 | 8.933 | 8.93 | 952.199 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.2039
2020-02-06 | 12:09 | 8.921 | 8.926 | 8.9235 | 951.5059 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.20405
2020-02-06 | 12:09 | 8.916 | 8.922 | 8.919 | 951.026 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.20416
2020-02-06 | 12:09 | 8.915 | 8.917 | 8.916 | 950.7062 | 1.783 | 1.784 | 1.7835 | 194.2116 | 0.20428
2020-02-06 | 12:09 | 8.912 | 8.916 | 8.914 | 950.4929 | 1.783 | 1.784 | 1.7835 | 194.2116 | 0.20433
2020-02-06 | 12:09 | 8.906 | 8.912 | 8.909 | 949.9597 | 1.782 | 1.783 | 1.7825 | 194.1027 | 0.20433
2020-02-06 | 12:09 | 8.906 | 8.909 | 8.9075 | 949.7998 | 1.781 | 1.782 | 1.7815 | 193.9938 | 0.20425
2020-02-06 | 12:09 | 8.91 | 8.913 | 8.9115 | 950.2263 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.20433
2020-02-06 | 12:09 | 8.91 | 8.915 | 8.9125 | 950.3329 | 1.782 | 1.785 | 1.7835 | 194.2116 | 0.20436
2020-02-06 | 12:09 | 8.917 | 8.919 | 8.918 | 950.9194 | 1.785 | 1.787 | 1.786 | 194.4838 | 0.20452
2020-02-06 | 12:10 | 8.921 | 8.927 | 8.924 | 951.5592 | 1.786 | 1.788 | 1.787 | 194.5927 | 0.2045
2020-02-06 | 12:10 | 8.929 | 8.936 | 8.9325 | 952.4655 | 1.787 | 1.79 | 1.7885 | 194.756 | 0.20448
2020-02-06 | 12:10 | 8.938 | 8.946 | 8.942 | 953.4785 | 1.789 | 1.791 | 1.79 | 194.9194 | 0.20443
2020-02-06 | 12:10 | 8.948 | 8.952 | 8.95 | 954.3315 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20448
2020-02-06 | 12:10 | 8.952 | 8.955 | 8.9535 | 954.7047 | 1.792 | 1.794 | 1.793 | 195.246 | 0.20451
2020-02-06 | 12:10 | 8.955 | 8.957 | 8.956 | 954.9713 | 1.795 | 1.796 | 1.7955 | 195.5183 | 0.20474
2020-02-06 | 12:10 | 8.959 | 8.961 | 8.96 | 955.3978 | 1.794 | 1.797 | 1.7955 | 195.5183 | 0.20465
2020-02-06 | 12:10 | 8.961 | 8.963 | 8.962 | 955.6111 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20477
2020-02-06 | 12:10 | 8.963 | 8.967 | 8.965 | 955.931 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20482
2020-02-06 | 12:10 | 8.962 | 8.968 | 8.965 | 955.931 | 1.796 | 1.799 | 1.7975 | 195.7361 | 0.20476
2020-02-06 | 12:10 | 8.96 | 8.963 | 8.9615 | 955.5578 | 1.797 | 1.798 | 1.7975 | 195.7361 | 0.20484
2020-02-06 | 12:10 | 8.954 | 8.96 | 8.957 | 955.078 | 1.795 | 1.798 | 1.7965 | 195.6272 | 0.20483
2020-02-06 | 12:10 | 8.954 | 8.957 | 8.9555 | 954.918 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20492
2020-02-06 | 12:11 | 8.955 | 8.957 | 8.956 | 954.9713 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20491
2020-02-06 | 12:11 | 8.959 | 8.961 | 8.96 | 955.3978 | 1.798 | 1.798 | 1.798 | 195.7905 | 0.20493
2020-02-06 | 12:11 | 8.956 | 8.96 | 8.958 | 955.1846 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20498
2020-02-06 | 12:11 | 8.961 | 8.964 | 8.9625 | 955.6644 | 1.798 | 1.8 | 1.799 | 195.8994 | 0.20499
2020-02-06 | 12:11 | 8.963 | 8.967 | 8.965 | 955.931 | 1.801 | 1.802 | 1.8015 | 196.1716 | 0.20522
2020-02-06 | 12:11 | 8.966 | 8.968 | 8.967 | 956.1442 | 1.802 | 1.803 | 1.8025 | 196.2805 | 0.20528
2020-02-06 | 12:11 | 8.966 | 8.973 | 8.9695 | 956.4108 | 1.803 | 1.805 | 1.804 | 196.4439 | 0.2054
2020-02-06 | 12:11 | 8.973 | 8.98 | 8.9765 | 957.1572 | 1.805 | 1.807 | 1.806 | 196.6617 | 0.20546
2020-02-06 | 12:11 | 8.979 | 8.982 | 8.9805 | 957.5837 | 1.805 | 1.808 | 1.8065 | 196.7161 | 0.20543
2020-02-06 | 12:11 | 8.978 | 8.98 | 8.979 | 957.4238 | 1.804 | 1.806 | 1.805 | 196.5528 | 0.20529
2020-02-06 | 12:11 | 8.979 | 8.984 | 8.9815 | 957.6904 | 1.806 | 1.807 | 1.8065 | 196.7161 | 0.20541
2020-02-06 | 12:12 | 8.986 | 8.994 | 8.99 | 958.5967 | 1.806 | 1.809 | 1.8075 | 196.825 | 0.20533
2020-02-06 | 12:12 | 8.996 | 9.005 | 9.0005 | 959.7163 | 1.809 | 1.813 | 1.811 | 197.2061 | 0.20548
2020-02-06 | 12:12 | 9.007 | 9.019 | 9.013 | 961.0492 | 1.812 | 1.815 | 1.8135 | 197.4784 | 0.20548
2020-02-06 | 12:12 | 9.022 | 9.025 | 9.0235 | 962.1688 | 1.814 | 1.816 | 1.815 | 197.6417 | 0.20541
2020-02-06 | 12:12 | 9.017 | 9.022 | 9.0195 | 961.7423 | 1.815 | 1.817 | 1.816 | 197.7506 | 0.20562
2020-02-06 | 12:12 | 9.012 | 9.019 | 9.0155 | 961.3158 | 1.815 | 1.815 | 1.815 | 197.6417 | 0.20559
2020-02-06 | 12:12 | 9.012 | 9.014 | 9.013 | 961.0492 | 1.814 | 1.816 | 1.815 | 197.6417 | 0.20565
2020-02-06 | 12:12 | 9.016 | 9.018 | 9.017 | 961.4757 | 1.815 | 1.817 | 1.816 | 197.7506 | 0.20567
2020-02-06 | 12:12 | 9.015 | 9.018 | 9.0165 | 961.4224 | 1.815 | 1.817 | 1.816 | 197.7506 | 0.20569
2020-02-06 | 12:12 | 9.016 | 9.018 | 9.017 | 961.4757 | 1.815 | 1.816 | 1.8155 | 197.6962 | 0.20562
2020-02-06 | 12:12 | 9.008 | 9.017 | 9.0125 | 960.9959 | 1.813 | 1.816 | 1.8145 | 197.5873 | 0.20561
2020-02-06 | 12:12 | 8.997 | 9.005 | 9.001 | 959.7696 | 1.812 | 1.814 | 1.813 | 197.4239 | 0.2057
2020-02-06 | 12:12 | 8.998 | 8.999 | 8.9985 | 959.5031 | 1.811 | 1.813 | 1.812 | 197.315 | 0.20564
2020-02-06 | 12:13 | 9.001 | 9.006 | 9.0035 | 960.0362 | 1.813 | 1.815 | 1.814 | 197.5328 | 0.20576
2020-02-06 | 12:13 | 9.005 | 9.009 | 9.007 | 960.4094 | 1.814 | 1.815 | 1.8145 | 197.5873 | 0.20573
2020-02-06 | 12:13 | 9.005 | 9.008 | 9.0065 | 960.3561 | 1.814 | 1.815 | 1.8145 | 197.5873 | 0.20574
2020-02-06 | 12:13 | 9.011 | 9.023 | 9.017 | 961.4757 | 1.816 | 1.819 | 1.8175 | 197.9139 | 0.20584
2020-02-06 | 12:13 | 9.025 | 9.032 | 9.0285 | 962.7019 | 1.817 | 1.819 | 1.818 | 197.9684 | 0.20564
2020-02-06 | 12:13 | 9.034 | 9.035 | 9.0345 | 963.3417 | 1.818 | 1.819 | 1.8185 | 198.0228 | 0.20556
2020-02-06 | 12:13 | 9.032 | 9.035 | 9.0335 | 963.2351 | 1.818 | 1.82 | 1.819 | 198.0773 | 0.20564
2020-02-06 | 12:13 | 9.029 | 9.033 | 9.031 | 962.9685 | 1.816 | 1.819 | 1.8175 | 197.9139 | 0.20552
2020-02-06 | 12:13 | 9.016 | 9.026 | 9.021 | 961.9022 | 1.814 | 1.817 | 1.8155 | 197.6962 | 0.20553
2020-02-06 | 12:13 | 9.015 | 9.018 | 9.0165 | 961.4224 | 1.813 | 1.816 | 1.8145 | 197.5873 | 0.20552
2020-02-06 | 12:13 | 9.018 | 9.02 | 9.019 | 961.689 | 1.814 | 1.815 | 1.8145 | 197.5873 | 0.20546
2020-02-06 | 12:14 | 9.016 | 9.018 | 9.017 | 961.4757 | 1.813 | 1.815 | 1.814 | 197.5328 | 0.20545
2020-02-06 | 12:14 | 9.01 | 9.016 | 9.013 | 961.0492 | 1.812 | 1.813 | 1.8125 | 197.3695 | 0.20537
2020-02-06 | 12:14 | 9.001 | 9.005 | 9.003 | 959.9829 | 1.812 | 1.813 | 1.8125 | 197.3695 | 0.2056
2020-02-06 | 12:14 | 8.996 | 9.001 | 8.9985 | 959.5031 | 1.811 | 1.812 | 1.8115 | 197.2606 | 0.20559
2020-02-06 | 12:14 | 8.992 | 8.996 | 8.994 | 959.0232 | 1.81 | 1.811 | 1.8105 | 197.1517 | 0.20558
2020-02-06 | 12:14 | 8.99 | 8.993 | 8.9915 | 958.7567 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20546
2020-02-06 | 12:14 | 8.99 | 8.992 | 8.991 | 958.7033 | 1.808 | 1.811 | 1.8095 | 197.0428 | 0.20553
2020-02-06 | 12:14 | 8.991 | 8.993 | 8.992 | 958.81 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20545
2020-02-06 | 12:14 | 8.989 | 8.992 | 8.9905 | 958.65 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20549
2020-02-06 | 12:14 | 8.988 | 8.992 | 8.99 | 958.5967 | 1.809 | 1.81 | 1.8095 | 197.0428 | 0.20555
2020-02-06 | 12:14 | 8.987 | 8.991 | 8.989 | 958.4901 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20552
2020-02-06 | 12:14 | 8.991 | 8.993 | 8.992 | 958.81 | 1.809 | 1.81 | 1.8095 | 197.0428 | 0.20551
2020-02-06 | 12:14 | 8.99 | 8.993 | 8.9915 | 958.7567 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20546
2020-02-06 | 12:15 | 8.985 | 8.99 | 8.9875 | 958.3301 | 1.807 | 1.808 | 1.8075 | 196.825 | 0.20538
2020-02-06 | 12:15 | 8.977 | 8.985 | 8.981 | 957.6371 | 1.807 | 1.808 | 1.8075 | 196.825 | 0.20553
2020-02-06 | 12:15 | 8.967 | 8.976 | 8.9715 | 956.6241 | 1.804 | 1.806 | 1.805 | 196.5528 | 0.20547
2020-02-06 | 12:15 | 8.96 | 8.966 | 8.963 | 955.7177 | 1.802 | 1.805 | 1.8035 | 196.3894 | 0.20549
2020-02-06 | 12:15 | 8.954 | 8.957 | 8.9555 | 954.918 | 1.801 | 1.802 | 1.8015 | 196.1716 | 0.20543
2020-02-06 | 12:15 | 8.943 | 8.952 | 8.9475 | 954.065 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20545
2020-02-06 | 12:15 | 8.944 | 8.947 | 8.9455 | 953.8517 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20549
2020-02-06 | 12:15 | 8.936 | 8.942 | 8.939 | 953.1586 | 1.799 | 1.799 | 1.799 | 195.8994 | 0.20553
2020-02-06 | 12:15 | 8.928 | 8.934 | 8.931 | 952.3056 | 1.797 | 1.798 | 1.7975 | 195.7361 | 0.20554
2020-02-06 | 12:15 | 8.92 | 8.926 | 8.923 | 951.4526 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20561
2020-02-06 | 12:15 | 8.916 | 8.921 | 8.9185 | 950.9727 | 1.793 | 1.795 | 1.794 | 195.3549 | 0.20543
2020-02-06 | 12:16 | 8.905 | 8.912 | 8.9085 | 949.9064 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20543
2020-02-06 | 12:16 | 8.904 | 8.905 | 8.9045 | 949.4799 | 1.791 | 1.792 | 1.7915 | 195.0827 | 0.20546
2020-02-06 | 12:16 | 8.904 | 8.906 | 8.905 | 949.5332 | 1.79 | 1.793 | 1.7915 | 195.0827 | 0.20545
2020-02-06 | 12:16 | 8.906 | 8.915 | 8.9105 | 950.1197 | 1.791 | 1.794 | 1.7925 | 195.1916 | 0.20544
2020-02-06 | 12:16 | 8.919 | 8.932 | 8.9255 | 951.7191 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20555
2020-02-06 | 12:16 | 8.934 | 8.941 | 8.9375 | 952.9987 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20545
2020-02-06 | 12:16 | 8.942 | 8.951 | 8.9465 | 953.9583 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.20547
2020-02-06 | 12:16 | 8.957 | 8.966 | 8.9615 | 955.5578 | 1.802 | 1.803 | 1.8025 | 196.2805 | 0.20541
2020-02-06 | 12:16 | 8.965 | 8.97 | 8.9675 | 956.1976 | 1.804 | 1.805 | 1.8045 | 196.4983 | 0.2055
2020-02-06 | 12:16 | 8.973 | 8.978 | 8.9755 | 957.0506 | 1.806 | 1.807 | 1.8065 | 196.7161 | 0.20554
2020-02-06 | 12:16 | 8.98 | 8.985 | 8.9825 | 957.797 | 1.807 | 1.808 | 1.8075 | 196.825 | 0.2055
2020-02-06 | 12:16 | 8.981 | 8.986 | 8.9835 | 957.9036 | 1.806 | 1.808 | 1.807 | 196.7706 | 0.20542
2020-02-06 | 12:16 | 8.985 | 8.986 | 8.9855 | 958.1169 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.2056
2020-02-06 | 12:17 | 8.985 | 8.988 | 8.9865 | 958.2235 | 1.808 | 1.809 | 1.8085 | 196.9339 | 0.20552
2020-02-06 | 12:17 | 8.987 | 8.992 | 8.9895 | 958.5434 | 1.81 | 1.812 | 1.811 | 197.2061 | 0.20574
2020-02-06 | 12:17 | 8.993 | 8.998 | 8.9955 | 959.1832 | 1.812 | 1.813 | 1.8125 | 197.3695 | 0.20577
2020-02-06 | 12:17 | 8.999 | 9.003 | 9.001 | 959.7696 | 1.813 | 1.814 | 1.8135 | 197.4784 | 0.20576
2020-02-06 | 12:17 | 9.001 | 9.003 | 9.002 | 959.8763 | 1.814 | 1.816 | 1.815 | 197.6417 | 0.2059
2020-02-06 | 12:17 | 8.991 | 8.999 | 8.995 | 959.1299 | 1.814 | 1.816 | 1.815 | 197.6417 | 0.20606
2020-02-06 | 12:17 | 8.976 | 8.987 | 8.9815 | 957.6904 | 1.812 | 1.815 | 1.8135 | 197.4784 | 0.2062
2020-02-06 | 12:17 | 8.96 | 8.971 | 8.9655 | 955.9843 | 1.81 | 1.813 | 1.8115 | 197.2606 | 0.20634
2020-02-06 | 12:17 | 8.939 | 8.956 | 8.9475 | 954.065 | 1.807 | 1.809 | 1.808 | 196.8795 | 0.20636
2020-02-06 | 12:17 | 8.922 | 8.935 | 8.9285 | 952.039 | 1.804 | 1.806 | 1.805 | 196.5528 | 0.20645
2020-02-06 | 12:17 | 8.899 | 8.918 | 8.9085 | 949.9064 | 1.8 | 1.804 | 1.802 | 196.2261 | 0.20657
2020-02-06 | 12:18 | 8.866 | 8.895 | 8.8805 | 946.9208 | 1.793 | 1.799 | 1.796 | 195.5727 | 0.20654
2020-02-06 | 12:18 | 8.843 | 8.861 | 8.852 | 943.8819 | 1.79 | 1.793 | 1.7915 | 195.0827 | 0.20668
2020-02-06 | 12:18 | 8.832 | 8.839 | 8.8355 | 942.1225 | 1.789 | 1.79 | 1.7895 | 194.8649 | 0.20684
2020-02-06 | 12:18 | 8.82 | 8.831 | 8.8255 | 941.0562 | 1.787 | 1.787 | 1.787 | 194.5927 | 0.20678
2020-02-06 | 12:18 | 8.805 | 8.819 | 8.812 | 939.6167 | 1.785 | 1.787 | 1.786 | 194.4838 | 0.20698
2020-02-06 | 12:18 | 8.787 | 8.802 | 8.7945 | 937.7507 | 1.781 | 1.785 | 1.783 | 194.1571 | 0.20705
2020-02-06 | 12:18 | 8.764 | 8.782 | 8.773 | 935.4582 | 1.776 | 1.78 | 1.778 | 193.6126 | 0.20697
2020-02-06 | 12:18 | 8.755 | 8.761 | 8.758 | 933.8587 | 1.775 | 1.778 | 1.7765 | 193.4493 | 0.20715
2020-02-06 | 12:18 | 8.744 | 8.757 | 8.7505 | 933.059 | 1.774 | 1.776 | 1.775 | 193.286 | 0.20715
2020-02-06 | 12:18 | 8.695 | 8.736 | 8.7155 | 929.327 | 1.763 | 1.771 | 1.767 | 192.4148 | 0.20705
2020-02-06 | 12:18 | 8.434 | 8.672 | 8.553 | 911.9997 | 1.7 | 1.757 | 1.7285 | 188.2224 | 0.20638
2020-02-06 | 12:18 | 7.739 | 8.328 | 8.0335 | 856.6059 | 1.538 | 1.676 | 1.607 | 174.9919 | 0.20429
2020-02-06 | 12:18 | 6.192 | 7.545 | 6.8685 | 732.3828 | 1.183 | 1.495 | 1.339 | 145.8084 | 0.19909
2020-02-06 | 12:19 | 4.51 | 5.664 | 5.087 | 542.4229 | 0.831 | 1.07 | 0.9505 | 103.5033 | 0.19082
2020-02-06 | 12:19 | 3.841 | 4.259 | 4.05 | 431.8484 | 0.696 | 0.775 | 0.7355 | 80.0912 | 0.18546
2020-02-06 | 12:19 | 3.783 | 3.821 | 3.802 | 405.4043 | 0.686 | 0.692 | 0.689 | 75.0276 | 0.18507
2020-02-06 | 12:19 | 3.737 | 3.776 | 3.7565 | 400.5527 | 0.677 | 0.685 | 0.681 | 74.1565 | 0.18514
2020-02-06 | 12:19 | 3.706 | 3.728 | 3.717 | 396.3408 | 0.673 | 0.676 | 0.6745 | 73.4487 | 0.18532
2020-02-06 | 12:19 | 3.684 | 3.702 | 3.693 | 393.7817 | 0.67 | 0.672 | 0.671 | 73.0675 | 0.18555
2020-02-06 | 12:19 | 3.668 | 3.679 | 3.6735 | 391.7025 | 0.666 | 0.668 | 0.667 | 72.632 | 0.18543
2020-02-06 | 12:19 | 3.664 | 3.667 | 3.6655 | 390.8494 | 0.664 | 0.666 | 0.665 | 72.4142 | 0.18527
2020-02-06 | 12:19 | 3.661 | 3.668 | 3.6645 | 390.7428 | 0.665 | 0.666 | 0.6655 | 72.4686 | 0.18546
2020-02-06 | 12:19 | 3.668 | 3.705 | 3.6865 | 393.0886 | 0.666 | 0.676 | 0.671 | 73.0675 | 0.18588
2020-02-06 | 12:19 | 3.729 | 3.798 | 3.7635 | 401.2991 | 0.682 | 0.696 | 0.689 | 75.0276 | 0.18696
2020-02-06 | 12:20 | 3.814 | 3.847 | 3.8305 | 408.4432 | 0.696 | 0.705 | 0.7005 | 76.2799 | 0.18676
2020-02-06 | 12:20 | 3.777 | 3.793 | 3.785 | 403.5916 | 0.69 | 0.694 | 0.692 | 75.3543 | 0.18671
2020-02-06 | 12:20 | 3.778 | 3.872 | 3.825 | 407.8568 | 0.689 | 0.71 | 0.6995 | 76.171 | 0.18676
2020-02-06 | 12:20 | 3.937 | 5.744 | 4.8405 | 516.1388 | 0.724 | 1.162 | 0.943 | 102.6866 | 0.19895
2020-02-06 | 12:20 | 6.349 | 7.314 | 6.8315 | 728.4375 | 1.275 | 1.509 | 1.392 | 151.5798 | 0.20809
2020-02-06 | 12:20 | 7.521 | 8.138 | 7.8295 | 834.8535 | 1.475 | 1.652 | 1.5635 | 170.255 | 0.20393
2020-02-06 | 12:20 | 5.955 | 6.929 | 6.442 | 686.9055 | 1.155 | 1.406 | 1.2805 | 139.4381 | 0.20299
2020-02-06 | 12:20 | 7.62 | 8.626 | 8.123 | 866.1492 | 1.557 | 1.759 | 1.658 | 180.5454 | 0.20845
2020-02-06 | 12:20 | 8.661 | 8.705 | 8.683 | 925.8615 | 1.764 | 1.77 | 1.767 | 192.4148 | 0.20782
2020-02-06 | 12:20 | 8.71 | 8.751 | 8.7305 | 930.9264 | 1.771 | 1.779 | 1.775 | 193.286 | 0.20763
2020-02-06 | 12:20 | 8.765 | 8.809 | 8.787 | 936.951 | 1.781 | 1.789 | 1.785 | 194.3749 | 0.20745
2020-02-06 | 12:20 | 8.817 | 8.833 | 8.825 | 941.0029 | 1.79 | 1.795 | 1.7925 | 195.1916 | 0.20743
2020-02-06 | 12:20 | 8.84 | 8.863 | 8.8515 | 943.8286 | 1.793 | 1.798 | 1.7955 | 195.5183 | 0.20715
2020-02-06 | 12:21 | 8.866 | 8.872 | 8.869 | 945.6946 | 1.798 | 1.799 | 1.7985 | 195.845 | 0.20709
2020-02-06 | 12:21 | 8.863 | 8.874 | 8.8685 | 945.6413 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20693
2020-02-06 | 12:21 | 8.853 | 8.86 | 8.8565 | 944.3617 | 1.793 | 1.796 | 1.7945 | 195.4094 | 0.20692
2020-02-06 | 12:21 | 8.846 | 8.852 | 8.849 | 943.562 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20681
2020-02-06 | 12:21 | 8.835 | 8.846 | 8.8405 | 942.6556 | 1.787 | 1.791 | 1.789 | 194.8105 | 0.20666
2020-02-06 | 12:21 | 8.835 | 8.837 | 8.836 | 942.1758 | 1.786 | 1.789 | 1.7875 | 194.6471 | 0.20659
2020-02-06 | 12:21 | 8.832 | 8.837 | 8.8345 | 942.0159 | 1.784 | 1.787 | 1.7855 | 194.4293 | 0.2064
2020-02-06 | 12:21 | 8.817 | 8.827 | 8.822 | 940.683 | 1.781 | 1.785 | 1.783 | 194.1571 | 0.2064
2020-02-06 | 12:21 | 8.794 | 8.813 | 8.8035 | 938.7104 | 1.776 | 1.78 | 1.778 | 193.6126 | 0.20625
2020-02-06 | 12:21 | 8.777 | 8.792 | 8.7845 | 936.6844 | 1.773 | 1.776 | 1.7745 | 193.2315 | 0.20629
2020-02-06 | 12:21 | 8.775 | 8.776 | 8.7755 | 935.7247 | 1.771 | 1.772 | 1.7715 | 192.9048 | 0.20616
2020-02-06 | 12:22 | 8.775 | 8.778 | 8.7765 | 935.8314 | 1.77 | 1.773 | 1.7715 | 192.9048 | 0.20613
2020-02-06 | 12:22 | 8.781 | 8.792 | 8.7865 | 936.8977 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.20596
2020-02-06 | 12:22 | 8.788 | 8.794 | 8.791 | 937.3775 | 1.77 | 1.773 | 1.7715 | 192.9048 | 0.20579
2020-02-06 | 12:22 | 8.788 | 8.791 | 8.7895 | 937.2176 | 1.77 | 1.771 | 1.7705 | 192.7959 | 0.20571
2020-02-06 | 12:22 | 8.792 | 8.795 | 8.7935 | 937.6441 | 1.769 | 1.771 | 1.77 | 192.7415 | 0.20556
2020-02-06 | 12:22 | 8.796 | 8.797 | 8.7965 | 937.964 | 1.769 | 1.771 | 1.77 | 192.7415 | 0.20549
2020-02-06 | 12:22 | 8.791 | 8.798 | 8.7945 | 937.7507 | 1.767 | 1.77 | 1.7685 | 192.5782 | 0.20536
2020-02-06 | 12:22 | 8.773 | 8.788 | 8.7805 | 936.2579 | 1.762 | 1.766 | 1.764 | 192.0881 | 0.20517
2020-02-06 | 12:22 | 8.771 | 8.773 | 8.772 | 935.3515 | 1.762 | 1.763 | 1.7625 | 191.9248 | 0.20519
2020-02-06 | 12:22 | 8.772 | 8.775 | 8.7735 | 935.5115 | 1.761 | 1.762 | 1.7615 | 191.8159 | 0.20504
2020-02-06 | 12:22 | 8.779 | 8.79 | 8.7845 | 936.6844 | 1.761 | 1.764 | 1.7625 | 191.9248 | 0.2049
2020-02-06 | 12:22 | 8.792 | 8.809 | 8.8005 | 938.3905 | 1.764 | 1.767 | 1.7655 | 192.2515 | 0.20487
2020-02-06 | 12:22 | 8.812 | 8.815 | 8.8135 | 939.7767 | 1.766 | 1.767 | 1.7665 | 192.3604 | 0.20469
2020-02-06 | 12:23 | 8.815 | 8.823 | 8.819 | 940.3631 | 1.765 | 1.767 | 1.766 | 192.3059 | 0.2045
2020-02-06 | 12:23 | 8.825 | 8.83 | 8.8275 | 941.2695 | 1.768 | 1.769 | 1.7685 | 192.5782 | 0.20459
2020-02-06 | 12:23 | 8.83 | 8.836 | 8.833 | 941.8559 | 1.768 | 1.769 | 1.7685 | 192.5782 | 0.20447
2020-02-06 | 12:23 | 8.837 | 8.842 | 8.8395 | 942.549 | 1.768 | 1.77 | 1.769 | 192.6326 | 0.20437
2020-02-06 | 12:23 | 8.843 | 8.848 | 8.8455 | 943.1888 | 1.77 | 1.772 | 1.771 | 192.8504 | 0.20447
2020-02-06 | 12:23 | 8.851 | 8.857 | 8.854 | 944.0951 | 1.77 | 1.771 | 1.7705 | 192.7959 | 0.20421
2020-02-06 | 12:23 | 8.858 | 8.865 | 8.8615 | 944.8949 | 1.771 | 1.772 | 1.7715 | 192.9048 | 0.20415
2020-02-06 | 12:23 | 8.866 | 8.872 | 8.869 | 945.6946 | 1.772 | 1.773 | 1.7725 | 193.0137 | 0.2041
2020-02-06 | 12:23 | 8.869 | 8.871 | 8.87 | 945.8012 | 1.772 | 1.773 | 1.7725 | 193.0137 | 0.20407
2020-02-06 | 12:23 | 8.864 | 8.868 | 8.866 | 945.3747 | 1.769 | 1.772 | 1.7705 | 192.7959 | 0.20394
2020-02-06 | 12:23 | 8.856 | 8.863 | 8.8595 | 944.6816 | 1.767 | 1.77 | 1.7685 | 192.5782 | 0.20386
2020-02-06 | 12:24 | 8.853 | 8.858 | 8.8555 | 944.2551 | 1.767 | 1.769 | 1.768 | 192.5237 | 0.20389
2020-02-06 | 12:24 | 8.841 | 8.85 | 8.8455 | 943.1888 | 1.765 | 1.767 | 1.766 | 192.3059 | 0.20389
2020-02-06 | 12:24 | 8.84 | 8.845 | 8.8425 | 942.8689 | 1.764 | 1.766 | 1.765 | 192.197 | 0.20384
2020-02-06 | 12:24 | 8.846 | 8.853 | 8.8495 | 943.6153 | 1.766 | 1.768 | 1.767 | 192.4148 | 0.20391
2020-02-06 | 12:24 | 8.852 | 8.86 | 8.856 | 944.3084 | 1.766 | 1.768 | 1.767 | 192.4148 | 0.20376
2020-02-06 | 12:24 | 8.858 | 8.865 | 8.8615 | 944.8949 | 1.767 | 1.768 | 1.7675 | 192.4693 | 0.20369
2020-02-06 | 12:24 | 8.868 | 8.874 | 8.871 | 945.9078 | 1.768 | 1.77 | 1.769 | 192.6326 | 0.20365
2020-02-06 | 12:24 | 8.871 | 8.874 | 8.8725 | 946.0678 | 1.769 | 1.77 | 1.7695 | 192.6871 | 0.20367
2020-02-06 | 12:24 | 8.874 | 8.88 | 8.877 | 946.5476 | 1.77 | 1.772 | 1.771 | 192.8504 | 0.20374
2020-02-06 | 12:24 | 8.88 | 8.882 | 8.881 | 946.9741 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.20376
2020-02-06 | 12:24 | 8.88 | 8.882 | 8.881 | 946.9741 | 1.771 | 1.772 | 1.7715 | 192.9048 | 0.20371
2020-02-06 | 12:24 | 8.881 | 8.884 | 8.8825 | 947.1341 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.20373
2020-02-06 | 12:24 | 8.888 | 8.892 | 8.89 | 947.9338 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.20356
2020-02-06 | 12:25 | 8.893 | 8.897 | 8.895 | 948.4669 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.20344
2020-02-06 | 12:25 | 8.9 | 8.903 | 8.9015 | 949.16 | 1.773 | 1.775 | 1.774 | 193.1771 | 0.20352
2020-02-06 | 12:25 | 8.899 | 8.903 | 8.901 | 949.1067 | 1.773 | 1.775 | 1.774 | 193.1771 | 0.20354
2020-02-06 | 12:25 | 8.898 | 8.902 | 8.9 | 949.0001 | 1.773 | 1.774 | 1.7735 | 193.1226 | 0.2035
2020-02-06 | 12:25 | 8.896 | 8.898 | 8.897 | 948.6802 | 1.774 | 1.776 | 1.775 | 193.286 | 0.20374
2020-02-06 | 12:25 | 8.898 | 8.904 | 8.901 | 949.1067 | 1.773 | 1.776 | 1.7745 | 193.2315 | 0.20359
2020-02-06 | 12:25 | 8.906 | 8.915 | 8.9105 | 950.1197 | 1.776 | 1.778 | 1.777 | 193.5038 | 0.20366
2020-02-06 | 12:25 | 8.916 | 8.919 | 8.9175 | 950.8661 | 1.779 | 1.78 | 1.7795 | 193.776 | 0.20379
2020-02-06 | 12:25 | 8.918 | 8.921 | 8.9195 | 951.0794 | 1.779 | 1.782 | 1.7805 | 193.8849 | 0.20386
2020-02-06 | 12:25 | 8.923 | 8.926 | 8.9245 | 951.6125 | 1.78 | 1.782 | 1.781 | 193.9393 | 0.2038
2020-02-06 | 12:25 | 8.927 | 8.931 | 8.929 | 952.0923 | 1.782 | 1.783 | 1.7825 | 194.1027 | 0.20387
2020-02-06 | 12:26 | 8.932 | 8.943 | 8.9375 | 952.9987 | 1.783 | 1.785 | 1.784 | 194.266 | 0.20385
2020-02-06 | 12:26 | 8.946 | 8.957 | 8.9515 | 954.4915 | 1.786 | 1.788 | 1.787 | 194.5927 | 0.20387
2020-02-06 | 12:26 | 8.96 | 8.97 | 8.965 | 955.931 | 1.788 | 1.792 | 1.79 | 194.9194 | 0.20391
2020-02-06 | 12:26 | 8.971 | 8.979 | 8.975 | 956.9973 | 1.79 | 1.794 | 1.792 | 195.1372 | 0.20391
2020-02-06 | 12:26 | 8.98 | 8.983 | 8.9815 | 957.6904 | 1.794 | 1.794 | 1.794 | 195.3549 | 0.20399
2020-02-06 | 12:26 | 8.985 | 8.988 | 8.9865 | 958.2235 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20416
2020-02-06 | 12:26 | 8.986 | 8.99 | 8.988 | 958.3835 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20407
2020-02-06 | 12:26 | 8.991 | 8.994 | 8.9925 | 958.8633 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20408
2020-02-06 | 12:26 | 8.995 | 9 | 8.9975 | 959.3964 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20408
2020-02-06 | 12:26 | 9.001 | 9.007 | 9.004 | 960.0895 | 1.799 | 1.802 | 1.8005 | 196.0628 | 0.20421
2020-02-06 | 12:26 | 9.006 | 9.013 | 9.0095 | 960.676 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.20426
2020-02-06 | 12:26 | 9.015 | 9.02 | 9.0175 | 961.529 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20419
2020-02-06 | 12:26 | 9.019 | 9.025 | 9.022 | 962.0088 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20409
2020-02-06 | 12:27 | 9.025 | 9.027 | 9.026 | 962.4354 | 1.803 | 1.804 | 1.8035 | 196.3894 | 0.20405
2020-02-06 | 12:27 | 9.01 | 9.023 | 9.0165 | 961.4224 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.2041
2020-02-06 | 12:27 | 9.002 | 9.008 | 9.005 | 960.1962 | 1.8 | 1.802 | 1.801 | 196.1172 | 0.20425
2020-02-06 | 12:27 | 9.008 | 9.027 | 9.0175 | 961.529 | 1.802 | 1.805 | 1.8035 | 196.3894 | 0.20425
2020-02-06 | 12:27 | 9.032 | 9.048 | 9.04 | 963.9282 | 1.805 | 1.807 | 1.806 | 196.6617 | 0.20402
2020-02-06 | 12:27 | 9.048 | 9.05 | 9.049 | 964.8878 | 1.807 | 1.809 | 1.808 | 196.8795 | 0.20404
2020-02-06 | 12:27 | 9.034 | 9.045 | 9.0395 | 963.8749 | 1.806 | 1.807 | 1.8065 | 196.7161 | 0.20409
2020-02-06 | 12:27 | 9.022 | 9.033 | 9.0275 | 962.5953 | 1.803 | 1.805 | 1.804 | 196.4439 | 0.20408
2020-02-06 | 12:27 | 8.998 | 9.018 | 9.008 | 960.516 | 1.8 | 1.804 | 1.802 | 196.2261 | 0.20429
2020-02-06 | 12:27 | 8.991 | 9 | 8.9955 | 959.1832 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20412
2020-02-06 | 12:27 | 8.977 | 8.991 | 8.984 | 957.9569 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.20416
2020-02-06 | 12:28 | 8.961 | 8.971 | 8.966 | 956.0376 | 1.792 | 1.795 | 1.7935 | 195.3005 | 0.20428
2020-02-06 | 12:28 | 8.906 | 8.957 | 8.9315 | 952.3589 | 1.781 | 1.79 | 1.7855 | 194.4293 | 0.20416
2020-02-06 | 12:28 | 8.884 | 8.923 | 8.9035 | 949.3733 | 1.778 | 1.786 | 1.782 | 194.0482 | 0.2044
2020-02-06 | 12:28 | 8.942 | 8.959 | 8.9505 | 954.3849 | 1.789 | 1.792 | 1.7905 | 194.9738 | 0.20429
2020-02-06 | 12:28 | 8.963 | 8.969 | 8.966 | 956.0376 | 1.791 | 1.794 | 1.7925 | 195.1916 | 0.20417
2020-02-06 | 12:28 | 8.965 | 8.968 | 8.9665 | 956.0909 | 1.792 | 1.794 | 1.793 | 195.246 | 0.20421
2020-02-06 | 12:28 | 8.952 | 8.965 | 8.9585 | 955.2379 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20417
2020-02-06 | 12:28 | 8.941 | 8.95 | 8.9455 | 953.8517 | 1.786 | 1.788 | 1.787 | 194.5927 | 0.20401
2020-02-06 | 12:28 | 8.929 | 8.935 | 8.932 | 952.4122 | 1.786 | 1.788 | 1.787 | 194.5927 | 0.20432
2020-02-06 | 12:28 | 8.928 | 8.944 | 8.936 | 952.8387 | 1.785 | 1.788 | 1.7865 | 194.5382 | 0.20417
2020-02-06 | 12:28 | 8.927 | 8.942 | 8.9345 | 952.6788 | 1.785 | 1.787 | 1.786 | 194.4838 | 0.20414
2020-02-06 | 12:28 | 8.923 | 8.926 | 8.9245 | 951.6125 | 1.784 | 1.786 | 1.785 | 194.3749 | 0.20426
2020-02-06 | 12:28 | 8.926 | 8.946 | 8.936 | 952.8387 | 1.785 | 1.79 | 1.7875 | 194.6471 | 0.20428
2020-02-06 | 12:29 | 8.951 | 8.954 | 8.9525 | 954.5981 | 1.789 | 1.791 | 1.79 | 194.9194 | 0.20419
2020-02-06 | 12:29 | 8.954 | 8.965 | 8.9595 | 955.3445 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20414
2020-02-06 | 12:29 | 8.964 | 8.971 | 8.9675 | 956.1976 | 1.792 | 1.793 | 1.7925 | 195.1916 | 0.20413
2020-02-06 | 12:29 | 8.973 | 8.975 | 8.974 | 956.8906 | 1.793 | 1.793 | 1.793 | 195.246 | 0.20404
2020-02-06 | 12:29 | 8.974 | 8.977 | 8.9755 | 957.0506 | 1.793 | 1.794 | 1.7935 | 195.3005 | 0.20406
2020-02-06 | 12:29 | 8.961 | 8.974 | 8.9675 | 956.1976 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20408
2020-02-06 | 12:29 | 8.946 | 8.956 | 8.951 | 954.4382 | 1.788 | 1.79 | 1.789 | 194.8105 | 0.20411
2020-02-06 | 12:29 | 8.937 | 8.944 | 8.9405 | 953.3186 | 1.786 | 1.787 | 1.7865 | 194.5382 | 0.20406
2020-02-06 | 12:29 | 8.931 | 8.937 | 8.934 | 952.6255 | 1.784 | 1.787 | 1.7855 | 194.4293 | 0.2041
2020-02-06 | 12:29 | 8.934 | 8.936 | 8.935 | 952.7321 | 1.783 | 1.785 | 1.784 | 194.266 | 0.2039
2020-02-06 | 12:29 | 8.939 | 8.95 | 8.9445 | 953.7451 | 1.785 | 1.787 | 1.786 | 194.4838 | 0.20392
2020-02-06 | 12:30 | 8.952 | 8.954 | 8.953 | 954.6514 | 1.786 | 1.787 | 1.7865 | 194.5382 | 0.20378
2020-02-06 | 12:30 | 8.953 | 8.96 | 8.9565 | 955.0246 | 1.786 | 1.79 | 1.788 | 194.7016 | 0.20387
2020-02-06 | 12:30 | 8.962 | 8.966 | 8.964 | 955.8244 | 1.789 | 1.79 | 1.7895 | 194.8649 | 0.20387
2020-02-06 | 12:30 | 8.966 | 8.973 | 8.9695 | 956.4108 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20392
2020-02-06 | 12:30 | 8.951 | 8.967 | 8.959 | 955.2912 | 1.787 | 1.791 | 1.789 | 194.8105 | 0.20393
2020-02-06 | 12:30 | 8.942 | 8.957 | 8.9495 | 954.2782 | 1.785 | 1.79 | 1.7875 | 194.6471 | 0.20397
2020-02-06 | 12:30 | 8.924 | 8.937 | 8.9305 | 952.2523 | 1.783 | 1.785 | 1.784 | 194.266 | 0.20401
2020-02-06 | 12:30 | 8.931 | 8.952 | 8.9415 | 953.4252 | 1.785 | 1.788 | 1.7865 | 194.5382 | 0.20404
2020-02-06 | 12:30 | 8.954 | 8.957 | 8.9555 | 954.918 | 1.788 | 1.79 | 1.789 | 194.8105 | 0.20401
2020-02-06 | 12:30 | 8.942 | 8.949 | 8.9455 | 953.8517 | 1.787 | 1.789 | 1.788 | 194.7016 | 0.20412
2020-02-06 | 12:30 | 8.944 | 8.956 | 8.95 | 954.3315 | 1.787 | 1.79 | 1.7885 | 194.756 | 0.20408
2020-02-06 | 12:30 | 8.96 | 8.966 | 8.963 | 955.7177 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20406
2020-02-06 | 12:30 | 8.964 | 8.971 | 8.9675 | 956.1976 | 1.791 | 1.793 | 1.792 | 195.1372 | 0.20408
2020-02-06 | 12:31 | 8.973 | 8.981 | 8.977 | 957.2105 | 1.793 | 1.795 | 1.794 | 195.3549 | 0.20409
2020-02-06 | 12:31 | 8.98 | 8.984 | 8.982 | 957.7437 | 1.794 | 1.796 | 1.795 | 195.4638 | 0.20409
2020-02-06 | 12:31 | 8.963 | 8.979 | 8.971 | 956.5708 | 1.792 | 1.796 | 1.794 | 195.3549 | 0.20422
2020-02-06 | 12:31 | 8.944 | 8.959 | 8.9515 | 954.4915 | 1.79 | 1.793 | 1.7915 | 195.0827 | 0.20438
2020-02-06 | 12:31 | 8.939 | 8.943 | 8.941 | 953.3719 | 1.79 | 1.791 | 1.7905 | 194.9738 | 0.20451
2020-02-06 | 12:31 | 8.926 | 8.938 | 8.932 | 952.4122 | 1.788 | 1.79 | 1.789 | 194.8105 | 0.20454
2020-02-06 | 12:31 | 8.927 | 8.942 | 8.9345 | 952.6788 | 1.789 | 1.791 | 1.79 | 194.9194 | 0.2046
2020-02-06 | 12:31 | 8.941 | 8.944 | 8.9425 | 953.5318 | 1.79 | 1.792 | 1.791 | 195.0283 | 0.20453
2020-02-06 | 12:31 | 8.948 | 8.958 | 8.953 | 954.6514 | 1.791 | 1.795 | 1.793 | 195.246 | 0.20452
2020-02-06 | 12:31 | 8.961 | 8.966 | 8.9635 | 955.771 | 1.795 | 1.798 | 1.7965 | 195.6272 | 0.20468
2020-02-06 | 12:31 | 8.967 | 8.975 | 8.971 | 956.5708 | 1.797 | 1.799 | 1.798 | 195.7905 | 0.20468
2020-02-06 | 12:32 | 8.977 | 8.991 | 8.984 | 957.9569 | 1.8 | 1.802 | 1.801 | 196.1172 | 0.20472
2020-02-06 | 12:32 | 8.992 | 8.999 | 8.9955 | 959.1832 | 1.802 | 1.804 | 1.803 | 196.335 | 0.20469
2020-02-06 | 12:32 | 9.001 | 9.005 | 9.003 | 959.9829 | 1.804 | 1.806 | 1.805 | 196.5528 | 0.20475
2020-02-06 | 12:32 | 9.005 | 9.009 | 9.007 | 960.4094 | 1.805 | 1.808 | 1.8065 | 196.7161 | 0.20483
2020-02-06 | 12:32 | 9.011 | 9.014 | 9.0125 | 960.9959 | 1.807 | 1.809 | 1.808 | 196.8795 | 0.20487
2020-02-06 | 12:32 | 9.009 | 9.013 | 9.011 | 960.8359 | 1.807 | 1.809 | 1.808 | 196.8795 | 0.2049
2020-02-06 | 12:32 | 9.008 | 9.013 | 9.0105 | 960.7826 | 1.807 | 1.808 | 1.8075 | 196.825 | 0.20486
2020-02-06 | 12:32 | 9.015 | 9.017 | 9.016 | 961.3691 | 1.809 | 1.81 | 1.8095 | 197.0428 | 0.20496
2020-02-06 | 12:32 | 9.012 | 9.016 | 9.014 | 961.1558 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.20495
2020-02-06 | 12:32 | 8.997 | 9.01 | 9.0035 | 960.0362 | 1.807 | 1.81 | 1.8085 | 196.9339 | 0.20513
2020-02-06 | 12:32 | 8.971 | 8.993 | 8.982 | 957.7437 | 1.802 | 1.807 | 1.8045 | 196.4983 | 0.20517
2020-02-06 | 12:32 | 8.94 | 8.964 | 8.952 | 954.5448 | 1.798 | 1.803 | 1.8005 | 196.0628 | 0.2054
2020-02-06 | 12:32 | 8.92 | 8.932 | 8.926 | 951.7724 | 1.794 | 1.797 | 1.7955 | 195.5183 | 0.20543
2020-02-06 | 12:33 | 8.925 | 8.935 | 8.93 | 952.199 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20545
2020-02-06 | 12:33 | 8.932 | 8.938 | 8.935 | 952.7321 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20539
2020-02-06 | 12:33 | 8.922 | 8.935 | 8.9285 | 952.039 | 1.794 | 1.797 | 1.7955 | 195.5183 | 0.20537
2020-02-06 | 12:33 | 8.894 | 8.92 | 8.907 | 949.7465 | 1.788 | 1.793 | 1.7905 | 194.9738 | 0.20529
2020-02-06 | 12:33 | 8.867 | 8.882 | 8.8745 | 946.281 | 1.784 | 1.788 | 1.786 | 194.4838 | 0.20552
2020-02-06 | 12:33 | 8.843 | 8.862 | 8.8525 | 943.9352 | 1.781 | 1.785 | 1.783 | 194.1571 | 0.20569
2020-02-06 | 12:33 | 8.828 | 8.838 | 8.833 | 941.8559 | 1.778 | 1.782 | 1.78 | 193.8304 | 0.2058
2020-02-06 | 12:33 | 8.834 | 8.844 | 8.839 | 942.4957 | 1.78 | 1.782 | 1.781 | 193.9393 | 0.20577
2020-02-06 | 12:33 | 8.849 | 8.855 | 8.852 | 943.8819 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.2057
2020-02-06 | 12:33 | 8.845 | 8.852 | 8.8485 | 943.5087 | 1.782 | 1.784 | 1.783 | 194.1571 | 0.20578
2020-02-06 | 12:33 | 8.839 | 8.846 | 8.8425 | 942.8689 | 1.78 | 1.782 | 1.781 | 193.9393 | 0.20569
2020-02-06 | 12:34 | 8.832 | 8.837 | 8.8345 | 942.0159 | 1.778 | 1.781 | 1.7795 | 193.776 | 0.2057
2020-02-06 | 12:34 | 8.825 | 8.833 | 8.829 | 941.4294 | 1.777 | 1.779 | 1.778 | 193.6126 | 0.20566
2020-02-06 | 12:34 | 8.824 | 8.826 | 8.825 | 941.0029 | 1.777 | 1.779 | 1.778 | 193.6126 | 0.20575
2020-02-06 | 12:34 | 8.812 | 8.823 | 8.8175 | 940.2032 | 1.774 | 1.776 | 1.775 | 193.286 | 0.20558
2020-02-06 | 12:34 | 8.784 | 8.806 | 8.795 | 937.804 | 1.768 | 1.773 | 1.7705 | 192.7959 | 0.20558
2020-02-06 | 12:34 | 8.768 | 8.781 | 8.7745 | 935.6181 | 1.766 | 1.767 | 1.7665 | 192.3604 | 0.2056
2020-02-06 | 12:34 | 8.745 | 8.768 | 8.7565 | 933.6988 | 1.762 | 1.766 | 1.764 | 192.0881 | 0.20573
2020-02-06 | 12:34 | 8.724 | 8.742 | 8.733 | 931.193 | 1.758 | 1.76 | 1.759 | 191.5437 | 0.2057
2020-02-06 | 12:34 | 8.711 | 8.72 | 8.7155 | 929.327 | 1.756 | 1.758 | 1.757 | 191.3259 | 0.20588
2020-02-06 | 12:34 | 8.706 | 8.71 | 8.708 | 928.5273 | 1.757 | 1.758 | 1.7575 | 191.3803 | 0.20611
2020-02-06 | 12:34 | 8.695 | 8.701 | 8.698 | 927.461 | 1.756 | 1.757 | 1.7565 | 191.2714 | 0.20623
2020-02-06 | 12:34 | 8.686 | 8.691 | 8.6885 | 926.448 | 1.756 | 1.757 | 1.7565 | 191.2714 | 0.20646
2020-02-06 | 12:34 | 8.675 | 8.687 | 8.681 | 925.6483 | 1.756 | 1.756 | 1.756 | 191.217 | 0.20658
2020-02-06 | 12:35 | 8.668 | 8.673 | 8.6705 | 924.5287 | 1.757 | 1.757 | 1.757 | 191.3259 | 0.20694
2020-02-06 | 12:35 | 8.676 | 8.688 | 8.682 | 925.7549 | 1.757 | 1.76 | 1.7585 | 191.4892 | 0.20685
2020-02-06 | 12:35 | 8.689 | 8.693 | 8.691 | 926.7146 | 1.759 | 1.761 | 1.76 | 191.6526 | 0.20681
2020-02-06 | 12:35 | 8.559 | 8.685 | 8.622 | 919.3572 | 1.724 | 1.759 | 1.7415 | 189.638 | 0.20627
2020-02-06 | 12:35 | 8.102 | 8.458 | 8.28 | 882.89 | 1.611 | 1.698 | 1.6545 | 180.1643 | 0.20406
2020-02-06 | 12:35 | 6.425 | 7.575 | 7 | 746.4046 | 1.265 | 1.471 | 1.368 | 148.9663 | 0.19958
2020-02-06 | 12:35 | 7.019 | 7.508 | 7.2635 | 774.5014 | 1.4 | 1.52 | 1.46 | 158.9845 | 0.20527
2020-02-06 | 12:35 | 5.25 | 7.195 | 6.2225 | 663.5003 | 0.972 | 1.418 | 1.195 | 130.1277 | 0.19612
2020-02-06 | 12:35 | 4.759 | 5.552 | 5.1555 | 549.727 | 0.891 | 1.082 | 0.9865 | 107.4234 | 0.19541
2020-02-06 | 12:35 | 5.915 | 6.502 | 6.2085 | 662.0075 | 1.166 | 1.283 | 1.2245 | 133.3401 | 0.20142
2020-02-06 | 12:35 | 4.959 | 6.04 | 5.4995 | 586.4074 | 0.948 | 1.17 | 1.059 | 115.3182 | 0.19665
2020-02-06 | 12:36 | 4.818 | 5.365 | 5.0915 | 542.9027 | 0.927 | 1.048 | 0.9875 | 107.5323 | 0.19807
2020-02-06 | 12:36 | 3.489 | 5.171 | 4.33 | 461.7045 | 0.615 | 0.973 | 0.794 | 86.4614 | 0.18727
2020-02-06 | 12:36 | 3.311 | 3.402 | 3.3565 | 357.901 | 0.589 | 0.601 | 0.595 | 64.7916 | 0.18103
2020-02-06 | 12:36 | 3.276 | 3.396 | 3.336 | 355.7151 | 0.584 | 0.61 | 0.597 | 65.0094 | 0.18276
2020-02-06 | 12:36 | 3.214 | 3.252 | 3.233 | 344.7323 | 0.573 | 0.58 | 0.5765 | 62.7771 | 0.1821
2020-02-06 | 12:36 | 3.19 | 3.207 | 3.1985 | 341.0536 | 0.569 | 0.572 | 0.5705 | 62.1237 | 0.18215
2020-02-06 | 12:36 | 3.165 | 3.185 | 3.175 | 338.5478 | 0.566 | 0.569 | 0.5675 | 61.7971 | 0.18254
2020-02-06 | 12:36 | 3.146 | 3.162 | 3.154 | 336.3086 | 0.562 | 0.565 | 0.5635 | 61.3615 | 0.18246
2020-02-06 | 12:36 | 3.127 | 3.142 | 3.1345 | 334.2293 | 0.557 | 0.562 | 0.5595 | 60.9259 | 0.18229
2020-02-06 | 12:36 | 3.113 | 3.125 | 3.119 | 332.5765 | 0.555 | 0.558 | 0.5565 | 60.5992 | 0.18221
2020-02-06 | 12:36 | 3.108 | 3.112 | 3.11 | 331.6169 | 0.553 | 0.555 | 0.554 | 60.327 | 0.18192
2020-02-06 | 12:36 | 3.107 | 3.111 | 3.109 | 331.5103 | 0.554 | 0.555 | 0.5545 | 60.3814 | 0.18214
2020-02-06 | 12:36 | 3.112 | 3.121 | 3.1165 | 332.31 | 0.555 | 0.557 | 0.556 | 60.5448 | 0.18219
2020-02-06 | 12:37 | 3.124 | 3.139 | 3.1315 | 333.9094 | 0.559 | 0.562 | 0.5605 | 61.0348 | 0.18279
2020-02-06 | 12:37 | 3.144 | 3.164 | 3.154 | 336.3086 | 0.563 | 0.568 | 0.5655 | 61.5793 | 0.1831
2020-02-06 | 12:37 | 3.17 | 3.198 | 3.184 | 339.5074 | 0.567 | 0.575 | 0.571 | 62.1782 | 0.18314
2020-02-06 | 12:37 | 3.204 | 3.236 | 3.22 | 343.3461 | 0.575 | 0.583 | 0.579 | 63.0493 | 0.18363
2020-02-06 | 12:37 | 3.247 | 3.306 | 3.2765 | 349.3706 | 0.585 | 0.595 | 0.59 | 64.2472 | 0.18389
2020-02-06 | 12:37 | 3.324 | 3.398 | 3.361 | 358.3808 | 0.6 | 0.613 | 0.6065 | 66.0439 | 0.18428
2020-02-06 | 12:37 | 3.414 | 3.483 | 3.4485 | 367.7109 | 0.617 | 0.63 | 0.6235 | 67.8951 | 0.18464
2020-02-06 | 12:37 | 3.5 | 3.562 | 3.531 | 376.5078 | 0.634 | 0.644 | 0.639 | 69.5829 | 0.18481
2020-02-06 | 12:37 | 3.577 | 3.623 | 3.6 | 383.8652 | 0.647 | 0.655 | 0.651 | 70.8897 | 0.18467
2020-02-06 | 12:37 | 3.633 | 3.68 | 3.6565 | 389.8898 | 0.657 | 0.666 | 0.6615 | 72.0331 | 0.18475
2020-02-06 | 12:37 | 3.695 | 3.741 | 3.718 | 396.4475 | 0.667 | 0.675 | 0.671 | 73.0675 | 0.18431
2020-02-06 | 12:38 | 3.749 | 3.797 | 3.773 | 402.3121 | 0.677 | 0.686 | 0.6815 | 74.2109 | 0.18446
2020-02-06 | 12:38 | 3.814 | 4.22 | 4.017 | 428.3296 | 0.69 | 0.79 | 0.74 | 80.5812 | 0.18813
2020-02-06 | 12:38 | 4.456 | 6.269 | 5.3625 | 571.7992 | 0.851 | 1.273 | 1.062 | 115.6449 | 0.20225
2020-02-06 | 12:38 | 6.834 | 8.365 | 7.5995 | 810.3288 | 1.398 | 1.713 | 1.5555 | 169.3838 | 0.20903
2020-02-06 | 12:38 | 8.493 | 8.686 | 8.5895 | 915.8917 | 1.735 | 1.766 | 1.7505 | 190.6181 | 0.20812
2020-02-06 | 12:38 | 8.709 | 8.779 | 8.744 | 932.3659 | 1.768 | 1.778 | 1.773 | 193.0682 | 0.20707
2020-02-06 | 12:38 | 8.786 | 8.815 | 8.8005 | 938.3905 | 1.779 | 1.783 | 1.781 | 193.9393 | 0.20667
2020-02-06 | 12:38 | 8.82 | 8.85 | 8.835 | 942.0692 | 1.785 | 1.789 | 1.787 | 194.5927 | 0.20656
2020-02-06 | 12:38 | 8.859 | 8.893 | 8.876 | 946.441 | 1.789 | 1.795 | 1.792 | 195.1372 | 0.20618
2020-02-06 | 12:38 | 8.901 | 8.905 | 8.903 | 949.32 | 1.796 | 1.797 | 1.7965 | 195.6272 | 0.20607
2020-02-06 | 12:38 | 8.906 | 8.912 | 8.909 | 949.9597 | 1.796 | 1.798 | 1.797 | 195.6816 | 0.20599
2020-02-06 | 12:38 | 8.906 | 8.913 | 8.9095 | 950.0131 | 1.794 | 1.796 | 1.795 | 195.4638 | 0.20575
2020-02-06 | 12:39 | 8.918 | 8.928 | 8.923 | 951.4526 | 1.795 | 1.796 | 1.7955 | 195.5183 | 0.20549
2020-02-06 | 12:39 | 8.902 | 8.926 | 8.914 | 950.4929 | 1.788 | 1.795 | 1.7915 | 195.0827 | 0.20524
2020-02-06 | 12:39 | 8.824 | 8.893 | 8.8585 | 944.575 | 1.768 | 1.786 | 1.777 | 193.5038 | 0.20486
2020-02-06 | 12:39 | 7.847 | 8.731 | 8.289 | 883.8496 | 1.562 | 1.744 | 1.653 | 180.001 | 0.20366
2020-02-06 | 12:39 | 8.016 | 8.654 | 8.335 | 888.7546 | 1.613 | 1.746 | 1.6795 | 182.8866 | 0.20578
2020-02-06 | 12:39 | 8.706 | 8.766 | 8.736 | 931.5129 | 1.754 | 1.764 | 1.759 | 191.5437 | 0.20563
2020-02-06 | 12:39 | 8.764 | 8.773 | 8.7685 | 934.9783 | 1.762 | 1.764 | 1.763 | 191.9792 | 0.20533
2020-02-06 | 12:39 | 8.773 | 8.781 | 8.777 | 935.8847 | 1.764 | 1.766 | 1.765 | 192.197 | 0.20536
2020-02-06 | 12:39 | 8.788 | 8.837 | 8.8125 | 939.67 | 1.768 | 1.776 | 1.772 | 192.9593 | 0.20535
2020-02-06 | 12:39 | 8.731 | 8.859 | 8.795 | 937.804 | 1.751 | 1.78 | 1.7655 | 192.2515 | 0.205
2020-02-06 | 12:39 | 8.744 | 8.854 | 8.799 | 938.2305 | 1.749 | 1.776 | 1.7625 | 191.9248 | 0.20456
2020-02-06 | 12:39 | 8.033 | 8.677 | 8.355 | 890.8872 | 1.567 | 1.735 | 1.651 | 179.7832 | 0.2018
2020-02-06 | 12:39 | 7.155 | 7.67 | 7.4125 | 790.3891 | 1.396 | 1.517 | 1.4565 | 158.6034 | 0.20066
2020-02-06 | 12:40 | 7.227 | 7.787 | 7.507 | 800.4656 | 1.412 | 1.544 | 1.478 | 160.9446 | 0.20106
2020-02-06 | 12:40 | 7.752 | 8.03 | 7.891 | 841.4112 | 1.528 | 1.598 | 1.563 | 170.2005 | 0.20228
2020-02-06 | 12:40 | 7.233 | 7.388 | 7.3105 | 779.5129 | 1.408 | 1.452 | 1.43 | 155.7177 | 0.19976
2020-02-06 | 12:40 | 7.368 | 7.674 | 7.521 | 801.9584 | 1.448 | 1.522 | 1.485 | 161.7069 | 0.20164
2020-02-06 | 12:40 | 7.463 | 7.889 | 7.676 | 818.4859 | 1.481 | 1.583 | 1.532 | 166.8248 | 0.20382
2020-02-06 | 12:40 | 8.187 | 8.447 | 8.317 | 886.8352 | 1.645 | 1.697 | 1.671 | 181.961 | 0.20518
2020-02-06 | 12:40 | 7.912 | 8.579 | 8.2455 | 879.2113 | 1.559 | 1.715 | 1.637 | 178.2587 | 0.20275
2020-02-06 | 12:40 | 8.043 | 8.628 | 8.3355 | 888.8079 | 1.608 | 1.734 | 1.671 | 181.961 | 0.20472
2020-02-06 | 12:40 | 8.643 | 8.674 | 8.6585 | 923.2491 | 1.737 | 1.741 | 1.739 | 189.3658 | 0.20511
2020-02-06 | 12:40 | 8.678 | 8.685 | 8.6815 | 925.7016 | 1.74 | 1.742 | 1.741 | 189.5836 | 0.2048
2020-02-06 | 12:40 | 8.657 | 8.674 | 8.6655 | 923.9955 | 1.737 | 1.74 | 1.7385 | 189.3114 | 0.20488
2020-02-06 | 12:41 | 8.653 | 8.663 | 8.658 | 923.1958 | 1.736 | 1.739 | 1.7375 | 189.2025 | 0.20494
2020-02-06 | 12:41 | 8.668 | 8.703 | 8.6855 | 926.1281 | 1.739 | 1.744 | 1.7415 | 189.638 | 0.20476
2020-02-06 | 12:41 | 8.71 | 8.726 | 8.718 | 929.5936 | 1.746 | 1.748 | 1.747 | 190.2369 | 0.20465
2020-02-06 | 12:41 | 8.717 | 8.725 | 8.721 | 929.9135 | 1.747 | 1.748 | 1.7475 | 190.2914 | 0.20463
2020-02-06 | 12:41 | 8.703 | 8.714 | 8.7085 | 928.5806 | 1.745 | 1.747 | 1.746 | 190.1281 | 0.20475
2020-02-06 | 12:41 | 8.702 | 8.705 | 8.7035 | 928.0474 | 1.745 | 1.746 | 1.7455 | 190.0736 | 0.20481
2020-02-06 | 12:41 | 8.706 | 8.71 | 8.708 | 928.5273 | 1.745 | 1.746 | 1.7455 | 190.0736 | 0.2047
2020-02-06 | 12:41 | 8.694 | 8.709 | 8.7015 | 927.8342 | 1.743 | 1.745 | 1.744 | 189.9103 | 0.20468
2020-02-06 | 12:41 | 8.688 | 8.69 | 8.689 | 926.5013 | 1.74 | 1.742 | 1.741 | 189.5836 | 0.20462
2020-02-06 | 12:41 | 8.683 | 8.689 | 8.686 | 926.1814 | 1.741 | 1.743 | 1.742 | 189.6925 | 0.20481
2020-02-06 | 12:41 | 8.667 | 8.683 | 8.675 | 925.0085 | 1.738 | 1.742 | 1.74 | 189.4747 | 0.20484
2020-02-06 | 12:41 | 8.657 | 8.661 | 8.659 | 923.3024 | 1.736 | 1.737 | 1.7365 | 189.0936 | 0.2048
2020-02-06 | 12:41 | 8.639 | 8.654 | 8.6465 | 921.9696 | 1.734 | 1.736 | 1.735 | 188.9302 | 0.20492
2020-02-06 | 12:42 | 8.608 | 8.629 | 8.6185 | 918.984 | 1.728 | 1.731 | 1.7295 | 188.3313 | 0.20493
2020-02-06 | 12:42 | 8.619 | 8.636 | 8.6275 | 919.9436 | 1.73 | 1.734 | 1.732 | 188.6035 | 0.20502
2020-02-06 | 12:42 | 8.62 | 8.634 | 8.627 | 919.8903 | 1.731 | 1.732 | 1.7315 | 188.5491 | 0.20497
2020-02-06 | 12:42 | 8.615 | 8.621 | 8.618 | 918.9306 | 1.729 | 1.731 | 1.73 | 188.3858 | 0.20501
2020-02-06 | 12:42 | 8.598 | 8.61 | 8.604 | 917.4378 | 1.726 | 1.729 | 1.7275 | 188.1135 | 0.20504
2020-02-06 | 12:42 | 8.598 | 8.605 | 8.6015 | 917.1713 | 1.727 | 1.728 | 1.7275 | 188.1135 | 0.2051
2020-02-06 | 12:42 | 8.607 | 8.622 | 8.6145 | 918.5574 | 1.729 | 1.731 | 1.73 | 188.3858 | 0.20509
2020-02-06 | 12:42 | 8.629 | 8.642 | 8.6355 | 920.7967 | 1.732 | 1.734 | 1.733 | 188.7124 | 0.20494
2020-02-06 | 12:42 | 8.64 | 8.642 | 8.641 | 921.3831 | 1.733 | 1.734 | 1.7335 | 188.7669 | 0.20487
2020-02-06 | 12:42 | 8.641 | 8.648 | 8.6445 | 921.7563 | 1.733 | 1.736 | 1.7345 | 188.8758 | 0.20491
2020-02-06 | 12:42 | 8.645 | 8.647 | 8.646 | 921.9163 | 1.736 | 1.737 | 1.7365 | 189.0936 | 0.20511
2020-02-06 | 12:43 | 8.642 | 8.648 | 8.645 | 921.8096 | 1.734 | 1.736 | 1.735 | 188.9302 | 0.20496
2020-02-06 | 12:43 | 8.607 | 8.638 | 8.6225 | 919.4105 | 1.729 | 1.733 | 1.731 | 188.4947 | 0.20502
2020-02-06 | 12:43 | 8.614 | 8.626 | 8.62 | 919.1439 | 1.731 | 1.733 | 1.732 | 188.6035 | 0.20519
2020-02-06 | 12:43 | 8.628 | 8.637 | 8.6325 | 920.4768 | 1.734 | 1.737 | 1.7355 | 188.9847 | 0.20531
2020-02-06 | 12:43 | 8.642 | 8.669 | 8.6555 | 922.9292 | 1.738 | 1.742 | 1.74 | 189.4747 | 0.2053
2020-02-06 | 12:43 | 8.666 | 8.668 | 8.667 | 924.1555 | 1.741 | 1.742 | 1.7415 | 189.638 | 0.2052
2020-02-06 | 12:43 | 8.652 | 8.663 | 8.6575 | 923.1425 | 1.738 | 1.741 | 1.7395 | 189.4202 | 0.20519
2020-02-06 | 12:43 | 8.653 | 8.66 | 8.6565 | 923.0359 | 1.739 | 1.741 | 1.74 | 189.4747 | 0.20527
2020-02-06 | 12:43 | 8.661 | 8.666 | 8.6635 | 923.7823 | 1.741 | 1.743 | 1.742 | 189.6925 | 0.20534
2020-02-06 | 12:43 | 8.666 | 8.675 | 8.6705 | 924.5287 | 1.742 | 1.745 | 1.7435 | 189.8558 | 0.20535
2020-02-06 | 12:43 | 8.677 | 8.683 | 8.68 | 925.5417 | 1.746 | 1.747 | 1.7465 | 190.1825 | 0.20548
2020-02-06 | 12:43 | 8.682 | 8.693 | 8.6875 | 926.3414 | 1.747 | 1.749 | 1.748 | 190.3458 | 0.20548
2020-02-06 | 12:43 | 8.672 | 8.695 | 8.6835 | 925.9149 | 1.745 | 1.749 | 1.747 | 190.2369 | 0.20546
2020-02-06 | 12:44 | 8.668 | 8.689 | 8.6785 | 925.3817 | 1.745 | 1.748 | 1.7465 | 190.1825 | 0.20552
2020-02-06 | 12:44 | 8.685 | 8.688 | 8.6865 | 926.2347 | 1.746 | 1.748 | 1.747 | 190.2369 | 0.20539
2020-02-06 | 12:44 | 8.683 | 8.689 | 8.686 | 926.1814 | 1.746 | 1.748 | 1.747 | 190.2369 | 0.2054
2020-02-06 | 12:44 | 8.657 | 8.685 | 8.671 | 924.582 | 1.74 | 1.746 | 1.743 | 189.8014 | 0.20528
2020-02-06 | 12:44 | 8.229 | 8.631 | 8.43 | 898.8843 | 1.641 | 1.735 | 1.688 | 183.8122 | 0.20449
2020-02-06 | 12:44 | 8.159 | 8.375 | 8.267 | 881.5038 | 1.632 | 1.686 | 1.659 | 180.6543 | 0.20494
2020-02-06 | 12:44 | 8.467 | 8.698 | 8.5825 | 915.1453 | 1.706 | 1.749 | 1.7275 | 188.1135 | 0.20556
2020-02-06 | 12:44 | 8.709 | 8.717 | 8.713 | 929.0604 | 1.75 | 1.751 | 1.7505 | 190.6181 | 0.20517
2020-02-06 | 12:44 | 8.662 | 8.702 | 8.682 | 925.7549 | 1.74 | 1.748 | 1.744 | 189.9103 | 0.20514
2020-02-06 | 12:44 | 8.65 | 8.665 | 8.6575 | 923.1425 | 1.738 | 1.742 | 1.74 | 189.4747 | 0.20525
2020-02-06 | 12:44 | 8.639 | 8.669 | 8.654 | 922.7693 | 1.738 | 1.742 | 1.74 | 189.4747 | 0.20533
2020-02-06 | 12:45 | 8.642 | 8.651 | 8.6465 | 921.9696 | 1.738 | 1.74 | 1.739 | 189.3658 | 0.20539
2020-02-06 | 12:45 | 8.593 | 8.646 | 8.6195 | 919.0906 | 1.73 | 1.739 | 1.7345 | 188.8758 | 0.2055
2020-02-06 | 12:45 | 8.591 | 8.678 | 8.6345 | 920.69 | 1.73 | 1.745 | 1.7375 | 189.2025 | 0.2055
2020-02-06 | 12:45 | 8.683 | 8.693 | 8.688 | 926.3947 | 1.745 | 1.748 | 1.7465 | 190.1825 | 0.20529
2020-02-06 | 12:45 | 8.657 | 8.679 | 8.668 | 924.2621 | 1.741 | 1.744 | 1.7425 | 189.7469 | 0.2053
2020-02-06 | 12:45 | 8.654 | 8.657 | 8.6555 | 922.9292 | 1.739 | 1.741 | 1.74 | 189.4747 | 0.2053
2020-02-06 | 12:45 | 8.641 | 8.653 | 8.647 | 922.0229 | 1.737 | 1.74 | 1.7385 | 189.3114 | 0.20532
2020-02-06 | 12:45 | 8.648 | 8.651 | 8.6495 | 922.2895 | 1.736 | 1.739 | 1.7375 | 189.2025 | 0.20514
2020-02-06 | 12:45 | 8.636 | 8.645 | 8.6405 | 921.3298 | 1.733 | 1.736 | 1.7345 | 188.8758 | 0.205
2020-02-06 | 12:45 | 8.589 | 8.628 | 8.6085 | 917.9177 | 1.725 | 1.732 | 1.7285 | 188.2224 | 0.20505
2020-02-06 | 12:45 | 8.572 | 8.582 | 8.577 | 914.5588 | 1.72 | 1.725 | 1.7225 | 187.5691 | 0.20509
2020-02-06 | 12:45 | 8.565 | 8.569 | 8.567 | 913.4926 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.20515
2020-02-06 | 12:45 | 8.555 | 8.567 | 8.561 | 912.8528 | 1.719 | 1.722 | 1.7205 | 187.3513 | 0.20524
2020-02-06 | 12:46 | 8.54 | 8.554 | 8.547 | 911.36 | 1.713 | 1.718 | 1.7155 | 186.8068 | 0.20498
2020-02-06 | 12:46 | 8.51 | 8.531 | 8.5205 | 908.5343 | 1.708 | 1.711 | 1.7095 | 186.1534 | 0.20489
2020-02-06 | 12:46 | 8.492 | 8.5 | 8.496 | 905.9219 | 1.705 | 1.707 | 1.706 | 185.7723 | 0.20506
2020-02-06 | 12:46 | 8.498 | 8.507 | 8.5025 | 906.615 | 1.706 | 1.708 | 1.707 | 185.8812 | 0.20503
2020-02-06 | 12:46 | 8.513 | 8.52 | 8.5165 | 908.1078 | 1.707 | 1.709 | 1.708 | 185.9901 | 0.20481
2020-02-06 | 12:46 | 8.509 | 8.519 | 8.514 | 907.8412 | 1.708 | 1.71 | 1.709 | 186.099 | 0.20499
2020-02-06 | 12:46 | 8.501 | 8.509 | 8.505 | 906.8815 | 1.706 | 1.708 | 1.707 | 185.8812 | 0.20497
2020-02-06 | 12:46 | 8.496 | 8.509 | 8.5025 | 906.615 | 1.706 | 1.709 | 1.7075 | 185.9357 | 0.20509
2020-02-06 | 12:46 | 8.448 | 8.479 | 8.4635 | 902.4564 | 1.697 | 1.702 | 1.6995 | 185.0645 | 0.20507
2020-02-06 | 12:46 | 8.423 | 8.459 | 8.441 | 900.0573 | 1.692 | 1.7 | 1.696 | 184.6834 | 0.20519
2020-02-06 | 12:46 | 8.483 | 8.524 | 8.5035 | 906.7216 | 1.705 | 1.712 | 1.7085 | 186.0445 | 0.20518
2020-02-06 | 12:47 | 8.511 | 8.525 | 8.518 | 908.2677 | 1.708 | 1.709 | 1.7085 | 186.0445 | 0.20483
2020-02-06 | 12:47 | 8.497 | 8.507 | 8.502 | 906.5617 | 1.706 | 1.708 | 1.707 | 185.8812 | 0.20504
2020-02-06 | 12:47 | 8.498 | 8.502 | 8.5 | 906.3484 | 1.706 | 1.71 | 1.708 | 185.9901 | 0.20521
2020-02-06 | 12:47 | 8.503 | 8.51 | 8.5065 | 907.0415 | 1.71 | 1.711 | 1.7105 | 186.2623 | 0.20535
2020-02-06 | 12:47 | 8.513 | 8.519 | 8.516 | 908.0545 | 1.712 | 1.713 | 1.7125 | 186.4801 | 0.20536
2020-02-06 | 12:47 | 8.508 | 8.515 | 8.5115 | 907.5746 | 1.71 | 1.712 | 1.711 | 186.3168 | 0.20529
2020-02-06 | 12:47 | 8.494 | 8.503 | 8.4985 | 906.1885 | 1.708 | 1.71 | 1.709 | 186.099 | 0.20536
2020-02-06 | 12:47 | 8.492 | 8.494 | 8.493 | 905.602 | 1.707 | 1.709 | 1.708 | 185.9901 | 0.20538
2020-02-06 | 12:47 | 8.489 | 8.492 | 8.4905 | 905.3354 | 1.707 | 1.708 | 1.7075 | 185.9357 | 0.20538
2020-02-06 | 12:47 | 8.495 | 8.515 | 8.505 | 906.8815 | 1.708 | 1.712 | 1.71 | 186.2079 | 0.20533
2020-02-06 | 12:47 | 8.511 | 8.519 | 8.515 | 907.9478 | 1.711 | 1.712 | 1.7115 | 186.3712 | 0.20527
2020-02-06 | 12:47 | 8.503 | 8.508 | 8.5055 | 906.9349 | 1.709 | 1.711 | 1.71 | 186.2079 | 0.20532
2020-02-06 | 12:47 | 8.489 | 8.501 | 8.495 | 905.8152 | 1.707 | 1.71 | 1.7085 | 186.0445 | 0.20539
2020-02-06 | 12:48 | 8.472 | 8.486 | 8.479 | 904.1092 | 1.705 | 1.708 | 1.7065 | 185.8268 | 0.20554
2020-02-06 | 12:48 | 8.458 | 8.466 | 8.462 | 902.2965 | 1.702 | 1.704 | 1.703 | 185.4456 | 0.20553
2020-02-06 | 12:48 | 8.433 | 8.453 | 8.443 | 900.2705 | 1.699 | 1.701 | 1.7 | 185.119 | 0.20563
2020-02-06 | 12:48 | 8.434 | 8.448 | 8.441 | 900.0573 | 1.699 | 1.702 | 1.7005 | 185.1734 | 0.20574
2020-02-06 | 12:48 | 8.453 | 8.471 | 8.462 | 902.2965 | 1.703 | 1.704 | 1.7035 | 185.5001 | 0.20559
2020-02-06 | 12:48 | 8.47 | 8.473 | 8.4715 | 903.3095 | 1.705 | 1.706 | 1.7055 | 185.7179 | 0.2056
2020-02-06 | 12:48 | 8.468 | 8.473 | 8.4705 | 903.2028 | 1.703 | 1.705 | 1.704 | 185.5545 | 0.20544
2020-02-06 | 12:48 | 8.451 | 8.465 | 8.458 | 901.87 | 1.7 | 1.703 | 1.7015 | 185.2823 | 0.20544
2020-02-06 | 12:48 | 8.442 | 8.45 | 8.446 | 900.5904 | 1.7 | 1.702 | 1.701 | 185.2278 | 0.20567
2020-02-06 | 12:48 | 8.429 | 8.441 | 8.435 | 899.4175 | 1.698 | 1.7 | 1.699 | 185.0101 | 0.2057
2020-02-06 | 12:48 | 8.427 | 8.43 | 8.4285 | 898.7244 | 1.698 | 1.7 | 1.699 | 185.0101 | 0.20586
2020-02-06 | 12:49 | 8.425 | 8.43 | 8.4275 | 898.6178 | 1.699 | 1.7 | 1.6995 | 185.0645 | 0.20594
2020-02-06 | 12:49 | 8.422 | 8.425 | 8.4235 | 898.1913 | 1.698 | 1.7 | 1.699 | 185.0101 | 0.20598
2020-02-06 | 12:49 | 8.421 | 8.423 | 8.422 | 898.0313 | 1.698 | 1.7 | 1.699 | 185.0101 | 0.20602
2020-02-06 | 12:49 | 8.418 | 8.425 | 8.4215 | 897.978 | 1.699 | 1.7 | 1.6995 | 185.0645 | 0.20609
2020-02-06 | 12:49 | 8.411 | 8.418 | 8.4145 | 897.2316 | 1.697 | 1.698 | 1.6975 | 184.8467 | 0.20602
2020-02-06 | 12:49 | 8.401 | 8.41 | 8.4055 | 896.2719 | 1.695 | 1.697 | 1.696 | 184.6834 | 0.20606
2020-02-06 | 12:49 | 8.394 | 8.401 | 8.3975 | 895.4189 | 1.693 | 1.695 | 1.694 | 184.4656 | 0.20601
2020-02-06 | 12:49 | 8.381 | 8.391 | 8.386 | 894.1927 | 1.689 | 1.693 | 1.691 | 184.1389 | 0.20593
2020-02-06 | 12:49 | 8.365 | 8.377 | 8.371 | 892.5932 | 1.687 | 1.69 | 1.6885 | 183.8667 | 0.20599
2020-02-06 | 12:49 | 8.357 | 8.361 | 8.359 | 891.3137 | 1.686 | 1.687 | 1.6865 | 183.6489 | 0.20604
2020-02-06 | 12:49 | 8.356 | 8.36 | 8.358 | 891.207 | 1.686 | 1.688 | 1.687 | 183.7033 | 0.20613
2020-02-06 | 12:49 | 8.361 | 8.368 | 8.3645 | 891.9001 | 1.686 | 1.689 | 1.6875 | 183.7578 | 0.20603
2020-02-06 | 12:49 | 8.351 | 8.371 | 8.361 | 891.5269 | 1.676 | 1.687 | 1.6815 | 183.1044 | 0.20538
2020-02-06 | 12:50 | 7.948 | 8.253 | 8.1005 | 863.75 | 1.582 | 1.646 | 1.614 | 175.7541 | 0.20348
2020-02-06 | 12:50 | 8.121 | 8.274 | 8.1975 | 874.0931 | 1.632 | 1.669 | 1.6505 | 179.7287 | 0.20562
2020-02-06 | 12:50 | 8.306 | 8.388 | 8.347 | 890.0341 | 1.679 | 1.696 | 1.6875 | 183.7578 | 0.20646
2020-02-06 | 12:50 | 8.393 | 8.412 | 8.4025 | 895.952 | 1.697 | 1.7 | 1.6985 | 184.9556 | 0.20643
2020-02-06 | 12:50 | 8.416 | 8.435 | 8.4255 | 898.4045 | 1.701 | 1.706 | 1.7035 | 185.5001 | 0.20648
2020-02-06 | 12:50 | 8.44 | 8.457 | 8.4485 | 900.857 | 1.708 | 1.71 | 1.709 | 186.099 | 0.20658
2020-02-06 | 12:50 | 8.462 | 8.482 | 8.472 | 903.3628 | 1.712 | 1.715 | 1.7135 | 186.589 | 0.20655
2020-02-06 | 12:50 | 8.488 | 8.504 | 8.496 | 905.9219 | 1.716 | 1.72 | 1.718 | 187.079 | 0.20651
2020-02-06 | 12:50 | 8.51 | 8.525 | 8.5175 | 908.2144 | 1.72 | 1.724 | 1.722 | 187.5146 | 0.20647
2020-02-06 | 12:50 | 8.531 | 8.544 | 8.5375 | 910.347 | 1.724 | 1.726 | 1.725 | 187.8413 | 0.20634
2020-02-06 | 12:50 | 8.544 | 8.548 | 8.546 | 911.2533 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.20637
2020-02-06 | 12:51 | 8.545 | 8.548 | 8.5465 | 911.3067 | 1.726 | 1.727 | 1.7265 | 188.0046 | 0.2063
2020-02-06 | 12:51 | 8.544 | 8.549 | 8.5465 | 911.3067 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.20636
2020-02-06 | 12:51 | 8.544 | 8.548 | 8.546 | 911.2533 | 1.725 | 1.727 | 1.726 | 187.9502 | 0.20625
2020-02-06 | 12:51 | 8.548 | 8.552 | 8.55 | 911.6799 | 1.725 | 1.726 | 1.7255 | 187.8957 | 0.2061
2020-02-06 | 12:51 | 8.55 | 8.553 | 8.5515 | 911.8398 | 1.724 | 1.726 | 1.725 | 187.8413 | 0.206
2020-02-06 | 12:51 | 8.548 | 8.554 | 8.551 | 911.7865 | 1.724 | 1.725 | 1.7245 | 187.7868 | 0.20595
2020-02-06 | 12:51 | 8.55 | 8.553 | 8.5515 | 911.8398 | 1.723 | 1.725 | 1.724 | 187.7324 | 0.20588
2020-02-06 | 12:51 | 8.541 | 8.547 | 8.544 | 911.0401 | 1.722 | 1.723 | 1.7225 | 187.5691 | 0.20588
2020-02-06 | 12:51 | 8.532 | 8.539 | 8.5355 | 910.1337 | 1.721 | 1.722 | 1.7215 | 187.4602 | 0.20597
2020-02-06 | 12:51 | 8.525 | 8.531 | 8.528 | 909.334 | 1.718 | 1.721 | 1.7195 | 187.2424 | 0.20591
2020-02-06 | 12:51 | 8.514 | 8.521 | 8.5175 | 908.2144 | 1.716 | 1.719 | 1.7175 | 187.0246 | 0.20593
2020-02-06 | 12:51 | 8.511 | 8.514 | 8.5125 | 907.6813 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20593
2020-02-06 | 12:51 | 8.514 | 8.518 | 8.516 | 908.0545 | 1.716 | 1.718 | 1.717 | 186.9701 | 0.2059
2020-02-06 | 12:52 | 8.499 | 8.51 | 8.5045 | 906.8282 | 1.713 | 1.715 | 1.714 | 186.6435 | 0.20582
2020-02-06 | 12:52 | 8.497 | 8.5 | 8.4985 | 906.1885 | 1.713 | 1.715 | 1.714 | 186.6435 | 0.20597
2020-02-06 | 12:52 | 8.487 | 8.494 | 8.4905 | 905.3354 | 1.71 | 1.713 | 1.7115 | 186.3712 | 0.20586
2020-02-06 | 12:52 | 8.474 | 8.486 | 8.48 | 904.2158 | 1.708 | 1.711 | 1.7095 | 186.1534 | 0.20587
2020-02-06 | 12:52 | 8.466 | 8.472 | 8.469 | 903.0429 | 1.705 | 1.708 | 1.7065 | 185.8268 | 0.20578
2020-02-06 | 12:52 | 8.461 | 8.467 | 8.464 | 902.5097 | 1.706 | 1.707 | 1.7065 | 185.8268 | 0.2059
2020-02-06 | 12:52 | 8.471 | 8.483 | 8.477 | 903.8959 | 1.708 | 1.708 | 1.708 | 185.9901 | 0.20576
2020-02-06 | 12:52 | 8.485 | 8.489 | 8.487 | 904.9622 | 1.709 | 1.711 | 1.71 | 186.2079 | 0.20576
2020-02-06 | 12:52 | 8.492 | 8.503 | 8.4975 | 906.0818 | 1.711 | 1.713 | 1.712 | 186.4257 | 0.20575
2020-02-06 | 12:52 | 8.502 | 8.507 | 8.5045 | 906.8282 | 1.711 | 1.712 | 1.7115 | 186.3712 | 0.20552
2020-02-06 | 12:52 | 8.509 | 8.512 | 8.5105 | 907.468 | 1.711 | 1.714 | 1.7125 | 186.4801 | 0.20549
2020-02-06 | 12:52 | 8.51 | 8.514 | 8.512 | 907.6279 | 1.712 | 1.714 | 1.713 | 186.5346 | 0.20552
2020-02-06 | 12:53 | 8.51 | 8.512 | 8.511 | 907.5213 | 1.712 | 1.713 | 1.7125 | 186.4801 | 0.20548
2020-02-06 | 12:53 | 8.504 | 8.511 | 8.5075 | 907.1481 | 1.712 | 1.714 | 1.713 | 186.5346 | 0.20563
2020-02-06 | 12:53 | 8.505 | 8.506 | 8.5055 | 906.9349 | 1.712 | 1.713 | 1.7125 | 186.4801 | 0.20562
2020-02-06 | 12:53 | 8.504 | 8.515 | 8.5095 | 907.3614 | 1.713 | 1.715 | 1.714 | 186.6435 | 0.2057
2020-02-06 | 12:53 | 8.513 | 8.516 | 8.5145 | 907.8945 | 1.714 | 1.715 | 1.7145 | 186.6979 | 0.20564
2020-02-06 | 12:53 | 8.515 | 8.518 | 8.5165 | 908.1078 | 1.714 | 1.716 | 1.715 | 186.7524 | 0.20565
2020-02-06 | 12:53 | 8.513 | 8.518 | 8.5155 | 908.0011 | 1.714 | 1.716 | 1.715 | 186.7524 | 0.20567
2020-02-06 | 12:53 | 8.509 | 8.511 | 8.51 | 907.4147 | 1.714 | 1.716 | 1.715 | 186.7524 | 0.20581
2020-02-06 | 12:53 | 8.509 | 8.517 | 8.513 | 907.7346 | 1.714 | 1.716 | 1.715 | 186.7524 | 0.20573
2020-02-06 | 12:53 | 8.514 | 8.517 | 8.5155 | 908.0011 | 1.714 | 1.717 | 1.7155 | 186.8068 | 0.20573
2020-02-06 | 12:53 | 8.512 | 8.515 | 8.5135 | 907.7879 | 1.715 | 1.716 | 1.7155 | 186.8068 | 0.20578
2020-02-06 | 12:54 | 8.511 | 8.512 | 8.5115 | 907.5746 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20589
2020-02-06 | 12:54 | 8.514 | 8.519 | 8.5165 | 908.1078 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20577
2020-02-06 | 12:54 | 8.519 | 8.525 | 8.522 | 908.6942 | 1.716 | 1.719 | 1.7175 | 187.0246 | 0.20582
2020-02-06 | 12:54 | 8.526 | 8.53 | 8.528 | 909.334 | 1.717 | 1.72 | 1.7185 | 187.1335 | 0.20579
2020-02-06 | 12:54 | 8.523 | 8.528 | 8.5255 | 909.0674 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20555
2020-02-06 | 12:54 | 8.519 | 8.524 | 8.5215 | 908.6409 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20565
2020-02-06 | 12:54 | 8.521 | 8.525 | 8.523 | 908.8009 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20567
2020-02-06 | 12:54 | 8.523 | 8.525 | 8.524 | 908.9075 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20577
2020-02-06 | 12:54 | 8.523 | 8.527 | 8.525 | 909.0141 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20574
2020-02-06 | 12:54 | 8.531 | 8.539 | 8.535 | 910.0804 | 1.718 | 1.72 | 1.719 | 187.1879 | 0.20568
2020-02-06 | 12:54 | 8.542 | 8.546 | 8.544 | 911.0401 | 1.72 | 1.721 | 1.7205 | 187.3513 | 0.20565
2020-02-06 | 12:54 | 8.546 | 8.552 | 8.549 | 911.5732 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.20558
2020-02-06 | 12:54 | 8.553 | 8.558 | 8.5555 | 912.2663 | 1.722 | 1.723 | 1.7225 | 187.5691 | 0.20561
2020-02-06 | 12:55 | 8.554 | 8.556 | 8.555 | 912.213 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.20556
2020-02-06 | 12:55 | 8.55 | 8.554 | 8.552 | 911.8931 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.20551
2020-02-06 | 12:55 | 8.547 | 8.55 | 8.5485 | 911.5199 | 1.719 | 1.72 | 1.7195 | 187.2424 | 0.20542
2020-02-06 | 12:55 | 8.545 | 8.551 | 8.548 | 911.4666 | 1.719 | 1.72 | 1.7195 | 187.2424 | 0.20543
2020-02-06 | 12:55 | 8.545 | 8.548 | 8.5465 | 911.3067 | 1.718 | 1.72 | 1.719 | 187.1879 | 0.20541
2020-02-06 | 12:55 | 8.54 | 8.545 | 8.5425 | 910.8801 | 1.718 | 1.719 | 1.7185 | 187.1335 | 0.20544
2020-02-06 | 12:55 | 8.537 | 8.541 | 8.539 | 910.5069 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20541
2020-02-06 | 12:55 | 8.536 | 8.538 | 8.537 | 910.2937 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20546
2020-02-06 | 12:55 | 8.537 | 8.538 | 8.5375 | 910.347 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20544
2020-02-06 | 12:55 | 8.538 | 8.539 | 8.5385 | 910.4536 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.2053
2020-02-06 | 12:55 | 8.534 | 8.537 | 8.5355 | 910.1337 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20537
2020-02-06 | 12:56 | 8.534 | 8.537 | 8.5355 | 910.1337 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20537
2020-02-06 | 12:56 | 8.539 | 8.543 | 8.541 | 910.7202 | 1.716 | 1.719 | 1.7175 | 187.0246 | 0.20536
2020-02-06 | 12:56 | 8.544 | 8.555 | 8.5495 | 911.6265 | 1.719 | 1.72 | 1.7195 | 187.2424 | 0.20539
2020-02-06 | 12:56 | 8.553 | 8.559 | 8.556 | 912.3196 | 1.719 | 1.721 | 1.72 | 187.2968 | 0.2053
2020-02-06 | 12:56 | 8.56 | 8.565 | 8.5625 | 913.0127 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.20526
2020-02-06 | 12:56 | 8.566 | 8.57 | 8.568 | 913.5992 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.20525
2020-02-06 | 12:56 | 8.568 | 8.572 | 8.57 | 913.8124 | 1.722 | 1.723 | 1.7225 | 187.5691 | 0.20526
2020-02-06 | 12:56 | 8.571 | 8.574 | 8.5725 | 914.079 | 1.722 | 1.725 | 1.7235 | 187.678 | 0.20532
2020-02-06 | 12:56 | 8.566 | 8.571 | 8.5685 | 913.6525 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.20524
2020-02-06 | 12:56 | 8.564 | 8.567 | 8.5655 | 913.3326 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.20531
2020-02-06 | 12:56 | 8.553 | 8.562 | 8.5575 | 912.4796 | 1.719 | 1.723 | 1.721 | 187.4057 | 0.20538
2020-02-06 | 12:56 | 8.548 | 8.553 | 8.5505 | 911.7332 | 1.719 | 1.721 | 1.72 | 187.2968 | 0.20543
2020-02-06 | 12:56 | 8.545 | 8.547 | 8.546 | 911.2533 | 1.717 | 1.72 | 1.7185 | 187.1335 | 0.20536
2020-02-06 | 12:57 | 8.543 | 8.545 | 8.544 | 911.0401 | 1.717 | 1.719 | 1.718 | 187.079 | 0.20535
2020-02-06 | 12:57 | 8.544 | 8.546 | 8.545 | 911.1467 | 1.718 | 1.719 | 1.7185 | 187.1335 | 0.20538
2020-02-06 | 12:57 | 8.547 | 8.55 | 8.5485 | 911.5199 | 1.718 | 1.719 | 1.7185 | 187.1335 | 0.2053
2020-02-06 | 12:57 | 8.55 | 8.552 | 8.551 | 911.7865 | 1.72 | 1.721 | 1.7205 | 187.3513 | 0.20548
2020-02-06 | 12:57 | 8.549 | 8.551 | 8.55 | 911.6799 | 1.72 | 1.721 | 1.7205 | 187.3513 | 0.2055
2020-02-06 | 12:57 | 8.551 | 8.556 | 8.5535 | 912.0531 | 1.719 | 1.722 | 1.7205 | 187.3513 | 0.20542
2020-02-06 | 12:57 | 8.552 | 8.555 | 8.5535 | 912.0531 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.2056
2020-02-06 | 12:57 | 8.554 | 8.559 | 8.5565 | 912.3729 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.20552
2020-02-06 | 12:57 | 8.532 | 8.549 | 8.5405 | 910.6669 | 1.719 | 1.722 | 1.7205 | 187.3513 | 0.20573
2020-02-06 | 12:57 | 8.534 | 8.54 | 8.537 | 910.2937 | 1.72 | 1.721 | 1.7205 | 187.3513 | 0.20581
2020-02-06 | 12:57 | 8.532 | 8.545 | 8.5385 | 910.4536 | 1.719 | 1.722 | 1.7205 | 187.3513 | 0.20578
2020-02-06 | 12:58 | 8.535 | 8.539 | 8.537 | 910.2937 | 1.72 | 1.72 | 1.72 | 187.2968 | 0.20575
2020-02-06 | 12:58 | 8.532 | 8.536 | 8.534 | 909.9738 | 1.719 | 1.72 | 1.7195 | 187.2424 | 0.20577
2020-02-06 | 12:58 | 8.538 | 8.542 | 8.54 | 910.6136 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.2058
2020-02-06 | 12:58 | 8.541 | 8.554 | 8.5475 | 911.4133 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.20574
2020-02-06 | 12:58 | 8.557 | 8.564 | 8.5605 | 912.7995 | 1.723 | 1.726 | 1.7245 | 187.7868 | 0.20573
2020-02-06 | 12:58 | 8.566 | 8.57 | 8.568 | 913.5992 | 1.725 | 1.727 | 1.726 | 187.9502 | 0.20572
2020-02-06 | 12:58 | 8.562 | 8.568 | 8.565 | 913.2793 | 1.726 | 1.727 | 1.7265 | 188.0046 | 0.20586
2020-02-06 | 12:58 | 8.559 | 8.562 | 8.5605 | 912.7995 | 1.725 | 1.726 | 1.7255 | 187.8957 | 0.20585
2020-02-06 | 12:58 | 8.559 | 8.564 | 8.5615 | 912.9061 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.206
2020-02-06 | 12:58 | 8.565 | 8.567 | 8.566 | 913.3859 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.20589
2020-02-06 | 12:58 | 8.566 | 8.576 | 8.571 | 913.9191 | 1.727 | 1.729 | 1.728 | 188.168 | 0.20589
2020-02-06 | 12:58 | 8.568 | 8.576 | 8.572 | 914.0257 | 1.726 | 1.73 | 1.728 | 188.168 | 0.20587
2020-02-06 | 12:58 | 8.553 | 8.563 | 8.558 | 912.5329 | 1.723 | 1.726 | 1.7245 | 187.7868 | 0.20579
2020-02-06 | 12:59 | 8.537 | 8.55 | 8.5435 | 910.9868 | 1.72 | 1.723 | 1.7215 | 187.4602 | 0.20578
2020-02-06 | 12:59 | 8.526 | 8.534 | 8.53 | 909.5473 | 1.718 | 1.72 | 1.719 | 187.1879 | 0.2058
2020-02-06 | 12:59 | 8.524 | 8.528 | 8.526 | 909.1208 | 1.718 | 1.72 | 1.719 | 187.1879 | 0.2059
2020-02-06 | 12:59 | 8.523 | 8.525 | 8.524 | 908.9075 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20577
2020-02-06 | 12:59 | 8.52 | 8.523 | 8.5215 | 908.6409 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20565
2020-02-06 | 12:59 | 8.523 | 8.525 | 8.524 | 908.9075 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20565
2020-02-06 | 12:59 | 8.518 | 8.523 | 8.5205 | 908.5343 | 1.715 | 1.716 | 1.7155 | 186.8068 | 0.20561
2020-02-06 | 12:59 | 8.519 | 8.52 | 8.5195 | 908.4277 | 1.714 | 1.717 | 1.7155 | 186.8068 | 0.20564
2020-02-06 | 12:59 | 8.508 | 8.519 | 8.5135 | 907.7879 | 1.712 | 1.715 | 1.7135 | 186.589 | 0.20554
2020-02-06 | 12:59 | 8.5 | 8.509 | 8.5045 | 906.8282 | 1.71 | 1.713 | 1.7115 | 186.3712 | 0.20552
2020-02-06 | 12:59 | 8.489 | 8.499 | 8.494 | 905.7086 | 1.709 | 1.711 | 1.71 | 186.2079 | 0.20559
2020-02-06 | 13:00 | 8.483 | 8.489 | 8.486 | 904.8556 | 1.708 | 1.709 | 1.7085 | 186.0445 | 0.20561
2020-02-06 | 13:00 | 8.483 | 8.485 | 8.484 | 904.6423 | 1.708 | 1.71 | 1.709 | 186.099 | 0.20572
2020-02-06 | 13:00 | 8.476 | 8.483 | 8.4795 | 904.1625 | 1.707 | 1.708 | 1.7075 | 185.9357 | 0.20564
2020-02-06 | 13:00 | 8.469 | 8.475 | 8.472 | 903.3628 | 1.706 | 1.707 | 1.7065 | 185.8268 | 0.20571
2020-02-06 | 13:00 | 8.456 | 8.468 | 8.462 | 902.2965 | 1.703 | 1.707 | 1.705 | 185.6634 | 0.20577
2020-02-06 | 13:00 | 8.449 | 8.456 | 8.4525 | 901.2835 | 1.702 | 1.703 | 1.7025 | 185.3912 | 0.2057
2020-02-06 | 13:00 | 8.453 | 8.462 | 8.4575 | 901.8167 | 1.703 | 1.705 | 1.704 | 185.5545 | 0.20576
2020-02-06 | 13:00 | 8.462 | 8.466 | 8.464 | 902.5097 | 1.703 | 1.704 | 1.7035 | 185.5001 | 0.20554
2020-02-06 | 13:00 | 8.464 | 8.469 | 8.4665 | 902.7763 | 1.704 | 1.705 | 1.7045 | 185.609 | 0.2056
2020-02-06 | 13:00 | 8.469 | 8.47 | 8.4695 | 903.0962 | 1.703 | 1.705 | 1.704 | 185.5545 | 0.20546
2020-02-06 | 13:00 | 8.466 | 8.469 | 8.4675 | 902.8829 | 1.703 | 1.705 | 1.704 | 185.5545 | 0.20551
2020-02-06 | 13:00 | 8.463 | 8.467 | 8.465 | 902.6164 | 1.703 | 1.705 | 1.704 | 185.5545 | 0.20557
2020-02-06 | 13:00 | 8.462 | 8.463 | 8.4625 | 902.3498 | 1.702 | 1.703 | 1.7025 | 185.3912 | 0.20545
2020-02-06 | 13:01 | 8.459 | 8.462 | 8.4605 | 902.1365 | 1.702 | 1.703 | 1.7025 | 185.3912 | 0.2055
2020-02-06 | 13:01 | 8.451 | 8.457 | 8.454 | 901.4435 | 1.701 | 1.703 | 1.702 | 185.3367 | 0.2056
2020-02-06 | 13:01 | 8.447 | 8.452 | 8.4495 | 900.9636 | 1.701 | 1.701 | 1.701 | 185.2278 | 0.20559
2020-02-06 | 13:01 | 8.444 | 8.447 | 8.4455 | 900.5371 | 1.7 | 1.702 | 1.701 | 185.2278 | 0.20569
2020-02-06 | 13:01 | 8.441 | 8.446 | 8.4435 | 900.3238 | 1.7 | 1.702 | 1.701 | 185.2278 | 0.20573
2020-02-06 | 13:01 | 8.435 | 8.443 | 8.439 | 899.844 | 1.698 | 1.701 | 1.6995 | 185.0645 | 0.20566
2020-02-06 | 13:01 | 8.435 | 8.436 | 8.4355 | 899.4708 | 1.698 | 1.7 | 1.699 | 185.0101 | 0.20569
2020-02-06 | 13:01 | 8.438 | 8.443 | 8.4405 | 900.004 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.20544
2020-02-06 | 13:01 | 8.449 | 8.458 | 8.4535 | 901.3901 | 1.695 | 1.698 | 1.6965 | 184.7378 | 0.20495
2020-02-06 | 13:01 | 8.454 | 8.466 | 8.46 | 902.0832 | 1.695 | 1.699 | 1.697 | 184.7923 | 0.20485
2020-02-06 | 13:01 | 8.468 | 8.481 | 8.4745 | 903.6294 | 1.702 | 1.704 | 1.703 | 185.4456 | 0.20522
2020-02-06 | 13:02 | 8.483 | 8.495 | 8.489 | 905.1755 | 1.704 | 1.706 | 1.705 | 185.6634 | 0.20511
2020-02-06 | 13:02 | 8.498 | 8.508 | 8.503 | 906.6683 | 1.706 | 1.709 | 1.7075 | 185.9357 | 0.20508
2020-02-06 | 13:02 | 8.507 | 8.516 | 8.5115 | 907.5746 | 1.708 | 1.711 | 1.7095 | 186.1534 | 0.20511
2020-02-06 | 13:02 | 8.516 | 8.521 | 8.5185 | 908.321 | 1.71 | 1.713 | 1.7115 | 186.3712 | 0.20518
2020-02-06 | 13:02 | 8.518 | 8.527 | 8.5225 | 908.7476 | 1.711 | 1.715 | 1.713 | 186.5346 | 0.20527
2020-02-06 | 13:02 | 8.528 | 8.537 | 8.5325 | 909.8138 | 1.714 | 1.716 | 1.715 | 186.7524 | 0.20526
2020-02-06 | 13:02 | 8.537 | 8.54 | 8.5385 | 910.4536 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20524
2020-02-06 | 13:02 | 8.536 | 8.539 | 8.5375 | 910.347 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20526
2020-02-06 | 13:02 | 8.536 | 8.542 | 8.539 | 910.5069 | 1.715 | 1.716 | 1.7155 | 186.8068 | 0.20517
2020-02-06 | 13:02 | 8.538 | 8.541 | 8.5395 | 910.5602 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20528
2020-02-06 | 13:02 | 8.534 | 8.537 | 8.5355 | 910.1337 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20531
2020-02-06 | 13:02 | 8.531 | 8.534 | 8.5325 | 909.8138 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.20544
2020-02-06 | 13:02 | 8.532 | 8.537 | 8.5345 | 910.0271 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20552
2020-02-06 | 13:03 | 8.534 | 8.538 | 8.536 | 910.187 | 1.717 | 1.719 | 1.718 | 187.079 | 0.20554
2020-02-06 | 13:03 | 8.537 | 8.541 | 8.539 | 910.5069 | 1.718 | 1.719 | 1.7185 | 187.1335 | 0.20553
2020-02-06 | 13:03 | 8.535 | 8.539 | 8.537 | 910.2937 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.20546
2020-02-06 | 13:03 | 8.53 | 8.544 | 8.537 | 910.2937 | 1.716 | 1.72 | 1.718 | 187.079 | 0.20552
2020-02-06 | 13:03 | 8.545 | 8.548 | 8.5465 | 911.3067 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.20565
2020-02-06 | 13:03 | 8.552 | 8.566 | 8.559 | 912.6395 | 1.722 | 1.723 | 1.7225 | 187.5691 | 0.20552
2020-02-06 | 13:03 | 8.569 | 8.57 | 8.5695 | 913.7591 | 1.724 | 1.725 | 1.7245 | 187.7868 | 0.20551
2020-02-06 | 13:03 | 8.571 | 8.579 | 8.575 | 914.3456 | 1.725 | 1.726 | 1.7255 | 187.8957 | 0.2055
2020-02-06 | 13:03 | 8.578 | 8.586 | 8.582 | 915.092 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.20551
2020-02-06 | 13:03 | 8.58 | 8.586 | 8.583 | 915.1986 | 1.726 | 1.729 | 1.7275 | 188.1135 | 0.20554
2020-02-06 | 13:03 | 8.589 | 8.592 | 8.5905 | 915.9983 | 1.728 | 1.731 | 1.7295 | 188.3313 | 0.2056
2020-02-06 | 13:04 | 8.591 | 8.592 | 8.5915 | 916.105 | 1.73 | 1.731 | 1.7305 | 188.4402 | 0.2057
2020-02-06 | 13:04 | 8.592 | 8.596 | 8.594 | 916.3715 | 1.73 | 1.732 | 1.731 | 188.4947 | 0.2057
2020-02-06 | 13:04 | 8.596 | 8.602 | 8.599 | 916.9047 | 1.732 | 1.733 | 1.7325 | 188.658 | 0.20576
2020-02-06 | 13:04 | 8.602 | 8.607 | 8.6045 | 917.4911 | 1.733 | 1.734 | 1.7335 | 188.7669 | 0.20574
2020-02-06 | 13:04 | 8.607 | 8.612 | 8.6095 | 918.0243 | 1.734 | 1.734 | 1.734 | 188.8213 | 0.20568
2020-02-06 | 13:04 | 8.611 | 8.615 | 8.613 | 918.3975 | 1.735 | 1.737 | 1.736 | 189.0391 | 0.20584
2020-02-06 | 13:04 | 8.578 | 8.607 | 8.5925 | 916.2116 | 1.729 | 1.735 | 1.732 | 188.6035 | 0.20585
2020-02-06 | 13:04 | 8.556 | 8.572 | 8.564 | 913.1727 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.20594
2020-02-06 | 13:04 | 8.559 | 8.573 | 8.566 | 913.3859 | 1.727 | 1.73 | 1.7285 | 188.2224 | 0.20607
2020-02-06 | 13:04 | 8.577 | 8.586 | 8.5815 | 915.0387 | 1.73 | 1.731 | 1.7305 | 188.4402 | 0.20594
2020-02-06 | 13:04 | 8.588 | 8.596 | 8.592 | 916.1583 | 1.731 | 1.733 | 1.732 | 188.6035 | 0.20586
2020-02-06 | 13:04 | 8.579 | 8.593 | 8.586 | 915.5185 | 1.729 | 1.732 | 1.7305 | 188.4402 | 0.20583
2020-02-06 | 13:04 | 8.495 | 8.578 | 8.5365 | 910.2404 | 1.706 | 1.729 | 1.7175 | 187.0246 | 0.20547
2020-02-06 | 13:05 | 8.125 | 8.416 | 8.2705 | 881.877 | 1.631 | 1.687 | 1.659 | 180.6543 | 0.20485
2020-02-06 | 13:05 | 8.195 | 8.425 | 8.31 | 886.0888 | 1.649 | 1.702 | 1.6755 | 182.4511 | 0.20591
2020-02-06 | 13:05 | 8.469 | 8.531 | 8.5 | 906.3484 | 1.711 | 1.721 | 1.716 | 186.8612 | 0.20617
2020-02-06 | 13:05 | 8.545 | 8.569 | 8.557 | 912.4263 | 1.724 | 1.728 | 1.726 | 187.9502 | 0.20599
2020-02-06 | 13:05 | 8.573 | 8.576 | 8.5745 | 914.2923 | 1.728 | 1.729 | 1.7285 | 188.2224 | 0.20587
2020-02-06 | 13:05 | 8.56 | 8.575 | 8.5675 | 913.5459 | 1.727 | 1.729 | 1.728 | 188.168 | 0.20598
2020-02-06 | 13:05 | 8.557 | 8.561 | 8.559 | 912.6395 | 1.725 | 1.727 | 1.726 | 187.9502 | 0.20594
2020-02-06 | 13:05 | 8.553 | 8.557 | 8.555 | 912.213 | 1.724 | 1.727 | 1.7255 | 187.8957 | 0.20598
2020-02-06 | 13:05 | 8.534 | 8.554 | 8.544 | 911.0401 | 1.722 | 1.726 | 1.724 | 187.7324 | 0.20606
2020-02-06 | 13:05 | 8.531 | 8.54 | 8.5355 | 910.1337 | 1.723 | 1.724 | 1.7235 | 187.678 | 0.20621
2020-02-06 | 13:05 | 8.542 | 8.594 | 8.568 | 913.5992 | 1.725 | 1.735 | 1.73 | 188.3858 | 0.2062
2020-02-06 | 13:06 | 8.597 | 8.607 | 8.602 | 917.2246 | 1.736 | 1.737 | 1.7365 | 189.0936 | 0.20616
2020-02-06 | 13:06 | 8.61 | 8.633 | 8.6215 | 919.3038 | 1.737 | 1.741 | 1.739 | 189.3658 | 0.20599
2020-02-06 | 13:06 | 8.631 | 8.634 | 8.6325 | 920.4768 | 1.742 | 1.743 | 1.7425 | 189.7469 | 0.20614
2020-02-06 | 13:06 | 8.628 | 8.631 | 8.6295 | 920.1569 | 1.741 | 1.742 | 1.7415 | 189.638 | 0.20609
2020-02-06 | 13:06 | 8.623 | 8.631 | 8.627 | 919.8903 | 1.739 | 1.742 | 1.7405 | 189.5291 | 0.20603
2020-02-06 | 13:06 | 8.618 | 8.628 | 8.623 | 919.4638 | 1.739 | 1.741 | 1.74 | 189.4747 | 0.20607
2020-02-06 | 13:06 | 8.595 | 8.617 | 8.606 | 917.6511 | 1.733 | 1.739 | 1.736 | 189.0391 | 0.206
2020-02-06 | 13:06 | 8.585 | 8.591 | 8.588 | 915.7318 | 1.733 | 1.734 | 1.7335 | 188.7669 | 0.20614
2020-02-06 | 13:06 | 8.557 | 8.583 | 8.57 | 913.8124 | 1.726 | 1.732 | 1.729 | 188.2769 | 0.20603
2020-02-06 | 13:06 | 8.514 | 8.572 | 8.543 | 910.9335 | 1.717 | 1.731 | 1.724 | 187.7324 | 0.20609
2020-02-06 | 13:06 | 8.585 | 8.599 | 8.592 | 916.1583 | 1.733 | 1.737 | 1.735 | 188.9302 | 0.20622
2020-02-06 | 13:06 | 8.596 | 8.61 | 8.603 | 917.3312 | 1.736 | 1.74 | 1.738 | 189.2569 | 0.20631
2020-02-06 | 13:06 | 8.609 | 8.62 | 8.6145 | 918.5574 | 1.741 | 1.743 | 1.742 | 189.6925 | 0.20651
2020-02-06 | 13:07 | 8.621 | 8.631 | 8.626 | 919.7837 | 1.742 | 1.745 | 1.7435 | 189.8558 | 0.20641
2020-02-06 | 13:07 | 8.641 | 8.664 | 8.6525 | 922.6094 | 1.747 | 1.751 | 1.749 | 190.4547 | 0.20643
2020-02-06 | 13:07 | 8.252 | 8.668 | 8.46 | 902.0832 | 1.649 | 1.751 | 1.7 | 185.119 | 0.20521
2020-02-06 | 13:07 | 8.082 | 8.214 | 8.148 | 868.8149 | 1.621 | 1.65 | 1.6355 | 178.0953 | 0.20499
2020-02-06 | 13:07 | 8.101 | 8.202 | 8.1515 | 869.1881 | 1.619 | 1.646 | 1.6325 | 177.7686 | 0.20452
2020-02-06 | 13:07 | 7.317 | 8.102 | 7.7095 | 822.058 | 1.427 | 1.619 | 1.523 | 165.8448 | 0.20174
2020-02-06 | 13:07 | 6.124 | 6.874 | 6.499 | 692.9833 | 1.181 | 1.353 | 1.267 | 137.9681 | 0.19909
2020-02-06 | 13:07 | 7.27 | 8.204 | 7.737 | 824.9903 | 1.473 | 1.664 | 1.5685 | 170.7995 | 0.20703
2020-02-06 | 13:07 | 8.13 | 8.299 | 8.2145 | 875.9058 | 1.635 | 1.677 | 1.656 | 180.3276 | 0.20588
2020-02-06 | 13:07 | 7.952 | 8.102 | 8.027 | 855.9128 | 1.594 | 1.635 | 1.6145 | 175.8086 | 0.2054
2020-02-06 | 13:07 | 8.232 | 8.301 | 8.2665 | 881.4505 | 1.661 | 1.674 | 1.6675 | 181.5799 | 0.206
2020-02-06 | 13:08 | 8.326 | 8.379 | 8.3525 | 890.6206 | 1.675 | 1.691 | 1.683 | 183.2678 | 0.20578
2020-02-06 | 13:08 | 7.533 | 8.226 | 7.8795 | 840.185 | 1.489 | 1.648 | 1.5685 | 170.7995 | 0.20329
2020-02-06 | 13:08 | 7.896 | 8.164 | 8.03 | 856.2327 | 1.602 | 1.645 | 1.6235 | 176.7886 | 0.20647
2020-02-06 | 13:08 | 7.779 | 7.91 | 7.8445 | 836.4529 | 1.546 | 1.584 | 1.565 | 170.4183 | 0.20374
2020-02-06 | 13:08 | 7.747 | 8.407 | 8.077 | 861.2442 | 1.545 | 1.703 | 1.624 | 176.843 | 0.20533
2020-02-06 | 13:08 | 8.466 | 8.526 | 8.496 | 905.9219 | 1.714 | 1.724 | 1.719 | 187.1879 | 0.20663
2020-02-06 | 13:08 | 8.493 | 8.523 | 8.508 | 907.2014 | 1.716 | 1.721 | 1.7185 | 187.1335 | 0.20628
2020-02-06 | 13:08 | 8.409 | 8.474 | 8.4415 | 900.1106 | 1.702 | 1.714 | 1.708 | 185.9901 | 0.20663
2020-02-06 | 13:08 | 8.409 | 8.448 | 8.4285 | 898.7244 | 1.701 | 1.71 | 1.7055 | 185.7179 | 0.20665
2020-02-06 | 13:08 | 8.463 | 8.511 | 8.487 | 904.9622 | 1.712 | 1.719 | 1.7155 | 186.8068 | 0.20642
2020-02-06 | 13:08 | 8.512 | 8.514 | 8.513 | 907.7346 | 1.72 | 1.721 | 1.7205 | 187.3513 | 0.20639
2020-02-06 | 13:08 | 8.502 | 8.515 | 8.5085 | 907.2547 | 1.717 | 1.72 | 1.7185 | 187.1335 | 0.20626
2020-02-06 | 13:08 | 8.488 | 8.496 | 8.492 | 905.4954 | 1.715 | 1.716 | 1.7155 | 186.8068 | 0.2063
2020-02-06 | 13:09 | 8.478 | 8.485 | 8.4815 | 904.3758 | 1.715 | 1.717 | 1.716 | 186.8612 | 0.20662
2020-02-06 | 13:09 | 8.459 | 8.475 | 8.467 | 902.8296 | 1.711 | 1.713 | 1.712 | 186.4257 | 0.20649
2020-02-06 | 13:09 | 8.455 | 8.462 | 8.4585 | 901.9233 | 1.711 | 1.712 | 1.7115 | 186.3712 | 0.20664
2020-02-06 | 13:09 | 8.462 | 8.466 | 8.464 | 902.5097 | 1.712 | 1.714 | 1.713 | 186.5346 | 0.20668
2020-02-06 | 13:09 | 8.465 | 8.466 | 8.4655 | 902.6697 | 1.713 | 1.714 | 1.7135 | 186.589 | 0.20671
2020-02-06 | 13:09 | 8.461 | 8.463 | 8.462 | 902.2965 | 1.713 | 1.714 | 1.7135 | 186.589 | 0.20679
2020-02-06 | 13:09 | 8.452 | 8.46 | 8.456 | 901.6567 | 1.712 | 1.713 | 1.7125 | 186.4801 | 0.20682
2020-02-06 | 13:09 | 8.449 | 8.454 | 8.4515 | 901.1769 | 1.71 | 1.714 | 1.712 | 186.4257 | 0.20687
2020-02-06 | 13:09 | 8.426 | 8.445 | 8.4355 | 899.4708 | 1.705 | 1.708 | 1.7065 | 185.8268 | 0.2066
2020-02-06 | 13:09 | 7.952 | 8.418 | 8.185 | 872.7602 | 1.582 | 1.702 | 1.642 | 178.8031 | 0.20487
2020-02-06 | 13:09 | 5.982 | 7.777 | 6.8795 | 733.5557 | 1.163 | 1.548 | 1.3555 | 147.6051 | 0.20122
2020-02-06 | 13:10 | 6.323 | 7.458 | 6.8905 | 734.7287 | 1.26 | 1.502 | 1.381 | 150.3819 | 0.20468
2020-02-06 | 13:10 | 5.216 | 7.127 | 6.1715 | 658.0622 | 0.936 | 1.412 | 1.174 | 127.841 | 0.19427
2020-02-06 | 13:10 | 3.544 | 4.502 | 4.023 | 428.9694 | 0.625 | 0.796 | 0.7105 | 77.3688 | 0.18036
2020-02-06 | 13:10 | 3.406 | 3.503 | 3.4545 | 368.3507 | 0.606 | 0.621 | 0.6135 | 66.8062 | 0.18137
2020-02-06 | 13:10 | 3.349 | 3.392 | 3.3705 | 359.3938 | 0.6 | 0.605 | 0.6025 | 65.6083 | 0.18255
2020-02-06 | 13:10 | 3.325 | 3.342 | 3.3335 | 355.4485 | 0.595 | 0.598 | 0.5965 | 64.955 | 0.18274
2020-02-06 | 13:10 | 3.317 | 3.322 | 3.3195 | 353.9557 | 0.594 | 0.595 | 0.5945 | 64.7372 | 0.1829
2020-02-06 | 13:10 | 3.317 | 3.318 | 3.3175 | 353.7424 | 0.594 | 0.596 | 0.595 | 64.7916 | 0.18316
2020-02-06 | 13:10 | 3.314 | 3.317 | 3.3155 | 353.5292 | 0.594 | 0.595 | 0.5945 | 64.7372 | 0.18312
2020-02-06 | 13:10 | 3.312 | 3.314 | 3.313 | 353.2626 | 0.593 | 0.594 | 0.5935 | 64.6283 | 0.18295
2020-02-06 | 13:10 | 3.313 | 3.317 | 3.315 | 353.4759 | 0.593 | 0.595 | 0.594 | 64.6827 | 0.18299
2020-02-06 | 13:10 | 3.319 | 3.331 | 3.325 | 354.5422 | 0.594 | 0.597 | 0.5955 | 64.8461 | 0.1829
2020-02-06 | 13:10 | 3.334 | 3.385 | 3.3595 | 358.2209 | 0.597 | 0.609 | 0.603 | 65.6628 | 0.1833
2020-02-06 | 13:11 | 3.379 | 3.407 | 3.393 | 361.793 | 0.604 | 0.612 | 0.608 | 66.2072 | 0.183
2020-02-06 | 13:11 | 3.452 | 3.719 | 3.5855 | 382.3191 | 0.626 | 0.686 | 0.656 | 71.4341 | 0.18684
2020-02-06 | 13:11 | 3.441 | 3.53 | 3.4855 | 371.6562 | 0.616 | 0.633 | 0.6245 | 68.004 | 0.18298
2020-02-06 | 13:11 | 3.463 | 3.533 | 3.498 | 372.989 | 0.619 | 0.636 | 0.6275 | 68.3307 | 0.1832
2020-02-06 | 13:11 | 3.445 | 3.453 | 3.449 | 367.7642 | 0.616 | 0.619 | 0.6175 | 67.2417 | 0.18284
2020-02-06 | 13:11 | 3.459 | 3.481 | 3.47 | 370.0034 | 0.62 | 0.626 | 0.623 | 67.8407 | 0.18335
2020-02-06 | 13:11 | 3.486 | 3.514 | 3.5 | 373.2023 | 0.628 | 0.633 | 0.6305 | 68.6574 | 0.18397
2020-02-06 | 13:11 | 3.522 | 3.559 | 3.5405 | 377.5208 | 0.636 | 0.642 | 0.639 | 69.5829 | 0.18432
2020-02-06 | 13:11 | 3.569 | 3.616 | 3.5925 | 383.0655 | 0.644 | 0.654 | 0.649 | 70.6719 | 0.18449
2020-02-06 | 13:11 | 3.632 | 3.72 | 3.676 | 391.969 | 0.656 | 0.676 | 0.666 | 72.5231 | 0.18502
2020-02-06 | 13:11 | 3.832 | 4.102 | 3.967 | 422.9981 | 0.704 | 0.755 | 0.7295 | 79.4378 | 0.1878
2020-02-06 | 13:12 | 3.935 | 4.541 | 4.238 | 451.8946 | 0.715 | 0.85 | 0.7825 | 85.2092 | 0.18856
2020-02-06 | 13:12 | 4.615 | 5.204 | 4.9095 | 523.4962 | 0.878 | 1.006 | 0.942 | 102.5777 | 0.19595
2020-02-06 | 13:12 | 5.416 | 7.922 | 6.669 | 711.1103 | 1.065 | 1.628 | 1.3465 | 146.6251 | 0.20619
2020-02-06 | 13:12 | 7.823 | 8.321 | 8.072 | 860.7111 | 1.562 | 1.692 | 1.627 | 177.1697 | 0.20584
2020-02-06 | 13:12 | 7.223 | 7.563 | 7.393 | 788.3098 | 1.436 | 1.505 | 1.4705 | 160.1279 | 0.20313
2020-02-06 | 13:12 | 6.399 | 7.188 | 6.7935 | 724.3856 | 1.244 | 1.422 | 1.333 | 145.155 | 0.20038
2020-02-06 | 13:12 | 5.801 | 6.2 | 6.0005 | 639.8287 | 1.126 | 1.206 | 1.166 | 126.9698 | 0.19844
2020-02-06 | 13:12 | 5.42 | 6.235 | 5.8275 | 621.3818 | 1.046 | 1.262 | 1.154 | 125.6631 | 0.20223
2020-02-06 | 13:12 | 6.922 | 8.422 | 7.672 | 818.0594 | 1.418 | 1.717 | 1.5675 | 170.6906 | 0.20865
2020-02-06 | 13:12 | 8.482 | 8.562 | 8.522 | 908.6942 | 1.727 | 1.737 | 1.732 | 188.6035 | 0.20755
2020-02-06 | 13:12 | 8.569 | 8.577 | 8.573 | 914.1323 | 1.737 | 1.74 | 1.7385 | 189.3114 | 0.20709
2020-02-06 | 13:12 | 8.52 | 8.563 | 8.5415 | 910.7735 | 1.725 | 1.736 | 1.7305 | 188.4402 | 0.2069
2020-02-06 | 13:12 | 8.507 | 8.52 | 8.5135 | 907.7879 | 1.725 | 1.727 | 1.726 | 187.9502 | 0.20704
2020-02-06 | 13:13 | 8.525 | 8.534 | 8.5295 | 909.494 | 1.728 | 1.73 | 1.729 | 188.2769 | 0.20701
2020-02-06 | 13:13 | 8.527 | 8.542 | 8.5345 | 910.0271 | 1.727 | 1.73 | 1.7285 | 188.2224 | 0.20683
2020-02-06 | 13:13 | 8.456 | 8.523 | 8.4895 | 905.2288 | 1.712 | 1.725 | 1.7185 | 187.1335 | 0.20673
2020-02-06 | 13:13 | 8.462 | 8.542 | 8.502 | 906.5617 | 1.713 | 1.729 | 1.721 | 187.4057 | 0.20672
2020-02-06 | 13:13 | 8.553 | 8.598 | 8.5755 | 914.3989 | 1.731 | 1.737 | 1.734 | 188.8213 | 0.2065
2020-02-06 | 13:13 | 8.61 | 8.638 | 8.624 | 919.5704 | 1.738 | 1.743 | 1.7405 | 189.5291 | 0.20611
2020-02-06 | 13:13 | 8.634 | 8.64 | 8.637 | 920.9566 | 1.742 | 1.745 | 1.7435 | 189.8558 | 0.20615
2020-02-06 | 13:13 | 8.558 | 8.632 | 8.595 | 916.4782 | 1.714 | 1.741 | 1.7275 | 188.1135 | 0.20526
2020-02-06 | 13:13 | 7.542 | 8.401 | 7.9715 | 849.9949 | 1.479 | 1.672 | 1.5755 | 171.5617 | 0.20184
2020-02-06 | 13:13 | 5.884 | 7.332 | 6.608 | 704.6059 | 1.119 | 1.419 | 1.269 | 138.1859 | 0.19612
2020-02-06 | 13:13 | 6.27 | 8.257 | 7.2635 | 774.5014 | 1.246 | 1.683 | 1.4645 | 159.4745 | 0.20591
2020-02-06 | 13:14 | 8.441 | 8.596 | 8.5185 | 908.321 | 1.716 | 1.737 | 1.7265 | 188.0046 | 0.20698
2020-02-06 | 13:14 | 8.489 | 8.598 | 8.5435 | 910.9868 | 1.712 | 1.737 | 1.7245 | 187.7868 | 0.20614
2020-02-06 | 13:14 | 8.474 | 8.517 | 8.4955 | 905.8686 | 1.711 | 1.72 | 1.7155 | 186.8068 | 0.20622
2020-02-06 | 13:14 | 8.318 | 8.517 | 8.4175 | 897.5515 | 1.656 | 1.72 | 1.688 | 183.8122 | 0.20479
2020-02-06 | 13:14 | 6.826 | 8.015 | 7.4205 | 791.2421 | 1.312 | 1.595 | 1.4535 | 158.2767 | 0.20004
2020-02-06 | 13:14 | 5.571 | 6.398 | 5.9845 | 638.1226 | 1.06 | 1.227 | 1.1435 | 124.5197 | 0.19513
2020-02-06 | 13:14 | 4.922 | 5.471 | 5.1965 | 554.0988 | 0.912 | 1.035 | 0.9735 | 106.0078 | 0.19132
2020-02-06 | 13:14 | 4.929 | 6.069 | 5.499 | 586.3541 | 0.922 | 1.189 | 1.0555 | 114.9371 | 0.19602
2020-02-06 | 13:14 | 5.979 | 6.239 | 6.109 | 651.3979 | 1.144 | 1.212 | 1.178 | 128.2765 | 0.19693
2020-02-06 | 13:14 | 5.865 | 6.495 | 6.18 | 658.9686 | 1.128 | 1.287 | 1.2075 | 131.4889 | 0.19954
2020-02-06 | 13:14 | 6.494 | 6.999 | 6.7465 | 719.3741 | 1.239 | 1.383 | 1.311 | 142.7594 | 0.19845
2020-02-06 | 13:14 | 4.875 | 6.084 | 5.4795 | 584.2748 | 0.89 | 1.149 | 1.0195 | 111.0169 | 0.19001
2020-02-06 | 13:14 | 4.737 | 4.914 | 4.8255 | 514.5393 | 0.869 | 0.917 | 0.893 | 97.2419 | 0.18899
2020-02-06 | 13:15 | 4.503 | 4.868 | 4.6855 | 499.6112 | 0.825 | 0.906 | 0.8655 | 94.2473 | 0.18864
2020-02-06 | 13:15 | 4.384 | 4.496 | 4.44 | 473.4337 | 0.799 | 0.822 | 0.8105 | 88.2582 | 0.18642
2020-02-06 | 13:15 | 4.44 | 4.778 | 4.609 | 491.4541 | 0.815 | 0.89 | 0.8525 | 92.8317 | 0.18889
2020-02-06 | 13:15 | 4.793 | 5.305 | 5.049 | 538.3709 | 0.896 | 1.026 | 0.961 | 104.6467 | 0.19438
2020-02-06 | 13:15 | 5.478 | 6.457 | 5.9675 | 636.3099 | 1.072 | 1.279 | 1.1755 | 128.0043 | 0.20117
2020-02-06 | 13:15 | 5.937 | 6.386 | 6.1615 | 656.996 | 1.16 | 1.254 | 1.207 | 131.4345 | 0.20005
2020-02-06 | 13:15 | 5.93 | 6.66 | 6.295 | 671.231 | 1.157 | 1.348 | 1.2525 | 136.3891 | 0.20319
2020-02-06 | 13:15 | 7.04 | 7.909 | 7.4745 | 797.0001 | 1.434 | 1.61 | 1.522 | 165.7359 | 0.20795
2020-02-06 | 13:15 | 7.844 | 7.9 | 7.872 | 839.3852 | 1.594 | 1.606 | 1.6 | 174.2296 | 0.20757
2020-02-06 | 13:15 | 7.879 | 8.144 | 8.0115 | 854.26 | 1.605 | 1.659 | 1.632 | 177.7142 | 0.20803
2020-02-06 | 13:15 | 8.201 | 8.315 | 8.258 | 880.5441 | 1.669 | 1.686 | 1.6775 | 182.6688 | 0.20745
2020-02-06 | 13:16 | 8.289 | 8.331 | 8.31 | 886.0888 | 1.674 | 1.688 | 1.681 | 183.05 | 0.20658
2020-02-06 | 13:16 | 8.038 | 8.198 | 8.118 | 865.616 | 1.624 | 1.653 | 1.6385 | 178.422 | 0.20612
2020-02-06 | 13:16 | 7.633 | 8.021 | 7.827 | 834.5869 | 1.521 | 1.616 | 1.5685 | 170.7995 | 0.20465
2020-02-06 | 13:16 | 7.465 | 7.869 | 7.667 | 817.5263 | 1.493 | 1.587 | 1.54 | 167.696 | 0.20513
2020-02-06 | 13:16 | 7.957 | 8.108 | 8.0325 | 856.4992 | 1.61 | 1.64 | 1.625 | 176.9519 | 0.2066
2020-02-06 | 13:16 | 7.79 | 8.006 | 7.898 | 842.1576 | 1.57 | 1.612 | 1.591 | 173.2496 | 0.20572
Average |  |  |  |  | 763.4172 |  |  |  | 154.8096 | 0.19995
[5 empty] | W/m^2 |  |  |  | W/m^2

## Measurements - Accra
Albedo Measurements
[7 empty] | Cloudless time periods |  | Average Albedo
Site: | Accra, Library | [5 empty] | 12:09pm-2:09pm |  | 0.26861
Location: | Latitude: 5.555164°; Longitude: -0.2045872° | [7 empty] | 0.26861 | Inclination
Date: | 11.02.2020; 12:09pm-2:09pm | [7 empty] | 0.26861 | 0%
|  | Top Pyranometer |  |  |  | Bottom Pyranometer
|  | Min. Value [V] | Max. Value [V] | Average Value [V] | Irradiance [W/m^2] | Min. Value [V] | Max. Value [V] | Average Value [V] | Irradiance [W/m^2] | Albedo | Average Albedo: | 0.26861
2020-02-11 | 12:09 | 6.412 | 6.416 | 6.414 | 683.9198 | 1.666 | 1.666 | 1.666 | 181.4166 | 0.26526
2020-02-11 | 12:09 | 6.415 | 6.418 | 6.4165 | 684.1864 | 1.666 | 1.668 | 1.667 | 181.5255 | 0.26532
2020-02-11 | 12:09 | 6.419 | 6.423 | 6.421 | 684.6662 | 1.667 | 1.669 | 1.668 | 181.6344 | 0.26529
2020-02-11 | 12:09 | 6.423 | 6.427 | 6.425 | 685.0928 | 1.668 | 1.671 | 1.6695 | 181.7977 | 0.26536
2020-02-11 | 12:09 | 6.426 | 6.428 | 6.427 | 685.306 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26552
2020-02-11 | 12:09 | 6.425 | 6.429 | 6.427 | 685.306 | 1.671 | 1.672 | 1.6715 | 182.0155 | 0.2656
2020-02-11 | 12:10 | 6.421 | 6.425 | 6.423 | 684.8795 | 1.67 | 1.67 | 1.67 | 181.8521 | 0.26552
2020-02-11 | 12:10 | 6.417 | 6.419 | 6.418 | 684.3464 | 1.669 | 1.67 | 1.6695 | 181.7977 | 0.26565
2020-02-11 | 12:10 | 6.421 | 6.425 | 6.423 | 684.8795 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.2656
2020-02-11 | 12:10 | 6.426 | 6.432 | 6.429 | 685.5193 | 1.672 | 1.672 | 1.672 | 182.0699 | 0.26559
2020-02-11 | 12:10 | 6.429 | 6.431 | 6.43 | 685.6259 | 1.672 | 1.673 | 1.6725 | 182.1244 | 0.26563
2020-02-11 | 12:10 | 6.429 | 6.43 | 6.4295 | 685.5726 | 1.672 | 1.673 | 1.6725 | 182.1244 | 0.26565
2020-02-11 | 12:10 | 6.422 | 6.431 | 6.4265 | 685.2527 | 1.671 | 1.674 | 1.6725 | 182.1244 | 0.26578
2020-02-11 | 12:10 | 6.416 | 6.419 | 6.4175 | 684.293 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26583
2020-02-11 | 12:10 | 6.412 | 6.415 | 6.4135 | 683.8665 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.26592
2020-02-11 | 12:10 | 6.404 | 6.41 | 6.407 | 683.1734 | 1.668 | 1.669 | 1.6685 | 181.6888 | 0.26595
2020-02-11 | 12:10 | 6.402 | 6.407 | 6.4045 | 682.9069 | 1.667 | 1.67 | 1.6685 | 181.6888 | 0.26605
2020-02-11 | 12:10 | 6.409 | 6.414 | 6.4115 | 683.6533 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.266
2020-02-11 | 12:10 | 6.413 | 6.418 | 6.4155 | 684.0798 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26599
2020-02-11 | 12:11 | 6.412 | 6.416 | 6.414 | 683.9198 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26606
2020-02-11 | 12:11 | 6.412 | 6.417 | 6.4145 | 683.9732 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26604
2020-02-11 | 12:11 | 6.411 | 6.412 | 6.4115 | 683.6533 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.266
2020-02-11 | 12:11 | 6.413 | 6.415 | 6.414 | 683.9198 | 1.671 | 1.672 | 1.6715 | 182.0155 | 0.26614
2020-02-11 | 12:11 | 6.411 | 6.415 | 6.413 | 683.8132 | 1.671 | 1.672 | 1.6715 | 182.0155 | 0.26618
2020-02-11 | 12:11 | 6.41 | 6.412 | 6.411 | 683.5999 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26618
2020-02-11 | 12:11 | 6.414 | 6.419 | 6.4165 | 684.1864 | 1.671 | 1.673 | 1.672 | 182.0699 | 0.26611
2020-02-11 | 12:11 | 6.423 | 6.431 | 6.427 | 685.306 | 1.673 | 1.678 | 1.6755 | 182.4511 | 0.26623
2020-02-11 | 12:11 | 6.436 | 6.444 | 6.44 | 686.6922 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26641
2020-02-11 | 12:11 | 6.444 | 6.448 | 6.446 | 687.332 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26616
2020-02-11 | 12:11 | 6.448 | 6.45 | 6.449 | 687.6519 | 1.681 | 1.682 | 1.6815 | 183.1044 | 0.26627
2020-02-11 | 12:12 | 6.449 | 6.459 | 6.454 | 688.185 | 1.682 | 1.686 | 1.684 | 183.3767 | 0.26646
2020-02-11 | 12:12 | 6.462 | 6.47 | 6.466 | 689.4646 | 1.687 | 1.689 | 1.688 | 183.8122 | 0.2666
2020-02-11 | 12:12 | 6.47 | 6.477 | 6.4735 | 690.2643 | 1.689 | 1.691 | 1.69 | 184.03 | 0.26661
2020-02-11 | 12:12 | 6.478 | 6.48 | 6.479 | 690.8507 | 1.691 | 1.693 | 1.692 | 184.2478 | 0.2667
2020-02-11 | 12:12 | 6.473 | 6.479 | 6.476 | 690.5308 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26666
2020-02-11 | 12:12 | 6.474 | 6.48 | 6.477 | 690.6375 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26694
2020-02-11 | 12:12 | 6.481 | 6.487 | 6.484 | 691.3839 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26696
2020-02-11 | 12:12 | 6.488 | 6.494 | 6.491 | 692.1303 | 1.696 | 1.698 | 1.697 | 184.7923 | 0.26699
2020-02-11 | 12:12 | 6.495 | 6.498 | 6.4965 | 692.7167 | 1.698 | 1.699 | 1.6985 | 184.9556 | 0.267
2020-02-11 | 12:12 | 6.489 | 6.497 | 6.493 | 692.3435 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.26707
2020-02-11 | 12:12 | 6.488 | 6.493 | 6.4905 | 692.077 | 1.695 | 1.696 | 1.6955 | 184.6289 | 0.26678
2020-02-11 | 12:12 | 6.493 | 6.501 | 6.497 | 692.7701 | 1.696 | 1.697 | 1.6965 | 184.7378 | 0.26667
2020-02-11 | 12:12 | 6.502 | 6.504 | 6.503 | 693.4098 | 1.698 | 1.699 | 1.6985 | 184.9556 | 0.26673
2020-02-11 | 12:13 | 6.501 | 6.502 | 6.5015 | 693.2499 | 1.697 | 1.7 | 1.6985 | 184.9556 | 0.2668
2020-02-11 | 12:13 | 6.503 | 6.516 | 6.5095 | 694.1029 | 1.699 | 1.703 | 1.701 | 185.2278 | 0.26686
2020-02-11 | 12:13 | 6.521 | 6.523 | 6.522 | 695.4358 | 1.704 | 1.705 | 1.7045 | 185.609 | 0.2669
2020-02-11 | 12:13 | 6.524 | 6.532 | 6.528 | 696.0756 | 1.705 | 1.708 | 1.7065 | 185.8268 | 0.26696
2020-02-11 | 12:13 | 6.53 | 6.532 | 6.531 | 696.3955 | 1.707 | 1.708 | 1.7075 | 185.9357 | 0.267
2020-02-11 | 12:13 | 6.526 | 6.531 | 6.5285 | 696.1289 | 1.707 | 1.708 | 1.7075 | 185.9357 | 0.2671
2020-02-11 | 12:13 | 6.52 | 6.526 | 6.523 | 695.5424 | 1.705 | 1.707 | 1.706 | 185.7723 | 0.26709
2020-02-11 | 12:13 | 6.511 | 6.518 | 6.5145 | 694.6361 | 1.702 | 1.705 | 1.7035 | 185.5001 | 0.26705
2020-02-11 | 12:13 | 6.499 | 6.507 | 6.503 | 693.4098 | 1.699 | 1.701 | 1.7 | 185.119 | 0.26697
2020-02-11 | 12:13 | 6.495 | 6.498 | 6.4965 | 692.7167 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.26692
2020-02-11 | 12:13 | 6.497 | 6.508 | 6.5025 | 693.3565 | 1.698 | 1.702 | 1.7 | 185.119 | 0.26699
2020-02-11 | 12:14 | 6.509 | 6.523 | 6.516 | 694.796 | 1.703 | 1.707 | 1.705 | 185.6634 | 0.26722
2020-02-11 | 12:14 | 6.522 | 6.525 | 6.5235 | 695.5957 | 1.705 | 1.707 | 1.706 | 185.7723 | 0.26707
2020-02-11 | 12:14 | 6.512 | 6.52 | 6.516 | 694.796 | 1.704 | 1.706 | 1.705 | 185.6634 | 0.26722
2020-02-11 | 12:14 | 6.502 | 6.508 | 6.505 | 693.6231 | 1.7 | 1.703 | 1.7015 | 185.2823 | 0.26712
2020-02-11 | 12:14 | 6.487 | 6.5 | 6.4935 | 692.3969 | 1.697 | 1.7 | 1.6985 | 184.9556 | 0.26712
2020-02-11 | 12:14 | 6.486 | 6.487 | 6.4865 | 691.6505 | 1.696 | 1.698 | 1.697 | 184.7923 | 0.26718
2020-02-11 | 12:14 | 6.476 | 6.485 | 6.4805 | 691.0107 | 1.693 | 1.696 | 1.6945 | 184.52 | 0.26703
2020-02-11 | 12:14 | 6.472 | 6.475 | 6.4735 | 690.2643 | 1.692 | 1.693 | 1.6925 | 184.3023 | 0.267
2020-02-11 | 12:14 | 6.471 | 6.472 | 6.4715 | 690.051 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26716
2020-02-11 | 12:14 | 6.471 | 6.476 | 6.4735 | 690.2643 | 1.691 | 1.694 | 1.6925 | 184.3023 | 0.267
2020-02-11 | 12:14 | 6.476 | 6.487 | 6.4815 | 691.1173 | 1.693 | 1.698 | 1.6955 | 184.6289 | 0.26715
2020-02-11 | 12:14 | 6.483 | 6.501 | 6.492 | 692.2369 | 1.697 | 1.701 | 1.699 | 185.0101 | 0.26726
2020-02-11 | 12:14 | 6.497 | 6.503 | 6.5 | 693.0899 | 1.699 | 1.702 | 1.7005 | 185.1734 | 0.26717
2020-02-11 | 12:15 | 6.491 | 6.496 | 6.4935 | 692.3969 | 1.698 | 1.701 | 1.6995 | 185.0645 | 0.26728
2020-02-11 | 12:15 | 6.483 | 6.49 | 6.4865 | 691.6505 | 1.697 | 1.698 | 1.6975 | 184.8467 | 0.26725
2020-02-11 | 12:15 | 6.482 | 6.484 | 6.483 | 691.2773 | 1.695 | 1.697 | 1.696 | 184.6834 | 0.26716
2020-02-11 | 12:15 | 6.471 | 6.481 | 6.476 | 690.5308 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26729
2020-02-11 | 12:15 | 6.469 | 6.471 | 6.47 | 689.8911 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26723
2020-02-11 | 12:15 | 6.464 | 6.469 | 6.4665 | 689.5179 | 1.691 | 1.693 | 1.692 | 184.2478 | 0.26721
2020-02-11 | 12:15 | 6.465 | 6.467 | 6.466 | 689.4646 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26708
2020-02-11 | 12:15 | 6.459 | 6.465 | 6.462 | 689.038 | 1.689 | 1.691 | 1.69 | 184.03 | 0.26708
2020-02-11 | 12:15 | 6.46 | 6.466 | 6.463 | 689.1447 | 1.689 | 1.692 | 1.6905 | 184.0845 | 0.26712
2020-02-11 | 12:15 | 6.467 | 6.473 | 6.47 | 689.8911 | 1.691 | 1.693 | 1.692 | 184.2478 | 0.26707
2020-02-11 | 12:15 | 6.47 | 6.473 | 6.4715 | 690.051 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26716
2020-02-11 | 12:16 | 6.471 | 6.472 | 6.4715 | 690.051 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26716
2020-02-11 | 12:16 | 6.468 | 6.471 | 6.4695 | 689.8378 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26701
2020-02-11 | 12:16 | 6.463 | 6.467 | 6.465 | 689.3579 | 1.69 | 1.691 | 1.6905 | 184.0845 | 0.26704
2020-02-11 | 12:16 | 6.463 | 6.467 | 6.465 | 689.3579 | 1.69 | 1.691 | 1.6905 | 184.0845 | 0.26704
2020-02-11 | 12:16 | 6.467 | 6.472 | 6.4695 | 689.8378 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26693
2020-02-11 | 12:16 | 6.472 | 6.473 | 6.4725 | 690.1576 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26689
2020-02-11 | 12:16 | 6.471 | 6.473 | 6.472 | 690.1043 | 1.69 | 1.691 | 1.6905 | 184.0845 | 0.26675
2020-02-11 | 12:16 | 6.474 | 6.478 | 6.476 | 690.5308 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26666
2020-02-11 | 12:16 | 6.477 | 6.479 | 6.478 | 690.7441 | 1.693 | 1.695 | 1.694 | 184.4656 | 0.26705
2020-02-11 | 12:16 | 6.479 | 6.487 | 6.483 | 691.2773 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26701
2020-02-11 | 12:16 | 6.487 | 6.49 | 6.4885 | 691.8637 | 1.697 | 1.698 | 1.6975 | 184.8467 | 0.26717
2020-02-11 | 12:16 | 6.489 | 6.492 | 6.4905 | 692.077 | 1.697 | 1.7 | 1.6985 | 184.9556 | 0.26725
2020-02-11 | 12:16 | 6.481 | 6.492 | 6.4865 | 691.6505 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.26733
2020-02-11 | 12:17 | 6.474 | 6.479 | 6.4765 | 690.5842 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26727
2020-02-11 | 12:17 | 6.469 | 6.473 | 6.471 | 689.9977 | 1.694 | 1.695 | 1.6945 | 184.52 | 0.26742
2020-02-11 | 12:17 | 6.468 | 6.471 | 6.4695 | 689.8378 | 1.694 | 1.695 | 1.6945 | 184.52 | 0.26748
2020-02-11 | 12:17 | 6.471 | 6.474 | 6.4725 | 690.1576 | 1.695 | 1.696 | 1.6955 | 184.6289 | 0.26752
2020-02-11 | 12:17 | 6.471 | 6.474 | 6.4725 | 690.1576 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26744
2020-02-11 | 12:17 | 6.474 | 6.482 | 6.478 | 690.7441 | 1.695 | 1.698 | 1.6965 | 184.7378 | 0.26745
2020-02-11 | 12:17 | 6.478 | 6.481 | 6.4795 | 690.9041 | 1.696 | 1.697 | 1.6965 | 184.7378 | 0.26739
2020-02-11 | 12:17 | 6.469 | 6.475 | 6.472 | 690.1043 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26746
2020-02-11 | 12:17 | 6.464 | 6.469 | 6.4665 | 689.5179 | 1.693 | 1.694 | 1.6935 | 184.4111 | 0.26745
2020-02-11 | 12:17 | 6.462 | 6.464 | 6.463 | 689.1447 | 1.692 | 1.693 | 1.6925 | 184.3023 | 0.26744
2020-02-11 | 12:17 | 6.459 | 6.464 | 6.4615 | 688.9847 | 1.692 | 1.693 | 1.6925 | 184.3023 | 0.2675
2020-02-11 | 12:18 | 6.464 | 6.466 | 6.465 | 689.3579 | 1.693 | 1.694 | 1.6935 | 184.4111 | 0.26751
2020-02-11 | 12:18 | 6.459 | 6.462 | 6.4605 | 688.8781 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26762
2020-02-11 | 12:18 | 6.457 | 6.46 | 6.4585 | 688.6648 | 1.692 | 1.693 | 1.6925 | 184.3023 | 0.26762
2020-02-11 | 12:18 | 6.454 | 6.455 | 6.4545 | 688.2383 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26763
2020-02-11 | 12:18 | 6.455 | 6.46 | 6.4575 | 688.5582 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26751
2020-02-11 | 12:18 | 6.452 | 6.457 | 6.4545 | 688.2383 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26755
2020-02-11 | 12:18 | 6.454 | 6.458 | 6.456 | 688.3983 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26749
2020-02-11 | 12:18 | 6.457 | 6.46 | 6.4585 | 688.6648 | 1.69 | 1.692 | 1.691 | 184.1389 | 0.26739
2020-02-11 | 12:18 | 6.457 | 6.461 | 6.459 | 688.7182 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26744
2020-02-11 | 12:18 | 6.46 | 6.462 | 6.461 | 688.9314 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.2676
2020-02-11 | 12:18 | 6.463 | 6.469 | 6.466 | 689.4646 | 1.693 | 1.695 | 1.694 | 184.4656 | 0.26755
2020-02-11 | 12:18 | 6.473 | 6.475 | 6.474 | 690.3176 | 1.696 | 1.697 | 1.6965 | 184.7378 | 0.26761
2020-02-11 | 12:18 | 6.467 | 6.474 | 6.4705 | 689.9444 | 1.694 | 1.697 | 1.6955 | 184.6289 | 0.2676
2020-02-11 | 12:19 | 6.463 | 6.465 | 6.464 | 689.2513 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26747
2020-02-11 | 12:19 | 6.456 | 6.46 | 6.458 | 688.6115 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26749
2020-02-11 | 12:19 | 6.455 | 6.459 | 6.457 | 688.5049 | 1.691 | 1.691 | 1.691 | 184.1389 | 0.26745
2020-02-11 | 12:19 | 6.456 | 6.458 | 6.457 | 688.5049 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26753
2020-02-11 | 12:19 | 6.452 | 6.457 | 6.4545 | 688.2383 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26763
2020-02-11 | 12:19 | 6.445 | 6.452 | 6.4485 | 687.5985 | 1.687 | 1.689 | 1.688 | 183.8122 | 0.26732
2020-02-11 | 12:19 | 6.442 | 6.446 | 6.444 | 687.1187 | 1.687 | 1.688 | 1.6875 | 183.7578 | 0.26743
2020-02-11 | 12:19 | 6.442 | 6.446 | 6.444 | 687.1187 | 1.687 | 1.689 | 1.688 | 183.8122 | 0.26751
2020-02-11 | 12:19 | 6.44 | 6.443 | 6.4415 | 686.8521 | 1.685 | 1.687 | 1.686 | 183.5944 | 0.2673
2020-02-11 | 12:19 | 6.437 | 6.441 | 6.439 | 686.5856 | 1.684 | 1.686 | 1.685 | 183.4856 | 0.26724
2020-02-11 | 12:19 | 6.441 | 6.448 | 6.4445 | 687.172 | 1.685 | 1.688 | 1.6865 | 183.6489 | 0.26725
2020-02-11 | 12:20 | 6.451 | 6.458 | 6.4545 | 688.2383 | 1.685 | 1.689 | 1.687 | 183.7033 | 0.26692
2020-02-11 | 12:20 | 6.457 | 6.46 | 6.4585 | 688.6648 | 1.68 | 1.682 | 1.681 | 183.05 | 0.2658
2020-02-11 | 12:20 | 6.456 | 6.458 | 6.457 | 688.5049 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.26563
2020-02-11 | 12:20 | 6.454 | 6.457 | 6.4555 | 688.3449 | 1.677 | 1.679 | 1.678 | 182.7233 | 0.26545
2020-02-11 | 12:20 | 6.439 | 6.452 | 6.4455 | 687.2787 | 1.674 | 1.676 | 1.675 | 182.3966 | 0.26539
2020-02-11 | 12:20 | 6.437 | 6.438 | 6.4375 | 686.4256 | 1.672 | 1.675 | 1.6735 | 182.2333 | 0.26548
2020-02-11 | 12:20 | 6.43 | 6.436 | 6.433 | 685.9458 | 1.66 | 1.672 | 1.666 | 181.4166 | 0.26448
2020-02-11 | 12:20 | 6.413 | 6.426 | 6.4195 | 684.5063 | 1.658 | 1.659 | 1.6585 | 180.5999 | 0.26384
2020-02-11 | 12:20 | 6.411 | 6.412 | 6.4115 | 683.6533 | 1.661 | 1.666 | 1.6635 | 181.1443 | 0.26497
2020-02-11 | 12:20 | 6.408 | 6.411 | 6.4095 | 683.44 | 1.666 | 1.667 | 1.6665 | 181.471 | 0.26553
2020-02-11 | 12:20 | 6.41 | 6.411 | 6.4105 | 683.5466 | 1.665 | 1.666 | 1.6655 | 181.3621 | 0.26533
2020-02-11 | 12:20 | 6.402 | 6.409 | 6.4055 | 683.0135 | 1.663 | 1.665 | 1.664 | 181.1988 | 0.26529
2020-02-11 | 12:20 | 6.393 | 6.401 | 6.397 | 682.1071 | 1.661 | 1.662 | 1.6615 | 180.9266 | 0.26525
2020-02-11 | 12:21 | 6.384 | 6.392 | 6.388 | 681.1475 | 1.658 | 1.66 | 1.659 | 180.6543 | 0.26522
2020-02-11 | 12:21 | 6.378 | 6.382 | 6.38 | 680.2944 | 1.655 | 1.658 | 1.6565 | 180.3821 | 0.26515
2020-02-11 | 12:21 | 6.376 | 6.378 | 6.377 | 679.9746 | 1.654 | 1.656 | 1.655 | 180.2187 | 0.26504
2020-02-11 | 12:21 | 6.373 | 6.375 | 6.374 | 679.6547 | 1.653 | 1.654 | 1.6535 | 180.0554 | 0.26492
2020-02-11 | 12:21 | 6.367 | 6.375 | 6.371 | 679.3348 | 1.652 | 1.654 | 1.653 | 180.001 | 0.26497
2020-02-11 | 12:21 | 6.362 | 6.367 | 6.3645 | 678.6417 | 1.65 | 1.651 | 1.6505 | 179.7287 | 0.26484
2020-02-11 | 12:21 | 6.359 | 6.363 | 6.361 | 678.2685 | 1.649 | 1.65 | 1.6495 | 179.6198 | 0.26482
2020-02-11 | 12:21 | 6.358 | 6.362 | 6.36 | 678.1619 | 1.651 | 1.652 | 1.6515 | 179.8376 | 0.26518
2020-02-11 | 12:21 | 6.358 | 6.362 | 6.36 | 678.1619 | 1.649 | 1.652 | 1.6505 | 179.7287 | 0.26502
2020-02-11 | 12:21 | 6.356 | 6.357 | 6.3565 | 677.7887 | 1.648 | 1.651 | 1.6495 | 179.6198 | 0.26501
2020-02-11 | 12:21 | 6.355 | 6.361 | 6.358 | 677.9486 | 1.65 | 1.651 | 1.6505 | 179.7287 | 0.26511
2020-02-11 | 12:22 | 6.359 | 6.362 | 6.3605 | 678.2152 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26508
2020-02-11 | 12:22 | 6.352 | 6.357 | 6.3545 | 677.5754 | 1.649 | 1.652 | 1.6505 | 179.7287 | 0.26525
2020-02-11 | 12:22 | 6.347 | 6.352 | 6.3495 | 677.0423 | 1.648 | 1.649 | 1.6485 | 179.5109 | 0.26514
2020-02-11 | 12:22 | 6.344 | 6.347 | 6.3455 | 676.6157 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26507
2020-02-11 | 12:22 | 6.336 | 6.342 | 6.339 | 675.9226 | 1.644 | 1.646 | 1.645 | 179.1298 | 0.26502
2020-02-11 | 12:22 | 6.334 | 6.338 | 6.336 | 675.6028 | 1.643 | 1.646 | 1.6445 | 179.0754 | 0.26506
2020-02-11 | 12:22 | 6.317 | 6.332 | 6.3245 | 674.3765 | 1.639 | 1.643 | 1.641 | 178.6942 | 0.26498
2020-02-11 | 12:22 | 6.308 | 6.316 | 6.312 | 673.0437 | 1.637 | 1.64 | 1.6385 | 178.422 | 0.2651
2020-02-11 | 12:22 | 6.297 | 6.305 | 6.301 | 671.8707 | 1.633 | 1.638 | 1.6355 | 178.0953 | 0.26507
2020-02-11 | 12:22 | 6.288 | 6.295 | 6.2915 | 670.8578 | 1.63 | 1.632 | 1.631 | 177.6053 | 0.26474
2020-02-11 | 12:22 | 6.277 | 6.285 | 6.281 | 669.7381 | 1.626 | 1.628 | 1.627 | 177.1697 | 0.26454
2020-02-11 | 12:22 | 6.263 | 6.274 | 6.2685 | 668.4053 | 1.623 | 1.626 | 1.6245 | 176.8975 | 0.26466
2020-02-11 | 12:22 | 6.25 | 6.263 | 6.2565 | 667.1257 | 1.619 | 1.624 | 1.6215 | 176.5708 | 0.26467
2020-02-11 | 12:23 | 6.239 | 6.244 | 6.2415 | 665.5263 | 1.617 | 1.618 | 1.6175 | 176.1352 | 0.26466
2020-02-11 | 12:23 | 6.24 | 6.243 | 6.2415 | 665.5263 | 1.617 | 1.619 | 1.618 | 176.1897 | 0.26474
2020-02-11 | 12:23 | 6.231 | 6.237 | 6.234 | 664.7266 | 1.615 | 1.618 | 1.6165 | 176.0263 | 0.26481
2020-02-11 | 12:23 | 6.213 | 6.227 | 6.22 | 663.2338 | 1.605 | 1.613 | 1.609 | 175.2096 | 0.26417
2020-02-11 | 12:23 | 6.198 | 6.208 | 6.203 | 661.4211 | 1.602 | 1.604 | 1.603 | 174.5563 | 0.26391
2020-02-11 | 12:23 | 6.196 | 6.2 | 6.198 | 660.8879 | 1.602 | 1.603 | 1.6025 | 174.5018 | 0.26404
2020-02-11 | 12:23 | 6.179 | 6.194 | 6.1865 | 659.6617 | 1.599 | 1.604 | 1.6015 | 174.3929 | 0.26437
2020-02-11 | 12:23 | 6.165 | 6.174 | 6.1695 | 657.849 | 1.595 | 1.599 | 1.597 | 173.9029 | 0.26435
2020-02-11 | 12:23 | 6.158 | 6.161 | 6.1595 | 656.7827 | 1.593 | 1.595 | 1.594 | 173.5762 | 0.26428
2020-02-11 | 12:23 | 6.153 | 6.157 | 6.155 | 656.3029 | 1.594 | 1.602 | 1.598 | 174.0118 | 0.26514
2020-02-11 | 12:23 | 6.148 | 6.151 | 6.1495 | 655.7164 | 1.604 | 1.604 | 1.604 | 174.6652 | 0.26637
2020-02-11 | 12:24 | 6.147 | 6.149 | 6.148 | 655.5565 | 1.603 | 1.605 | 1.604 | 174.6652 | 0.26644
2020-02-11 | 12:24 | 6.142 | 6.146 | 6.144 | 655.1299 | 1.602 | 1.604 | 1.603 | 174.5563 | 0.26645
2020-02-11 | 12:24 | 6.144 | 6.155 | 6.1495 | 655.7164 | 1.601 | 1.605 | 1.603 | 174.5563 | 0.26621
2020-02-11 | 12:24 | 6.157 | 6.162 | 6.1595 | 656.7827 | 1.606 | 1.607 | 1.6065 | 174.9374 | 0.26636
2020-02-11 | 12:24 | 6.164 | 6.171 | 6.1675 | 657.6357 | 1.607 | 1.611 | 1.609 | 175.2096 | 0.26642
2020-02-11 | 12:24 | 6.171 | 6.172 | 6.1715 | 658.0622 | 1.61 | 1.612 | 1.611 | 175.4274 | 0.26658
2020-02-11 | 12:24 | 6.172 | 6.18 | 6.176 | 658.5421 | 1.61 | 1.613 | 1.6115 | 175.4819 | 0.26647
2020-02-11 | 12:24 | 6.184 | 6.192 | 6.188 | 659.8216 | 1.614 | 1.617 | 1.6155 | 175.9175 | 0.26661
2020-02-11 | 12:24 | 6.193 | 6.196 | 6.1945 | 660.5147 | 1.616 | 1.618 | 1.617 | 176.0808 | 0.26658
2020-02-11 | 12:24 | 6.187 | 6.193 | 6.19 | 660.0349 | 1.615 | 1.617 | 1.616 | 175.9719 | 0.26661
2020-02-11 | 12:24 | 6.186 | 6.19 | 6.188 | 659.8216 | 1.613 | 1.616 | 1.6145 | 175.8086 | 0.26645
2020-02-11 | 12:24 | 6.191 | 6.197 | 6.194 | 660.4614 | 1.616 | 1.617 | 1.6165 | 176.0263 | 0.26652
2020-02-11 | 12:24 | 6.198 | 6.208 | 6.203 | 661.4211 | 1.617 | 1.62 | 1.6185 | 176.2441 | 0.26646
2020-02-11 | 12:25 | 6.208 | 6.22 | 6.214 | 662.594 | 1.621 | 1.622 | 1.6215 | 176.5708 | 0.26648
2020-02-11 | 12:25 | 6.22 | 6.222 | 6.221 | 663.3404 | 1.622 | 1.624 | 1.623 | 176.7342 | 0.26643
2020-02-11 | 12:25 | 6.223 | 6.224 | 6.2235 | 663.607 | 1.621 | 1.624 | 1.6225 | 176.6797 | 0.26624
2020-02-11 | 12:25 | 6.224 | 6.227 | 6.2255 | 663.8202 | 1.623 | 1.625 | 1.624 | 176.843 | 0.2664
2020-02-11 | 12:25 | 6.225 | 6.226 | 6.2255 | 663.8202 | 1.621 | 1.625 | 1.623 | 176.7342 | 0.26624
2020-02-11 | 12:25 | 6.221 | 6.223 | 6.222 | 663.447 | 1.615 | 1.62 | 1.6175 | 176.1352 | 0.26549
2020-02-11 | 12:25 | 6.223 | 6.231 | 6.227 | 663.9802 | 1.61 | 1.613 | 1.6115 | 175.4819 | 0.26429
2020-02-11 | 12:25 | 6.227 | 6.233 | 6.23 | 664.3001 | 1.61 | 1.612 | 1.611 | 175.4274 | 0.26408
2020-02-11 | 12:25 | 6.212 | 6.224 | 6.218 | 663.0205 | 1.593 | 1.608 | 1.6005 | 174.284 | 0.26286
2020-02-11 | 12:25 | 6.211 | 6.212 | 6.2115 | 662.3274 | 1.59 | 1.606 | 1.598 | 174.0118 | 0.26273
2020-02-11 | 12:25 | 6.21 | 6.214 | 6.212 | 662.3807 | 1.608 | 1.612 | 1.61 | 175.3185 | 0.26468
2020-02-11 | 12:26 | 6.215 | 6.222 | 6.2185 | 663.0738 | 1.612 | 1.616 | 1.614 | 175.7541 | 0.26506
2020-02-11 | 12:26 | 6.217 | 6.22 | 6.2185 | 663.0738 | 1.614 | 1.618 | 1.616 | 175.9719 | 0.26539
2020-02-11 | 12:26 | 6.217 | 6.219 | 6.218 | 663.0205 | 1.616 | 1.618 | 1.617 | 176.0808 | 0.26557
2020-02-11 | 12:26 | 6.218 | 6.234 | 6.226 | 663.8735 | 1.617 | 1.619 | 1.618 | 176.1897 | 0.2654
2020-02-11 | 12:26 | 6.239 | 6.248 | 6.2435 | 665.7396 | 1.601 | 1.623 | 1.612 | 175.5363 | 0.26367
2020-02-11 | 12:26 | 6.248 | 6.262 | 6.255 | 666.9658 | 1.598 | 1.604 | 1.601 | 174.3385 | 0.26139
2020-02-11 | 12:26 | 6.266 | 6.283 | 6.2745 | 669.0451 | 1.605 | 1.614 | 1.6095 | 175.2641 | 0.26196
2020-02-11 | 12:26 | 6.287 | 6.301 | 6.294 | 671.1243 | 1.615 | 1.619 | 1.617 | 176.0808 | 0.26237
2020-02-11 | 12:26 | 6.304 | 6.313 | 6.3085 | 672.6705 | 1.621 | 1.623 | 1.622 | 176.6253 | 0.26257
2020-02-11 | 12:26 | 6.316 | 6.322 | 6.319 | 673.7901 | 1.623 | 1.627 | 1.625 | 176.9519 | 0.26262
2020-02-11 | 12:26 | 6.323 | 6.337 | 6.33 | 674.963 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26233
2020-02-11 | 12:26 | 6.339 | 6.343 | 6.341 | 676.1359 | 1.627 | 1.63 | 1.6285 | 177.3331 | 0.26227
2020-02-11 | 12:26 | 6.345 | 6.352 | 6.3485 | 676.9356 | 1.629 | 1.632 | 1.6305 | 177.5509 | 0.26229
2020-02-11 | 12:27 | 6.354 | 6.363 | 6.3585 | 678.0019 | 1.635 | 1.639 | 1.637 | 178.2587 | 0.26292
2020-02-11 | 12:27 | 6.365 | 6.371 | 6.368 | 679.0149 | 1.638 | 1.639 | 1.6385 | 178.422 | 0.26277
2020-02-11 | 12:27 | 6.364 | 6.368 | 6.366 | 678.8016 | 1.637 | 1.638 | 1.6375 | 178.3131 | 0.26269
2020-02-11 | 12:27 | 6.359 | 6.363 | 6.361 | 678.2685 | 1.633 | 1.636 | 1.6345 | 177.9864 | 0.26241
2020-02-11 | 12:27 | 6.346 | 6.354 | 6.35 | 677.0956 | 1.626 | 1.63 | 1.628 | 177.2786 | 0.26182
2020-02-11 | 12:27 | 6.344 | 6.346 | 6.345 | 676.5624 | 1.624 | 1.627 | 1.6255 | 177.0064 | 0.26163
2020-02-11 | 12:27 | 6.346 | 6.352 | 6.349 | 676.9889 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26162
2020-02-11 | 12:27 | 6.346 | 6.35 | 6.348 | 676.8823 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.2619
2020-02-11 | 12:27 | 6.349 | 6.358 | 6.3535 | 677.4688 | 1.629 | 1.632 | 1.6305 | 177.5509 | 0.26208
2020-02-11 | 12:27 | 6.353 | 6.356 | 6.3545 | 677.5754 | 1.631 | 1.633 | 1.632 | 177.7142 | 0.26228
2020-02-11 | 12:27 | 6.342 | 6.35 | 6.346 | 676.669 | 1.631 | 1.633 | 1.632 | 177.7142 | 0.26263
2020-02-11 | 12:28 | 6.336 | 6.341 | 6.3385 | 675.8693 | 1.633 | 1.634 | 1.6335 | 177.8775 | 0.26318
2020-02-11 | 12:28 | 6.327 | 6.334 | 6.3305 | 675.0163 | 1.632 | 1.634 | 1.633 | 177.8231 | 0.26344
2020-02-11 | 12:28 | 6.322 | 6.325 | 6.3235 | 674.2699 | 1.629 | 1.632 | 1.6305 | 177.5509 | 0.26332
2020-02-11 | 12:28 | 6.311 | 6.32 | 6.3155 | 673.4169 | 1.628 | 1.63 | 1.629 | 177.3875 | 0.26341
2020-02-11 | 12:28 | 6.305 | 6.309 | 6.307 | 672.5105 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.26353
2020-02-11 | 12:28 | 6.3 | 6.304 | 6.302 | 671.9774 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26382
2020-02-11 | 12:28 | 6.295 | 6.298 | 6.2965 | 671.3909 | 1.626 | 1.628 | 1.627 | 177.1697 | 0.26388
2020-02-11 | 12:28 | 6.29 | 6.295 | 6.2925 | 670.9644 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26397
2020-02-11 | 12:28 | 6.284 | 6.288 | 6.286 | 670.2713 | 1.622 | 1.626 | 1.624 | 176.843 | 0.26384
2020-02-11 | 12:28 | 6.285 | 6.286 | 6.2855 | 670.218 | 1.624 | 1.627 | 1.6255 | 177.0064 | 0.2641
2020-02-11 | 12:28 | 6.282 | 6.29 | 6.286 | 670.2713 | 1.626 | 1.628 | 1.627 | 177.1697 | 0.26433
2020-02-11 | 12:28 | 6.288 | 6.291 | 6.2895 | 670.6445 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26402
2020-02-11 | 12:28 | 6.291 | 6.292 | 6.2915 | 670.8578 | 1.624 | 1.626 | 1.625 | 176.9519 | 0.26377
2020-02-11 | 12:29 | 6.294 | 6.302 | 6.298 | 671.5508 | 1.624 | 1.629 | 1.6265 | 177.1153 | 0.26374
2020-02-11 | 12:29 | 6.305 | 6.321 | 6.313 | 673.1503 | 1.629 | 1.634 | 1.6315 | 177.6597 | 0.26392
2020-02-11 | 12:29 | 6.326 | 6.333 | 6.3295 | 674.9097 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.26404
2020-02-11 | 12:29 | 6.307 | 6.337 | 6.322 | 674.1099 | 1.634 | 1.638 | 1.636 | 178.1498 | 0.26427
2020-02-11 | 12:29 | 6.329 | 6.342 | 6.3355 | 675.5494 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.26379
2020-02-11 | 12:29 | 6.342 | 6.344 | 6.343 | 676.3492 | 1.638 | 1.64 | 1.639 | 178.4764 | 0.26388
2020-02-11 | 12:29 | 6.337 | 6.341 | 6.339 | 675.9226 | 1.635 | 1.649 | 1.642 | 178.8031 | 0.26453
2020-02-11 | 12:29 | 6.339 | 6.344 | 6.3415 | 676.1892 | 1.651 | 1.654 | 1.6525 | 179.9465 | 0.26612
2020-02-11 | 12:29 | 6.324 | 6.344 | 6.334 | 675.3895 | 1.65 | 1.655 | 1.6525 | 179.9465 | 0.26643
2020-02-11 | 12:29 | 6.325 | 6.329 | 6.327 | 674.6431 | 1.649 | 1.651 | 1.65 | 179.6743 | 0.26632
2020-02-11 | 12:29 | 6.314 | 6.321 | 6.3175 | 673.6301 | 1.646 | 1.649 | 1.6475 | 179.402 | 0.26632
2020-02-11 | 12:30 | 6.309 | 6.313 | 6.311 | 672.937 | 1.645 | 1.646 | 1.6455 | 179.1843 | 0.26627
2020-02-11 | 12:30 | 6.308 | 6.312 | 6.31 | 672.8304 | 1.645 | 1.647 | 1.646 | 179.2387 | 0.2664
2020-02-11 | 12:30 | 6.313 | 6.316 | 6.3145 | 673.3102 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26637
2020-02-11 | 12:30 | 6.307 | 6.312 | 6.3095 | 672.7771 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.2665
2020-02-11 | 12:30 | 6.306 | 6.314 | 6.31 | 672.8304 | 1.645 | 1.647 | 1.646 | 179.2387 | 0.2664
2020-02-11 | 12:30 | 6.302 | 6.322 | 6.312 | 673.0437 | 1.645 | 1.651 | 1.648 | 179.4565 | 0.26663
2020-02-11 | 12:30 | 6.326 | 6.334 | 6.33 | 674.963 | 1.651 | 1.653 | 1.652 | 179.8921 | 0.26652
2020-02-11 | 12:30 | 6.332 | 6.335 | 6.3335 | 675.3362 | 1.652 | 1.653 | 1.6525 | 179.9465 | 0.26645
2020-02-11 | 12:30 | 6.327 | 6.331 | 6.329 | 674.8564 | 1.651 | 1.653 | 1.652 | 179.8921 | 0.26656
2020-02-11 | 12:30 | 6.326 | 6.328 | 6.327 | 674.6431 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26649
2020-02-11 | 12:30 | 6.319 | 6.325 | 6.322 | 674.1099 | 1.65 | 1.651 | 1.6505 | 179.7287 | 0.26662
2020-02-11 | 12:30 | 6.303 | 6.316 | 6.3095 | 672.7771 | 1.645 | 1.649 | 1.647 | 179.3476 | 0.26658
2020-02-11 | 12:30 | 6.291 | 6.298 | 6.2945 | 671.1776 | 1.641 | 1.644 | 1.6425 | 178.8576 | 0.26648
2020-02-11 | 12:31 | 6.281 | 6.287 | 6.284 | 670.058 | 1.638 | 1.638 | 1.638 | 178.3676 | 0.2662
2020-02-11 | 12:31 | 6.29 | 6.3 | 6.295 | 671.231 | 1.64 | 1.643 | 1.6415 | 178.7487 | 0.2663
2020-02-11 | 12:31 | 6.3 | 6.303 | 6.3015 | 671.924 | 1.641 | 1.644 | 1.6425 | 178.8576 | 0.26619
2020-02-11 | 12:31 | 6.304 | 6.309 | 6.3065 | 672.4572 | 1.642 | 1.645 | 1.6435 | 178.9665 | 0.26614
2020-02-11 | 12:31 | 6.303 | 6.31 | 6.3065 | 672.4572 | 1.642 | 1.645 | 1.6435 | 178.9665 | 0.26614
2020-02-11 | 12:31 | 6.286 | 6.299 | 6.2925 | 670.9644 | 1.637 | 1.641 | 1.639 | 178.4764 | 0.266
2020-02-11 | 12:31 | 6.282 | 6.284 | 6.283 | 669.9514 | 1.636 | 1.638 | 1.637 | 178.2587 | 0.26608
2020-02-11 | 12:31 | 6.268 | 6.279 | 6.2735 | 668.9384 | 1.63 | 1.635 | 1.6325 | 177.7686 | 0.26575
2020-02-11 | 12:31 | 6.254 | 6.263 | 6.2585 | 667.339 | 1.627 | 1.63 | 1.6285 | 177.3331 | 0.26573
2020-02-11 | 12:31 | 6.25 | 6.253 | 6.2515 | 666.5926 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26562
2020-02-11 | 12:31 | 6.244 | 6.251 | 6.2475 | 666.1661 | 1.624 | 1.626 | 1.625 | 176.9519 | 0.26563
2020-02-11 | 12:32 | 6.245 | 6.253 | 6.249 | 666.326 | 1.624 | 1.627 | 1.6255 | 177.0064 | 0.26565
2020-02-11 | 12:32 | 6.224 | 6.239 | 6.2315 | 664.46 | 1.618 | 1.622 | 1.62 | 176.4075 | 0.26549
2020-02-11 | 12:32 | 6.213 | 6.231 | 6.222 | 663.447 | 1.616 | 1.619 | 1.6175 | 176.1352 | 0.26549
2020-02-11 | 12:32 | 6.237 | 6.256 | 6.2465 | 666.0594 | 1.621 | 1.624 | 1.6225 | 176.6797 | 0.26526
2020-02-11 | 12:32 | 6.259 | 6.268 | 6.2635 | 667.8721 | 1.623 | 1.624 | 1.6235 | 176.7886 | 0.2647
2020-02-11 | 12:32 | 6.269 | 6.276 | 6.2725 | 668.8318 | 1.622 | 1.627 | 1.6245 | 176.8975 | 0.26449
2020-02-11 | 12:32 | 6.278 | 6.293 | 6.2855 | 670.218 | 1.628 | 1.631 | 1.6295 | 177.442 | 0.26475
2020-02-11 | 12:32 | 6.294 | 6.299 | 6.2965 | 671.3909 | 1.631 | 1.633 | 1.632 | 177.7142 | 0.2647
2020-02-11 | 12:32 | 6.287 | 6.299 | 6.293 | 671.0177 | 1.629 | 1.632 | 1.6305 | 177.5509 | 0.2646
2020-02-11 | 12:32 | 6.285 | 6.29 | 6.2875 | 670.4312 | 1.622 | 1.629 | 1.6255 | 177.0064 | 0.26402
2020-02-11 | 12:32 | 6.292 | 6.299 | 6.2955 | 671.2843 | 1.615 | 1.624 | 1.6195 | 176.353 | 0.26271
2020-02-11 | 12:32 | 6.286 | 6.295 | 6.2905 | 670.7511 | 1.625 | 1.626 | 1.6255 | 177.0064 | 0.26389
2020-02-11 | 12:32 | 6.269 | 6.284 | 6.2765 | 669.2583 | 1.62 | 1.626 | 1.623 | 176.7342 | 0.26407
2020-02-11 | 12:33 | 6.261 | 6.267 | 6.264 | 667.9255 | 1.619 | 1.621 | 1.62 | 176.4075 | 0.26411
2020-02-11 | 12:33 | 6.26 | 6.264 | 6.262 | 667.7122 | 1.618 | 1.619 | 1.6185 | 176.2441 | 0.26395
2020-02-11 | 12:33 | 6.258 | 6.265 | 6.2615 | 667.6589 | 1.618 | 1.623 | 1.6205 | 176.4619 | 0.2643
2020-02-11 | 12:33 | 6.266 | 6.278 | 6.272 | 668.7785 | 1.626 | 1.631 | 1.6285 | 177.3331 | 0.26516
2020-02-11 | 12:33 | 6.282 | 6.292 | 6.287 | 670.3779 | 1.63 | 1.634 | 1.632 | 177.7142 | 0.2651
2020-02-11 | 12:33 | 6.29 | 6.292 | 6.291 | 670.8044 | 1.633 | 1.635 | 1.634 | 177.932 | 0.26525
2020-02-11 | 12:33 | 6.294 | 6.305 | 6.2995 | 671.7108 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.2653
2020-02-11 | 12:33 | 6.307 | 6.315 | 6.311 | 672.937 | 1.639 | 1.643 | 1.641 | 178.6942 | 0.26554
2020-02-11 | 12:33 | 6.314 | 6.316 | 6.315 | 673.3635 | 1.642 | 1.643 | 1.6425 | 178.8576 | 0.26562
2020-02-11 | 12:33 | 6.298 | 6.311 | 6.3045 | 672.2439 | 1.636 | 1.64 | 1.638 | 178.3676 | 0.26533
2020-02-11 | 12:33 | 6.28 | 6.292 | 6.286 | 670.2713 | 1.632 | 1.635 | 1.6335 | 177.8775 | 0.26538
2020-02-11 | 12:34 | 6.278 | 6.282 | 6.28 | 669.6315 | 1.631 | 1.633 | 1.632 | 177.7142 | 0.26539
2020-02-11 | 12:34 | 6.285 | 6.294 | 6.2895 | 670.6445 | 1.633 | 1.636 | 1.6345 | 177.9864 | 0.2654
2020-02-11 | 12:34 | 6.295 | 6.31 | 6.3025 | 672.0307 | 1.636 | 1.64 | 1.638 | 178.3676 | 0.26542
2020-02-11 | 12:34 | 6.315 | 6.32 | 6.3175 | 673.6301 | 1.641 | 1.643 | 1.642 | 178.8031 | 0.26543
2020-02-11 | 12:34 | 6.313 | 6.317 | 6.315 | 673.3635 | 1.641 | 1.642 | 1.6415 | 178.7487 | 0.26546
2020-02-11 | 12:34 | 6.321 | 6.33 | 6.3255 | 674.4831 | 1.643 | 1.645 | 1.644 | 179.0209 | 0.26542
2020-02-11 | 12:34 | 6.333 | 6.343 | 6.338 | 675.816 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26554
2020-02-11 | 12:34 | 6.335 | 6.347 | 6.341 | 676.1359 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26541
2020-02-11 | 12:34 | 6.319 | 6.331 | 6.325 | 674.4298 | 1.642 | 1.646 | 1.644 | 179.0209 | 0.26544
2020-02-11 | 12:34 | 6.31 | 6.316 | 6.313 | 673.1503 | 1.64 | 1.642 | 1.641 | 178.6942 | 0.26546
2020-02-11 | 12:34 | 6.318 | 6.322 | 6.32 | 673.8967 | 1.643 | 1.644 | 1.6435 | 178.9665 | 0.26557
2020-02-11 | 12:34 | 6.321 | 6.323 | 6.322 | 674.1099 | 1.643 | 1.645 | 1.644 | 179.0209 | 0.26557
2020-02-11 | 12:34 | 6.323 | 6.328 | 6.3255 | 674.4831 | 1.644 | 1.647 | 1.6455 | 179.1843 | 0.26566
2020-02-11 | 12:35 | 6.327 | 6.331 | 6.329 | 674.8564 | 1.646 | 1.65 | 1.648 | 179.4565 | 0.26592
2020-02-11 | 12:35 | 6.331 | 6.34 | 6.3355 | 675.5494 | 1.649 | 1.652 | 1.6505 | 179.7287 | 0.26605
2020-02-11 | 12:35 | 6.346 | 6.364 | 6.355 | 677.6287 | 1.653 | 1.659 | 1.656 | 180.3276 | 0.26612
2020-02-11 | 12:35 | 6.366 | 6.376 | 6.371 | 679.3348 | 1.658 | 1.661 | 1.6595 | 180.7088 | 0.26601
2020-02-11 | 12:35 | 6.38 | 6.401 | 6.3905 | 681.414 | 1.661 | 1.667 | 1.664 | 181.1988 | 0.26592
2020-02-11 | 12:35 | 6.405 | 6.406 | 6.4055 | 683.0135 | 1.667 | 1.669 | 1.668 | 181.6344 | 0.26593
2020-02-11 | 12:35 | 6.407 | 6.412 | 6.4095 | 683.44 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.26608
2020-02-11 | 12:35 | 6.406 | 6.41 | 6.408 | 683.2801 | 1.667 | 1.669 | 1.668 | 181.6344 | 0.26583
2020-02-11 | 12:35 | 6.408 | 6.411 | 6.4095 | 683.44 | 1.668 | 1.67 | 1.669 | 181.7433 | 0.26592
2020-02-11 | 12:35 | 6.412 | 6.419 | 6.4155 | 684.0798 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26591
2020-02-11 | 12:35 | 6.411 | 6.419 | 6.415 | 684.0265 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26601
2020-02-11 | 12:36 | 6.409 | 6.411 | 6.41 | 683.4933 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26614
2020-02-11 | 12:36 | 6.414 | 6.422 | 6.418 | 684.3464 | 1.671 | 1.673 | 1.672 | 182.0699 | 0.26605
2020-02-11 | 12:36 | 6.423 | 6.432 | 6.4275 | 685.3593 | 1.674 | 1.675 | 1.6745 | 182.3422 | 0.26605
2020-02-11 | 12:36 | 6.429 | 6.43 | 6.4295 | 685.5726 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.26613
2020-02-11 | 12:36 | 6.423 | 6.426 | 6.4245 | 685.0394 | 1.672 | 1.675 | 1.6735 | 182.2333 | 0.26602
2020-02-11 | 12:36 | 6.411 | 6.422 | 6.4165 | 684.1864 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26587
2020-02-11 | 12:36 | 6.402 | 6.409 | 6.4055 | 683.0135 | 1.668 | 1.669 | 1.6685 | 181.6888 | 0.26601
2020-02-11 | 12:36 | 6.403 | 6.411 | 6.407 | 683.1734 | 1.668 | 1.67 | 1.669 | 181.7433 | 0.26603
2020-02-11 | 12:36 | 6.401 | 6.413 | 6.407 | 683.1734 | 1.668 | 1.671 | 1.6695 | 181.7977 | 0.26611
2020-02-11 | 12:36 | 6.399 | 6.404 | 6.4015 | 682.587 | 1.667 | 1.668 | 1.6675 | 181.5799 | 0.26602
2020-02-11 | 12:36 | 6.398 | 6.404 | 6.401 | 682.5337 | 1.665 | 1.667 | 1.666 | 181.4166 | 0.2658
2020-02-11 | 12:36 | 6.398 | 6.402 | 6.4 | 682.427 | 1.666 | 1.667 | 1.6665 | 181.471 | 0.26592
2020-02-11 | 12:36 | 6.401 | 6.405 | 6.403 | 682.7469 | 1.667 | 1.669 | 1.668 | 181.6344 | 0.26603
2020-02-11 | 12:37 | 6.407 | 6.41 | 6.4085 | 683.3334 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.26613
2020-02-11 | 12:37 | 6.399 | 6.408 | 6.4035 | 682.8002 | 1.668 | 1.672 | 1.67 | 181.8521 | 0.26633
2020-02-11 | 12:37 | 6.4 | 6.406 | 6.403 | 682.7469 | 1.668 | 1.671 | 1.6695 | 181.7977 | 0.26627
2020-02-11 | 12:37 | 6.39 | 6.404 | 6.397 | 682.1071 | 1.667 | 1.67 | 1.6685 | 181.6888 | 0.26636
2020-02-11 | 12:37 | 6.389 | 6.392 | 6.3905 | 681.414 | 1.665 | 1.666 | 1.6655 | 181.3621 | 0.26616
2020-02-11 | 12:37 | 6.392 | 6.395 | 6.3935 | 681.7339 | 1.666 | 1.668 | 1.667 | 181.5255 | 0.26627
2020-02-11 | 12:37 | 6.397 | 6.409 | 6.403 | 682.7469 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26643
2020-02-11 | 12:37 | 6.394 | 6.41 | 6.402 | 682.6403 | 1.668 | 1.671 | 1.6695 | 181.7977 | 0.26632
2020-02-11 | 12:37 | 6.392 | 6.398 | 6.395 | 681.8939 | 1.667 | 1.669 | 1.668 | 181.6344 | 0.26637
2020-02-11 | 12:37 | 6.395 | 6.4 | 6.3975 | 682.1605 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.26618
2020-02-11 | 12:37 | 6.397 | 6.404 | 6.4005 | 682.4803 | 1.668 | 1.67 | 1.669 | 181.7433 | 0.2663
2020-02-11 | 12:38 | 6.392 | 6.399 | 6.3955 | 681.9472 | 1.665 | 1.669 | 1.667 | 181.5255 | 0.26619
2020-02-11 | 12:38 | 6.393 | 6.401 | 6.397 | 682.1071 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.2662
2020-02-11 | 12:38 | 6.403 | 6.418 | 6.4105 | 683.5466 | 1.67 | 1.673 | 1.6715 | 182.0155 | 0.26628
2020-02-11 | 12:38 | 6.42 | 6.422 | 6.421 | 684.6662 | 1.674 | 1.675 | 1.6745 | 182.3422 | 0.26632
2020-02-11 | 12:38 | 6.422 | 6.427 | 6.4245 | 685.0394 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.26634
2020-02-11 | 12:38 | 6.425 | 6.426 | 6.4255 | 685.1461 | 1.674 | 1.676 | 1.675 | 182.3966 | 0.26622
2020-02-11 | 12:38 | 6.427 | 6.436 | 6.4315 | 685.7858 | 1.675 | 1.678 | 1.6765 | 182.56 | 0.26621
2020-02-11 | 12:38 | 6.437 | 6.439 | 6.438 | 686.4789 | 1.678 | 1.679 | 1.6785 | 182.7777 | 0.26625
2020-02-11 | 12:38 | 6.438 | 6.44 | 6.439 | 686.5856 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.26637
2020-02-11 | 12:38 | 6.422 | 6.435 | 6.4285 | 685.466 | 1.675 | 1.679 | 1.677 | 182.6144 | 0.26641
2020-02-11 | 12:38 | 6.42 | 6.423 | 6.4215 | 684.7196 | 1.674 | 1.676 | 1.675 | 182.3966 | 0.26638
2020-02-11 | 12:38 | 6.423 | 6.427 | 6.425 | 685.0928 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.26632
2020-02-11 | 12:39 | 6.423 | 6.429 | 6.426 | 685.1994 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.26627
2020-02-11 | 12:39 | 6.414 | 6.421 | 6.4175 | 684.293 | 1.671 | 1.674 | 1.6725 | 182.1244 | 0.26615
2020-02-11 | 12:39 | 6.407 | 6.414 | 6.4105 | 683.5466 | 1.67 | 1.672 | 1.671 | 181.961 | 0.2662
2020-02-11 | 12:39 | 6.406 | 6.409 | 6.4075 | 683.2267 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.26617
2020-02-11 | 12:39 | 6.404 | 6.405 | 6.4045 | 682.9069 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26637
2020-02-11 | 12:39 | 6.404 | 6.409 | 6.4065 | 683.1201 | 1.671 | 1.672 | 1.6715 | 182.0155 | 0.26645
2020-02-11 | 12:39 | 6.412 | 6.415 | 6.4135 | 683.8665 | 1.672 | 1.673 | 1.6725 | 182.1244 | 0.26632
2020-02-11 | 12:39 | 6.412 | 6.416 | 6.414 | 683.9198 | 1.672 | 1.673 | 1.6725 | 182.1244 | 0.26629
2020-02-11 | 12:39 | 6.42 | 6.431 | 6.4255 | 685.1461 | 1.675 | 1.678 | 1.6765 | 182.56 | 0.26645
2020-02-11 | 12:39 | 6.415 | 6.429 | 6.422 | 684.7729 | 1.673 | 1.677 | 1.675 | 182.3966 | 0.26636
2020-02-11 | 12:39 | 6.41 | 6.413 | 6.4115 | 683.6533 | 1.671 | 1.673 | 1.672 | 182.0699 | 0.26632
2020-02-11 | 12:39 | 6.411 | 6.423 | 6.417 | 684.2397 | 1.672 | 1.675 | 1.6735 | 182.2333 | 0.26633
2020-02-11 | 12:39 | 6.411 | 6.422 | 6.4165 | 684.1864 | 1.671 | 1.675 | 1.673 | 182.1788 | 0.26627
2020-02-11 | 12:40 | 6.407 | 6.409 | 6.408 | 683.2801 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.26615
2020-02-11 | 12:40 | 6.41 | 6.423 | 6.4165 | 684.1864 | 1.672 | 1.675 | 1.6735 | 182.2333 | 0.26635
2020-02-11 | 12:40 | 6.424 | 6.427 | 6.4255 | 685.1461 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.2663
2020-02-11 | 12:40 | 6.418 | 6.423 | 6.4205 | 684.6129 | 1.674 | 1.676 | 1.675 | 182.3966 | 0.26642
2020-02-11 | 12:40 | 6.42 | 6.423 | 6.4215 | 684.7196 | 1.674 | 1.675 | 1.6745 | 182.3422 | 0.2663
2020-02-11 | 12:40 | 6.423 | 6.427 | 6.425 | 685.0928 | 1.675 | 1.677 | 1.676 | 182.5055 | 0.2664
2020-02-11 | 12:40 | 6.428 | 6.435 | 6.4315 | 685.7858 | 1.677 | 1.68 | 1.6785 | 182.7777 | 0.26652
2020-02-11 | 12:40 | 6.44 | 6.443 | 6.4415 | 686.8521 | 1.681 | 1.683 | 1.682 | 183.1589 | 0.26666
2020-02-11 | 12:40 | 6.44 | 6.443 | 6.4415 | 686.8521 | 1.682 | 1.683 | 1.6825 | 183.2133 | 0.26674
2020-02-11 | 12:40 | 6.438 | 6.442 | 6.44 | 686.6922 | 1.68 | 1.683 | 1.6815 | 183.1044 | 0.26665
2020-02-11 | 12:40 | 6.434 | 6.439 | 6.4365 | 686.319 | 1.679 | 1.682 | 1.6805 | 182.9955 | 0.26663
2020-02-11 | 12:41 | 6.432 | 6.436 | 6.434 | 686.0524 | 1.68 | 1.68 | 1.68 | 182.9411 | 0.26666
2020-02-11 | 12:41 | 6.432 | 6.438 | 6.435 | 686.159 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26662
2020-02-11 | 12:41 | 6.43 | 6.432 | 6.431 | 685.7325 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.2667
2020-02-11 | 12:41 | 6.417 | 6.429 | 6.423 | 684.8795 | 1.674 | 1.678 | 1.676 | 182.5055 | 0.26648
2020-02-11 | 12:41 | 6.414 | 6.421 | 6.4175 | 684.293 | 1.673 | 1.676 | 1.6745 | 182.3422 | 0.26647
2020-02-11 | 12:41 | 6.421 | 6.426 | 6.4235 | 684.9328 | 1.675 | 1.677 | 1.676 | 182.5055 | 0.26646
2020-02-11 | 12:41 | 6.427 | 6.442 | 6.4345 | 686.1057 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26664
2020-02-11 | 12:41 | 6.439 | 6.441 | 6.44 | 686.6922 | 1.681 | 1.682 | 1.6815 | 183.1044 | 0.26665
2020-02-11 | 12:41 | 6.437 | 6.439 | 6.438 | 686.4789 | 1.681 | 1.682 | 1.6815 | 183.1044 | 0.26673
2020-02-11 | 12:41 | 6.426 | 6.433 | 6.4295 | 685.5726 | 1.677 | 1.68 | 1.6785 | 182.7777 | 0.26661
2020-02-11 | 12:41 | 6.426 | 6.43 | 6.428 | 685.4126 | 1.677 | 1.678 | 1.6775 | 182.6688 | 0.26651
2020-02-11 | 12:41 | 6.425 | 6.43 | 6.4275 | 685.3593 | 1.677 | 1.678 | 1.6775 | 182.6688 | 0.26653
2020-02-11 | 12:41 | 6.432 | 6.442 | 6.437 | 686.3723 | 1.678 | 1.683 | 1.6805 | 182.9955 | 0.26661
2020-02-11 | 12:42 | 6.446 | 6.454 | 6.45 | 687.7585 | 1.683 | 1.686 | 1.6845 | 183.4311 | 0.26671
2020-02-11 | 12:42 | 6.439 | 6.454 | 6.4465 | 687.3853 | 1.683 | 1.685 | 1.684 | 183.3767 | 0.26677
2020-02-11 | 12:42 | 6.425 | 6.434 | 6.4295 | 685.5726 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26684
2020-02-11 | 12:42 | 6.427 | 6.43 | 6.4285 | 685.466 | 1.678 | 1.681 | 1.6795 | 182.8866 | 0.26681
2020-02-11 | 12:42 | 6.421 | 6.426 | 6.4235 | 684.9328 | 1.677 | 1.678 | 1.6775 | 182.6688 | 0.2667
2020-02-11 | 12:42 | 6.419 | 6.421 | 6.42 | 684.5596 | 1.676 | 1.677 | 1.6765 | 182.56 | 0.26668
2020-02-11 | 12:42 | 6.419 | 6.421 | 6.42 | 684.5596 | 1.676 | 1.678 | 1.677 | 182.6144 | 0.26676
2020-02-11 | 12:42 | 6.411 | 6.417 | 6.414 | 683.9198 | 1.673 | 1.677 | 1.675 | 182.3966 | 0.26669
2020-02-11 | 12:42 | 6.406 | 6.411 | 6.4085 | 683.3334 | 1.672 | 1.674 | 1.673 | 182.1788 | 0.2666
2020-02-11 | 12:42 | 6.406 | 6.42 | 6.413 | 683.8132 | 1.673 | 1.678 | 1.6755 | 182.4511 | 0.26681
2020-02-11 | 12:42 | 6.424 | 6.431 | 6.4275 | 685.3593 | 1.678 | 1.68 | 1.679 | 182.8322 | 0.26677
2020-02-11 | 12:43 | 6.433 | 6.435 | 6.434 | 686.0524 | 1.68 | 1.682 | 1.681 | 183.05 | 0.26682
2020-02-11 | 12:43 | 6.432 | 6.435 | 6.4335 | 685.9991 | 1.679 | 1.682 | 1.6805 | 182.9955 | 0.26676
2020-02-11 | 12:43 | 6.431 | 6.434 | 6.4325 | 685.8925 | 1.68 | 1.681 | 1.6805 | 182.9955 | 0.2668
2020-02-11 | 12:43 | 6.432 | 6.437 | 6.4345 | 686.1057 | 1.68 | 1.682 | 1.681 | 183.05 | 0.2668
2020-02-11 | 12:43 | 6.431 | 6.434 | 6.4325 | 685.8925 | 1.678 | 1.683 | 1.6805 | 182.9955 | 0.2668
2020-02-11 | 12:43 | 6.432 | 6.433 | 6.4325 | 685.8925 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26672
2020-02-11 | 12:43 | 6.433 | 6.444 | 6.4385 | 686.5323 | 1.68 | 1.684 | 1.682 | 183.1589 | 0.26679
2020-02-11 | 12:43 | 6.445 | 6.46 | 6.4525 | 688.0251 | 1.684 | 1.687 | 1.6855 | 183.54 | 0.26676
2020-02-11 | 12:43 | 6.462 | 6.474 | 6.468 | 689.6778 | 1.687 | 1.692 | 1.6895 | 183.9756 | 0.26676
2020-02-11 | 12:43 | 6.475 | 6.482 | 6.4785 | 690.7974 | 1.692 | 1.695 | 1.6935 | 184.4111 | 0.26695
2020-02-11 | 12:43 | 6.481 | 6.485 | 6.483 | 691.2773 | 1.693 | 1.697 | 1.695 | 184.5745 | 0.26701
2020-02-11 | 12:43 | 6.488 | 6.493 | 6.4905 | 692.077 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.26717
2020-02-11 | 12:43 | 6.487 | 6.49 | 6.4885 | 691.8637 | 1.696 | 1.699 | 1.6975 | 184.8467 | 0.26717
2020-02-11 | 12:44 | 6.491 | 6.502 | 6.4965 | 692.7167 | 1.698 | 1.7 | 1.699 | 185.0101 | 0.26708
2020-02-11 | 12:44 | 6.488 | 6.498 | 6.493 | 692.3435 | 1.698 | 1.701 | 1.6995 | 185.0645 | 0.2673
2020-02-11 | 12:44 | 6.482 | 6.489 | 6.4855 | 691.5438 | 1.696 | 1.697 | 1.6965 | 184.7378 | 0.26714
2020-02-11 | 12:44 | 6.48 | 6.483 | 6.4815 | 691.1173 | 1.695 | 1.697 | 1.696 | 184.6834 | 0.26722
2020-02-11 | 12:44 | 6.479 | 6.482 | 6.4805 | 691.0107 | 1.694 | 1.695 | 1.6945 | 184.52 | 0.26703
2020-02-11 | 12:44 | 6.48 | 6.481 | 6.4805 | 691.0107 | 1.692 | 1.696 | 1.694 | 184.4656 | 0.26695
2020-02-11 | 12:44 | 6.476 | 6.483 | 6.4795 | 690.9041 | 1.693 | 1.694 | 1.6935 | 184.4111 | 0.26691
2020-02-11 | 12:44 | 6.488 | 6.491 | 6.4895 | 691.9703 | 1.694 | 1.696 | 1.695 | 184.5745 | 0.26674
2020-02-11 | 12:44 | 6.495 | 6.498 | 6.4965 | 692.7167 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.26692
2020-02-11 | 12:44 | 6.5 | 6.504 | 6.502 | 693.3032 | 1.699 | 1.7 | 1.6995 | 185.0645 | 0.26693
2020-02-11 | 12:44 | 6.499 | 6.503 | 6.501 | 693.1966 | 1.698 | 1.699 | 1.6985 | 184.9556 | 0.26682
2020-02-11 | 12:45 | 6.496 | 6.5 | 6.498 | 692.8767 | 1.697 | 1.699 | 1.698 | 184.9012 | 0.26686
2020-02-11 | 12:45 | 6.495 | 6.499 | 6.497 | 692.7701 | 1.697 | 1.698 | 1.6975 | 184.8467 | 0.26682
2020-02-11 | 12:45 | 6.494 | 6.496 | 6.495 | 692.5568 | 1.696 | 1.697 | 1.6965 | 184.7378 | 0.26675
2020-02-11 | 12:45 | 6.495 | 6.497 | 6.496 | 692.6634 | 1.696 | 1.698 | 1.697 | 184.7923 | 0.26679
2020-02-11 | 12:45 | 6.489 | 6.494 | 6.4915 | 692.1836 | 1.695 | 1.697 | 1.696 | 184.6834 | 0.26681
2020-02-11 | 12:45 | 6.48 | 6.486 | 6.483 | 691.2773 | 1.692 | 1.696 | 1.694 | 184.4656 | 0.26685
2020-02-11 | 12:45 | 6.481 | 6.485 | 6.483 | 691.2773 | 1.692 | 1.694 | 1.693 | 184.3567 | 0.26669
2020-02-11 | 12:45 | 6.474 | 6.483 | 6.4785 | 690.7974 | 1.691 | 1.693 | 1.692 | 184.2478 | 0.26672
2020-02-11 | 12:45 | 6.471 | 6.473 | 6.472 | 690.1043 | 1.689 | 1.691 | 1.69 | 184.03 | 0.26667
2020-02-11 | 12:45 | 6.467 | 6.471 | 6.469 | 689.7844 | 1.689 | 1.691 | 1.69 | 184.03 | 0.26679
2020-02-11 | 12:45 | 6.472 | 6.478 | 6.475 | 690.4242 | 1.691 | 1.692 | 1.6915 | 184.1934 | 0.26678
2020-02-11 | 12:45 | 6.469 | 6.477 | 6.473 | 690.211 | 1.689 | 1.691 | 1.69 | 184.03 | 0.26663
2020-02-11 | 12:45 | 6.462 | 6.466 | 6.464 | 689.2513 | 1.687 | 1.689 | 1.688 | 183.8122 | 0.26668
2020-02-11 | 12:46 | 6.46 | 6.461 | 6.4605 | 688.8781 | 1.686 | 1.687 | 1.6865 | 183.6489 | 0.26659
2020-02-11 | 12:46 | 6.462 | 6.468 | 6.465 | 689.3579 | 1.687 | 1.689 | 1.688 | 183.8122 | 0.26664
2020-02-11 | 12:46 | 6.463 | 6.469 | 6.466 | 689.4646 | 1.689 | 1.69 | 1.6895 | 183.9756 | 0.26684
2020-02-11 | 12:46 | 6.456 | 6.464 | 6.46 | 688.8248 | 1.687 | 1.69 | 1.6885 | 183.8667 | 0.26693
2020-02-11 | 12:46 | 6.437 | 6.453 | 6.445 | 687.2253 | 1.684 | 1.688 | 1.686 | 183.5944 | 0.26715
2020-02-11 | 12:46 | 6.437 | 6.442 | 6.4395 | 686.6389 | 1.683 | 1.685 | 1.684 | 183.3767 | 0.26706
2020-02-11 | 12:46 | 6.44 | 6.443 | 6.4415 | 686.8521 | 1.684 | 1.685 | 1.6845 | 183.4311 | 0.26706
2020-02-11 | 12:46 | 6.441 | 6.443 | 6.442 | 686.9055 | 1.684 | 1.686 | 1.685 | 183.4856 | 0.26712
2020-02-11 | 12:46 | 6.438 | 6.443 | 6.4405 | 686.7455 | 1.682 | 1.686 | 1.684 | 183.3767 | 0.26702
2020-02-11 | 12:46 | 6.437 | 6.439 | 6.438 | 686.4789 | 1.684 | 1.686 | 1.685 | 183.4856 | 0.26729
2020-02-11 | 12:46 | 6.435 | 6.437 | 6.436 | 686.2657 | 1.683 | 1.684 | 1.6835 | 183.3222 | 0.26713
2020-02-11 | 12:47 | 6.44 | 6.45 | 6.445 | 687.2253 | 1.684 | 1.689 | 1.6865 | 183.6489 | 0.26723
2020-02-11 | 12:47 | 6.45 | 6.46 | 6.455 | 688.2916 | 1.687 | 1.691 | 1.689 | 183.9211 | 0.26721
2020-02-11 | 12:47 | 6.461 | 6.463 | 6.462 | 689.038 | 1.689 | 1.69 | 1.6895 | 183.9756 | 0.267
2020-02-11 | 12:47 | 6.459 | 6.463 | 6.461 | 688.9314 | 1.688 | 1.69 | 1.689 | 183.9211 | 0.26697
2020-02-11 | 12:47 | 6.452 | 6.458 | 6.455 | 688.2916 | 1.686 | 1.689 | 1.6875 | 183.7578 | 0.26698
2020-02-11 | 12:47 | 6.45 | 6.452 | 6.451 | 687.8651 | 1.686 | 1.687 | 1.6865 | 183.6489 | 0.26698
2020-02-11 | 12:47 | 6.452 | 6.456 | 6.454 | 688.185 | 1.686 | 1.688 | 1.687 | 183.7033 | 0.26694
2020-02-11 | 12:47 | 6.448 | 6.453 | 6.4505 | 687.8118 | 1.687 | 1.688 | 1.6875 | 183.7578 | 0.26716
2020-02-11 | 12:47 | 6.437 | 6.447 | 6.442 | 686.9055 | 1.684 | 1.686 | 1.685 | 183.4856 | 0.26712
2020-02-11 | 12:47 | 6.432 | 6.44 | 6.436 | 686.2657 | 1.683 | 1.686 | 1.6845 | 183.4311 | 0.26729
2020-02-11 | 12:47 | 6.444 | 6.453 | 6.4485 | 687.5985 | 1.686 | 1.688 | 1.687 | 183.7033 | 0.26717
2020-02-11 | 12:47 | 6.442 | 6.452 | 6.447 | 687.4386 | 1.686 | 1.688 | 1.687 | 183.7033 | 0.26723
2020-02-11 | 12:47 | 6.441 | 6.442 | 6.4415 | 686.8521 | 1.684 | 1.686 | 1.685 | 183.4856 | 0.26714
2020-02-11 | 12:48 | 6.44 | 6.442 | 6.441 | 686.7988 | 1.685 | 1.687 | 1.686 | 183.5944 | 0.26732
2020-02-11 | 12:48 | 6.431 | 6.441 | 6.436 | 686.2657 | 1.684 | 1.686 | 1.685 | 183.4856 | 0.26737
2020-02-11 | 12:48 | 6.419 | 6.427 | 6.423 | 684.8795 | 1.68 | 1.683 | 1.6815 | 183.1044 | 0.26735
2020-02-11 | 12:48 | 6.419 | 6.421 | 6.42 | 684.5596 | 1.68 | 1.681 | 1.6805 | 182.9955 | 0.26732
2020-02-11 | 12:48 | 6.413 | 6.417 | 6.415 | 684.0265 | 1.677 | 1.678 | 1.6775 | 182.6688 | 0.26705
2020-02-11 | 12:48 | 6.418 | 6.421 | 6.4195 | 684.5063 | 1.677 | 1.682 | 1.6795 | 182.8866 | 0.26718
2020-02-11 | 12:48 | 6.42 | 6.426 | 6.423 | 684.8795 | 1.68 | 1.682 | 1.681 | 183.05 | 0.26727
2020-02-11 | 12:48 | 6.43 | 6.435 | 6.4325 | 685.8925 | 1.683 | 1.683 | 1.683 | 183.2678 | 0.2672
2020-02-11 | 12:48 | 6.415 | 6.428 | 6.4215 | 684.7196 | 1.678 | 1.682 | 1.68 | 182.9411 | 0.26718
2020-02-11 | 12:48 | 6.399 | 6.413 | 6.406 | 683.0668 | 1.673 | 1.677 | 1.675 | 182.3966 | 0.26703
2020-02-11 | 12:48 | 6.39 | 6.396 | 6.393 | 681.6806 | 1.671 | 1.674 | 1.6725 | 182.1244 | 0.26717
2020-02-11 | 12:49 | 6.385 | 6.387 | 6.386 | 680.9342 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26722
2020-02-11 | 12:49 | 6.384 | 6.394 | 6.389 | 681.2541 | 1.669 | 1.674 | 1.6715 | 182.0155 | 0.26718
2020-02-11 | 12:49 | 6.399 | 6.405 | 6.402 | 682.6403 | 1.675 | 1.677 | 1.676 | 182.5055 | 0.26735
2020-02-11 | 12:49 | 6.406 | 6.411 | 6.4085 | 683.3334 | 1.677 | 1.679 | 1.678 | 182.7233 | 0.2674
2020-02-11 | 12:49 | 6.412 | 6.415 | 6.4135 | 683.8665 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.26743
2020-02-11 | 12:49 | 6.409 | 6.416 | 6.4125 | 683.7599 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.26747
2020-02-11 | 12:49 | 6.405 | 6.409 | 6.407 | 683.1734 | 1.676 | 1.677 | 1.6765 | 182.56 | 0.26722
2020-02-11 | 12:49 | 6.406 | 6.414 | 6.41 | 683.4933 | 1.676 | 1.681 | 1.6785 | 182.7777 | 0.26742
2020-02-11 | 12:49 | 6.418 | 6.428 | 6.423 | 684.8795 | 1.68 | 1.683 | 1.6815 | 183.1044 | 0.26735
2020-02-11 | 12:49 | 6.425 | 6.428 | 6.4265 | 685.2527 | 1.683 | 1.684 | 1.6835 | 183.3222 | 0.26752
2020-02-11 | 12:49 | 6.426 | 6.429 | 6.4275 | 685.3593 | 1.684 | 1.684 | 1.684 | 183.3767 | 0.26756
2020-02-11 | 12:49 | 6.425 | 6.428 | 6.4265 | 685.2527 | 1.683 | 1.685 | 1.684 | 183.3767 | 0.2676
2020-02-11 | 12:49 | 6.413 | 6.427 | 6.42 | 684.5596 | 1.68 | 1.685 | 1.6825 | 183.2133 | 0.26764
2020-02-11 | 12:50 | 6.398 | 6.41 | 6.404 | 682.8535 | 1.676 | 1.68 | 1.678 | 182.7233 | 0.26759
2020-02-11 | 12:50 | 6.394 | 6.398 | 6.396 | 682.0005 | 1.675 | 1.676 | 1.6755 | 182.4511 | 0.26752
2020-02-11 | 12:50 | 6.394 | 6.406 | 6.4 | 682.427 | 1.675 | 1.678 | 1.6765 | 182.56 | 0.26752
2020-02-11 | 12:50 | 6.406 | 6.409 | 6.4075 | 683.2267 | 1.678 | 1.68 | 1.679 | 182.8322 | 0.2676
2020-02-11 | 12:50 | 6.396 | 6.403 | 6.3995 | 682.3737 | 1.675 | 1.679 | 1.677 | 182.6144 | 0.26762
2020-02-11 | 12:50 | 6.395 | 6.402 | 6.3985 | 682.2671 | 1.676 | 1.679 | 1.6775 | 182.6688 | 0.26774
2020-02-11 | 12:50 | 6.402 | 6.405 | 6.4035 | 682.8002 | 1.678 | 1.68 | 1.679 | 182.8322 | 0.26777
2020-02-11 | 12:50 | 6.403 | 6.406 | 6.4045 | 682.9069 | 1.679 | 1.68 | 1.6795 | 182.8866 | 0.26781
2020-02-11 | 12:50 | 6.406 | 6.412 | 6.409 | 683.3867 | 1.68 | 1.681 | 1.6805 | 182.9955 | 0.26778
2020-02-11 | 12:50 | 6.412 | 6.414 | 6.413 | 683.8132 | 1.681 | 1.683 | 1.682 | 183.1589 | 0.26785
2020-02-11 | 12:50 | 6.408 | 6.415 | 6.4115 | 683.6533 | 1.68 | 1.683 | 1.6815 | 183.1044 | 0.26783
2020-02-11 | 12:51 | 6.419 | 6.426 | 6.4225 | 684.8262 | 1.683 | 1.687 | 1.685 | 183.4856 | 0.26793
2020-02-11 | 12:51 | 6.413 | 6.424 | 6.4185 | 684.3997 | 1.682 | 1.685 | 1.6835 | 183.3222 | 0.26786
2020-02-11 | 12:51 | 6.4 | 6.411 | 6.4055 | 683.0135 | 1.677 | 1.681 | 1.679 | 182.8322 | 0.26768
2020-02-11 | 12:51 | 6.389 | 6.399 | 6.394 | 681.7873 | 1.676 | 1.678 | 1.677 | 182.6144 | 0.26785
2020-02-11 | 12:51 | 6.384 | 6.389 | 6.3865 | 680.9875 | 1.674 | 1.676 | 1.675 | 182.3966 | 0.26784
2020-02-11 | 12:51 | 6.37 | 6.381 | 6.3755 | 679.8146 | 1.672 | 1.674 | 1.673 | 182.1788 | 0.26798
2020-02-11 | 12:51 | 6.371 | 6.374 | 6.3725 | 679.4947 | 1.672 | 1.673 | 1.6725 | 182.1244 | 0.26803
2020-02-11 | 12:51 | 6.376 | 6.379 | 6.3775 | 680.0279 | 1.672 | 1.674 | 1.673 | 182.1788 | 0.2679
2020-02-11 | 12:51 | 6.377 | 6.384 | 6.3805 | 680.3478 | 1.672 | 1.676 | 1.674 | 182.2877 | 0.26793
2020-02-11 | 12:51 | 6.384 | 6.392 | 6.388 | 681.1475 | 1.675 | 1.677 | 1.676 | 182.5055 | 0.26794
2020-02-11 | 12:51 | 6.394 | 6.401 | 6.3975 | 682.1605 | 1.677 | 1.68 | 1.6785 | 182.7777 | 0.26794
2020-02-11 | 12:51 | 6.399 | 6.406 | 6.4025 | 682.6936 | 1.679 | 1.681 | 1.68 | 182.9411 | 0.26797
2020-02-11 | 12:51 | 6.395 | 6.406 | 6.4005 | 682.4803 | 1.646 | 1.683 | 1.6645 | 181.2532 | 0.26558
2020-02-11 | 12:52 | 6.386 | 6.392 | 6.389 | 681.2541 | 1.592 | 1.67 | 1.631 | 177.6053 | 0.2607
2020-02-11 | 12:52 | 6.381 | 6.385 | 6.383 | 680.6143 | 1.672 | 1.675 | 1.6735 | 182.2333 | 0.26775
2020-02-11 | 12:52 | 6.371 | 6.378 | 6.3745 | 679.708 | 1.671 | 1.673 | 1.672 | 182.0699 | 0.26786
2020-02-11 | 12:52 | 6.372 | 6.375 | 6.3735 | 679.6014 | 1.671 | 1.672 | 1.6715 | 182.0155 | 0.26783
2020-02-11 | 12:52 | 6.368 | 6.373 | 6.3705 | 679.2815 | 1.67 | 1.672 | 1.671 | 181.961 | 0.26787
2020-02-11 | 12:52 | 6.352 | 6.365 | 6.3585 | 678.0019 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.26782
2020-02-11 | 12:52 | 6.342 | 6.349 | 6.3455 | 676.6157 | 1.662 | 1.665 | 1.6635 | 181.1443 | 0.26772
2020-02-11 | 12:52 | 6.341 | 6.344 | 6.3425 | 676.2958 | 1.662 | 1.663 | 1.6625 | 181.0354 | 0.26769
2020-02-11 | 12:52 | 6.346 | 6.347 | 6.3465 | 676.7224 | 1.663 | 1.664 | 1.6635 | 181.1443 | 0.26768
2020-02-11 | 12:52 | 6.349 | 6.352 | 6.3505 | 677.1489 | 1.665 | 1.666 | 1.6655 | 181.3621 | 0.26783
2020-02-11 | 12:52 | 6.355 | 6.359 | 6.357 | 677.842 | 1.666 | 1.668 | 1.667 | 181.5255 | 0.2678
2020-02-11 | 12:52 | 6.362 | 6.365 | 6.3635 | 678.5351 | 1.667 | 1.67 | 1.6685 | 181.6888 | 0.26777
2020-02-11 | 12:53 | 6.362 | 6.363 | 6.3625 | 678.4284 | 1.667 | 1.668 | 1.6675 | 181.5799 | 0.26765
2020-02-11 | 12:53 | 6.355 | 6.362 | 6.3585 | 678.0019 | 1.664 | 1.666 | 1.665 | 181.3077 | 0.26741
2020-02-11 | 12:53 | 6.352 | 6.355 | 6.3535 | 677.4688 | 1.663 | 1.665 | 1.664 | 181.1988 | 0.26746
2020-02-11 | 12:53 | 6.356 | 6.366 | 6.361 | 678.2685 | 1.665 | 1.669 | 1.667 | 181.5255 | 0.26763
2020-02-11 | 12:53 | 6.368 | 6.378 | 6.373 | 679.548 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26769
2020-02-11 | 12:53 | 6.381 | 6.387 | 6.384 | 680.721 | 1.673 | 1.674 | 1.6735 | 182.2333 | 0.26771
2020-02-11 | 12:53 | 6.383 | 6.387 | 6.385 | 680.8276 | 1.672 | 1.674 | 1.673 | 182.1788 | 0.26758
2020-02-11 | 12:53 | 6.373 | 6.381 | 6.377 | 679.9746 | 1.669 | 1.672 | 1.6705 | 181.9066 | 0.26752
2020-02-11 | 12:53 | 6.369 | 6.371 | 6.37 | 679.2282 | 1.669 | 1.669 | 1.669 | 181.7433 | 0.26757
2020-02-11 | 12:53 | 6.37 | 6.372 | 6.371 | 679.3348 | 1.669 | 1.671 | 1.67 | 181.8521 | 0.26769
2020-02-11 | 12:53 | 6.369 | 6.372 | 6.3705 | 679.2815 | 1.668 | 1.67 | 1.669 | 181.7433 | 0.26755
2020-02-11 | 12:54 | 6.369 | 6.371 | 6.37 | 679.2282 | 1.668 | 1.669 | 1.6685 | 181.6888 | 0.26749
2020-02-11 | 12:54 | 6.354 | 6.367 | 6.3605 | 678.2152 | 1.663 | 1.669 | 1.666 | 181.4166 | 0.26749
2020-02-11 | 12:54 | 6.342 | 6.353 | 6.3475 | 676.829 | 1.66 | 1.663 | 1.6615 | 180.9266 | 0.26732
2020-02-11 | 12:54 | 6.34 | 6.344 | 6.342 | 676.2425 | 1.659 | 1.661 | 1.66 | 180.7632 | 0.26731
2020-02-11 | 12:54 | 6.345 | 6.346 | 6.3455 | 676.6157 | 1.66 | 1.662 | 1.661 | 180.8721 | 0.26732
2020-02-11 | 12:54 | 6.34 | 6.343 | 6.3415 | 676.1892 | 1.659 | 1.66 | 1.6595 | 180.7088 | 0.26725
2020-02-11 | 12:54 | 6.337 | 6.339 | 6.338 | 675.816 | 1.658 | 1.66 | 1.659 | 180.6543 | 0.26731
2020-02-11 | 12:54 | 6.337 | 6.339 | 6.338 | 675.816 | 1.658 | 1.66 | 1.659 | 180.6543 | 0.26731
2020-02-11 | 12:54 | 6.334 | 6.336 | 6.335 | 675.4961 | 1.657 | 1.66 | 1.6585 | 180.5999 | 0.26736
2020-02-11 | 12:54 | 6.334 | 6.342 | 6.338 | 675.816 | 1.658 | 1.661 | 1.6595 | 180.7088 | 0.26739
2020-02-11 | 12:54 | 6.344 | 6.349 | 6.3465 | 676.7224 | 1.661 | 1.662 | 1.6615 | 180.9266 | 0.26736
2020-02-11 | 12:54 | 6.318 | 6.342 | 6.33 | 674.963 | 1.654 | 1.661 | 1.6575 | 180.491 | 0.26741
2020-02-11 | 12:54 | 6.302 | 6.313 | 6.3075 | 672.5638 | 1.648 | 1.652 | 1.65 | 179.6743 | 0.26715
2020-02-11 | 12:55 | 6.299 | 6.303 | 6.301 | 671.8707 | 1.648 | 1.649 | 1.6485 | 179.5109 | 0.26718
2020-02-11 | 12:55 | 6.3 | 6.307 | 6.3035 | 672.1373 | 1.647 | 1.65 | 1.6485 | 179.5109 | 0.26707
2020-02-11 | 12:55 | 6.309 | 6.316 | 6.3125 | 673.097 | 1.651 | 1.652 | 1.6515 | 179.8376 | 0.26718
2020-02-11 | 12:55 | 6.314 | 6.323 | 6.3185 | 673.7367 | 1.651 | 1.655 | 1.653 | 180.001 | 0.26717
2020-02-11 | 12:55 | 6.327 | 6.339 | 6.333 | 675.2829 | 1.655 | 1.66 | 1.6575 | 180.491 | 0.26728
2020-02-11 | 12:55 | 6.344 | 6.356 | 6.35 | 677.0956 | 1.661 | 1.665 | 1.663 | 181.0899 | 0.26745
2020-02-11 | 12:55 | 6.357 | 6.366 | 6.3615 | 678.3218 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.26769
2020-02-11 | 12:55 | 6.359 | 6.366 | 6.3625 | 678.4284 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.26765
2020-02-11 | 12:55 | 6.358 | 6.36 | 6.359 | 678.0552 | 1.665 | 1.667 | 1.666 | 181.4166 | 0.26755
2020-02-11 | 12:55 | 6.356 | 6.361 | 6.3585 | 678.0019 | 1.665 | 1.666 | 1.6655 | 181.3621 | 0.2675
2020-02-11 | 12:55 | 6.342 | 6.352 | 6.347 | 676.7757 | 1.661 | 1.664 | 1.6625 | 181.0354 | 0.2675
2020-02-11 | 12:56 | 6.334 | 6.339 | 6.3365 | 675.6561 | 1.659 | 1.66 | 1.6595 | 180.7088 | 0.26746
2020-02-11 | 12:56 | 6.33 | 6.334 | 6.332 | 675.1762 | 1.657 | 1.659 | 1.658 | 180.5454 | 0.2674
2020-02-11 | 12:56 | 6.327 | 6.33 | 6.3285 | 674.803 | 1.656 | 1.657 | 1.6565 | 180.3821 | 0.26731
2020-02-11 | 12:56 | 6.327 | 6.342 | 6.3345 | 675.4428 | 1.656 | 1.661 | 1.6585 | 180.5999 | 0.26738
2020-02-11 | 12:56 | 6.345 | 6.354 | 6.3495 | 677.0423 | 1.661 | 1.664 | 1.6625 | 181.0354 | 0.26739
2020-02-11 | 12:56 | 6.357 | 6.358 | 6.3575 | 677.8953 | 1.663 | 1.665 | 1.664 | 181.1988 | 0.2673
2020-02-11 | 12:56 | 6.345 | 6.354 | 6.3495 | 677.0423 | 1.661 | 1.663 | 1.662 | 180.981 | 0.26731
2020-02-11 | 12:56 | 6.34 | 6.342 | 6.341 | 676.1359 | 1.659 | 1.66 | 1.6595 | 180.7088 | 0.26727
2020-02-11 | 12:56 | 6.335 | 6.342 | 6.3385 | 675.8693 | 1.658 | 1.659 | 1.6585 | 180.5999 | 0.26721
2020-02-11 | 12:56 | 6.323 | 6.333 | 6.328 | 674.7497 | 1.654 | 1.657 | 1.6555 | 180.2732 | 0.26717
2020-02-11 | 12:56 | 6.316 | 6.32 | 6.318 | 673.6834 | 1.649 | 1.653 | 1.651 | 179.7832 | 0.26687
2020-02-11 | 12:56 | 6.304 | 6.313 | 6.3085 | 672.6705 | 1.647 | 1.65 | 1.6485 | 179.5109 | 0.26686
2020-02-11 | 12:56 | 6.288 | 6.301 | 6.2945 | 671.1776 | 1.642 | 1.647 | 1.6445 | 179.0754 | 0.26681
2020-02-11 | 12:57 | 6.279 | 6.285 | 6.282 | 669.8448 | 1.639 | 1.643 | 1.641 | 178.6942 | 0.26677
2020-02-11 | 12:57 | 6.271 | 6.276 | 6.2735 | 668.9384 | 1.639 | 1.64 | 1.6395 | 178.5309 | 0.26689
2020-02-11 | 12:57 | 6.258 | 6.269 | 6.2635 | 667.8721 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.26682
2020-02-11 | 12:57 | 6.223 | 6.253 | 6.238 | 665.1531 | 1.625 | 1.633 | 1.629 | 177.3875 | 0.26669
2020-02-11 | 12:57 | 6.2 | 6.213 | 6.2065 | 661.7943 | 1.618 | 1.623 | 1.6205 | 176.4619 | 0.26664
2020-02-11 | 12:57 | 6.204 | 6.209 | 6.2065 | 661.7943 | 1.618 | 1.62 | 1.619 | 176.2986 | 0.26639
2020-02-11 | 12:57 | 6.21 | 6.215 | 6.2125 | 662.434 | 1.621 | 1.622 | 1.6215 | 176.5708 | 0.26655
2020-02-11 | 12:57 | 6.214 | 6.22 | 6.217 | 662.9139 | 1.621 | 1.624 | 1.6225 | 176.6797 | 0.26652
2020-02-11 | 12:57 | 6.213 | 6.221 | 6.217 | 662.9139 | 1.622 | 1.624 | 1.623 | 176.7342 | 0.2666
2020-02-11 | 12:57 | 6.209 | 6.211 | 6.21 | 662.1675 | 1.621 | 1.623 | 1.622 | 176.6253 | 0.26674
2020-02-11 | 12:57 | 6.204 | 6.209 | 6.2065 | 661.7943 | 1.62 | 1.621 | 1.6205 | 176.4619 | 0.26664
2020-02-11 | 12:58 | 6.205 | 6.216 | 6.2105 | 662.2208 | 1.62 | 1.624 | 1.622 | 176.6253 | 0.26672
2020-02-11 | 12:58 | 6.22 | 6.23 | 6.225 | 663.7669 | 1.624 | 1.628 | 1.626 | 177.0608 | 0.26675
2020-02-11 | 12:58 | 6.231 | 6.241 | 6.236 | 664.9398 | 1.628 | 1.631 | 1.6295 | 177.442 | 0.26685
2020-02-11 | 12:58 | 6.247 | 6.263 | 6.255 | 666.9658 | 1.632 | 1.635 | 1.6335 | 177.8775 | 0.2667
2020-02-11 | 12:58 | 6.264 | 6.267 | 6.2655 | 668.0854 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.26674
2020-02-11 | 12:58 | 6.267 | 6.278 | 6.2725 | 668.8318 | 1.635 | 1.641 | 1.638 | 178.3676 | 0.26669
2020-02-11 | 12:58 | 6.282 | 6.289 | 6.2855 | 670.218 | 1.641 | 1.644 | 1.6425 | 178.8576 | 0.26686
2020-02-11 | 12:58 | 6.287 | 6.288 | 6.2875 | 670.4312 | 1.643 | 1.644 | 1.6435 | 178.9665 | 0.26694
2020-02-11 | 12:58 | 6.286 | 6.289 | 6.2875 | 670.4312 | 1.643 | 1.645 | 1.644 | 179.0209 | 0.26702
2020-02-11 | 12:58 | 6.288 | 6.304 | 6.296 | 671.3376 | 1.643 | 1.647 | 1.645 | 179.1298 | 0.26683
2020-02-11 | 12:58 | 6.307 | 6.316 | 6.3115 | 672.9903 | 1.648 | 1.651 | 1.6495 | 179.6198 | 0.2669
2020-02-11 | 12:58 | 6.316 | 6.319 | 6.3175 | 673.6301 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26689
2020-02-11 | 12:58 | 6.317 | 6.322 | 6.3195 | 673.8434 | 1.651 | 1.652 | 1.6515 | 179.8376 | 0.26688
2020-02-11 | 12:59 | 6.317 | 6.32 | 6.3185 | 673.7367 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26684
2020-02-11 | 12:59 | 6.319 | 6.322 | 6.3205 | 673.95 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26676
2020-02-11 | 12:59 | 6.31 | 6.32 | 6.315 | 673.3635 | 1.649 | 1.651 | 1.65 | 179.6743 | 0.26683
2020-02-11 | 12:59 | 6.305 | 6.307 | 6.306 | 672.4039 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.26665
2020-02-11 | 12:59 | 6.306 | 6.313 | 6.3095 | 672.7771 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.2665
2020-02-11 | 12:59 | 6.313 | 6.325 | 6.319 | 673.7901 | 1.648 | 1.652 | 1.65 | 179.6743 | 0.26666
2020-02-11 | 12:59 | 6.329 | 6.335 | 6.332 | 675.1762 | 1.653 | 1.655 | 1.654 | 180.1099 | 0.26676
2020-02-11 | 12:59 | 6.331 | 6.335 | 6.333 | 675.2829 | 1.652 | 1.654 | 1.653 | 180.001 | 0.26656
2020-02-11 | 12:59 | 6.33 | 6.332 | 6.331 | 675.0696 | 1.653 | 1.655 | 1.654 | 180.1099 | 0.2668
2020-02-11 | 12:59 | 6.33 | 6.333 | 6.3315 | 675.1229 | 1.654 | 1.656 | 1.655 | 180.2187 | 0.26694
2020-02-11 | 12:59 | 6.334 | 6.338 | 6.336 | 675.6028 | 1.655 | 1.656 | 1.6555 | 180.2732 | 0.26683
2020-02-11 | 13:00 | 6.331 | 6.332 | 6.3315 | 675.1229 | 1.654 | 1.655 | 1.6545 | 180.1643 | 0.26686
2020-02-11 | 13:00 | 6.323 | 6.33 | 6.3265 | 674.5898 | 1.652 | 1.654 | 1.653 | 180.001 | 0.26683
2020-02-11 | 13:00 | 6.315 | 6.319 | 6.317 | 673.5768 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26691
2020-02-11 | 13:00 | 6.315 | 6.317 | 6.316 | 673.4702 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26695
2020-02-11 | 13:00 | 6.313 | 6.316 | 6.3145 | 673.3102 | 1.65 | 1.651 | 1.6505 | 179.7287 | 0.26693
2020-02-11 | 13:00 | 6.313 | 6.317 | 6.315 | 673.3635 | 1.649 | 1.651 | 1.65 | 179.6743 | 0.26683
2020-02-11 | 13:00 | 6.307 | 6.314 | 6.3105 | 672.8837 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.2667
2020-02-11 | 13:00 | 6.305 | 6.307 | 6.306 | 672.4039 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26689
2020-02-11 | 13:00 | 6.302 | 6.306 | 6.304 | 672.1906 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26681
2020-02-11 | 13:00 | 6.301 | 6.305 | 6.303 | 672.084 | 1.645 | 1.646 | 1.6455 | 179.1843 | 0.26661
2020-02-11 | 13:00 | 6.294 | 6.305 | 6.2995 | 671.7108 | 1.645 | 1.645 | 1.645 | 179.1298 | 0.26668
2020-02-11 | 13:00 | 6.271 | 6.291 | 6.281 | 669.7381 | 1.638 | 1.642 | 1.64 | 178.5853 | 0.26665
2020-02-11 | 13:00 | 6.232 | 6.264 | 6.248 | 666.2194 | 1.627 | 1.637 | 1.632 | 177.7142 | 0.26675
2020-02-11 | 13:01 | 6.224 | 6.229 | 6.2265 | 663.9269 | 1.624 | 1.625 | 1.6245 | 176.8975 | 0.26644
2020-02-11 | 13:01 | 6.23 | 6.247 | 6.2385 | 665.2064 | 1.625 | 1.633 | 1.629 | 177.3875 | 0.26667
2020-02-11 | 13:01 | 6.253 | 6.27 | 6.2615 | 667.6589 | 1.634 | 1.639 | 1.6365 | 178.2042 | 0.26691
2020-02-11 | 13:01 | 6.273 | 6.282 | 6.2775 | 669.3649 | 1.639 | 1.644 | 1.6415 | 178.7487 | 0.26704
2020-02-11 | 13:01 | 6.279 | 6.283 | 6.281 | 669.7381 | 1.642 | 1.644 | 1.643 | 178.912 | 0.26714
2020-02-11 | 13:01 | 6.281 | 6.289 | 6.285 | 670.1647 | 1.644 | 1.646 | 1.645 | 179.1298 | 0.26729
2020-02-11 | 13:01 | 6.29 | 6.295 | 6.2925 | 670.9644 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.26722
2020-02-11 | 13:01 | 6.29 | 6.299 | 6.2945 | 671.1776 | 1.645 | 1.648 | 1.6465 | 179.2932 | 0.26713
2020-02-11 | 13:01 | 6.292 | 6.299 | 6.2955 | 671.2843 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26733
2020-02-11 | 13:01 | 6.285 | 6.29 | 6.2875 | 670.4312 | 1.645 | 1.647 | 1.646 | 179.2387 | 0.26735
2020-02-11 | 13:01 | 6.286 | 6.289 | 6.2875 | 670.4312 | 1.645 | 1.646 | 1.6455 | 179.1843 | 0.26727
2020-02-11 | 13:02 | 6.283 | 6.287 | 6.285 | 670.1647 | 1.644 | 1.645 | 1.6445 | 179.0754 | 0.26721
2020-02-11 | 13:02 | 6.283 | 6.291 | 6.287 | 670.3779 | 1.644 | 1.647 | 1.6455 | 179.1843 | 0.26729
2020-02-11 | 13:02 | 6.291 | 6.295 | 6.293 | 671.0177 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26744
2020-02-11 | 13:02 | 6.29 | 6.294 | 6.292 | 670.9111 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.26724
2020-02-11 | 13:02 | 6.293 | 6.302 | 6.2975 | 671.4975 | 1.646 | 1.649 | 1.6475 | 179.402 | 0.26717
2020-02-11 | 13:02 | 6.301 | 6.305 | 6.303 | 672.084 | 1.648 | 1.65 | 1.649 | 179.5654 | 0.26718
2020-02-11 | 13:02 | 6.298 | 6.301 | 6.2995 | 671.7108 | 1.648 | 1.649 | 1.6485 | 179.5109 | 0.26724
2020-02-11 | 13:02 | 6.297 | 6.299 | 6.298 | 671.5508 | 1.649 | 1.65 | 1.6495 | 179.6198 | 0.26747
2020-02-11 | 13:02 | 6.288 | 6.295 | 6.2915 | 670.8578 | 1.643 | 1.649 | 1.646 | 179.2387 | 0.26718
2020-02-11 | 13:02 | 6.282 | 6.288 | 6.285 | 670.1647 | 1.643 | 1.646 | 1.6445 | 179.0754 | 0.26721
2020-02-11 | 13:02 | 6.285 | 6.29 | 6.2875 | 670.4312 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26751
2020-02-11 | 13:02 | 6.292 | 6.293 | 6.2925 | 670.9644 | 1.647 | 1.648 | 1.6475 | 179.402 | 0.26738
2020-02-11 | 13:02 | 6.291 | 6.294 | 6.2925 | 670.9644 | 1.643 | 1.646 | 1.6445 | 179.0754 | 0.26689
2020-02-11 | 13:03 | 6.295 | 6.3 | 6.2975 | 671.4975 | 1.646 | 1.651 | 1.6485 | 179.5109 | 0.26733
2020-02-11 | 13:03 | 6.3 | 6.304 | 6.302 | 671.9774 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26754
2020-02-11 | 13:03 | 6.296 | 6.299 | 6.2975 | 671.4975 | 1.648 | 1.651 | 1.6495 | 179.6198 | 0.26749
2020-02-11 | 13:03 | 6.296 | 6.302 | 6.299 | 671.6575 | 1.65 | 1.651 | 1.6505 | 179.7287 | 0.26759
2020-02-11 | 13:03 | 6.303 | 6.309 | 6.306 | 672.4039 | 1.651 | 1.655 | 1.653 | 180.001 | 0.2677
2020-02-11 | 13:03 | 6.311 | 6.316 | 6.3135 | 673.2036 | 1.654 | 1.655 | 1.6545 | 180.1643 | 0.26762
2020-02-11 | 13:03 | 6.313 | 6.314 | 6.3135 | 673.2036 | 1.654 | 1.656 | 1.655 | 180.2187 | 0.2677
2020-02-11 | 13:03 | 6.312 | 6.314 | 6.313 | 673.1503 | 1.653 | 1.655 | 1.654 | 180.1099 | 0.26756
2020-02-11 | 13:03 | 6.309 | 6.313 | 6.311 | 672.937 | 1.652 | 1.655 | 1.6535 | 180.0554 | 0.26757
2020-02-11 | 13:03 | 6.307 | 6.31 | 6.3085 | 672.6705 | 1.653 | 1.655 | 1.654 | 180.1099 | 0.26775
2020-02-11 | 13:03 | 6.307 | 6.308 | 6.3075 | 672.5638 | 1.654 | 1.655 | 1.6545 | 180.1643 | 0.26788
2020-02-11 | 13:04 | 6.305 | 6.308 | 6.3065 | 672.4572 | 1.653 | 1.653 | 1.653 | 180.001 | 0.26768
2020-02-11 | 13:04 | 6.308 | 6.318 | 6.313 | 673.1503 | 1.654 | 1.657 | 1.6555 | 180.2732 | 0.26781
2020-02-11 | 13:04 | 6.32 | 6.323 | 6.3215 | 674.0566 | 1.656 | 1.658 | 1.657 | 180.4365 | 0.26769
2020-02-11 | 13:04 | 6.32 | 6.324 | 6.322 | 674.1099 | 1.657 | 1.658 | 1.6575 | 180.491 | 0.26775
2020-02-11 | 13:04 | 6.317 | 6.322 | 6.3195 | 673.8434 | 1.657 | 1.658 | 1.6575 | 180.491 | 0.26785
2020-02-11 | 13:04 | 6.319 | 6.325 | 6.322 | 674.1099 | 1.657 | 1.659 | 1.658 | 180.5454 | 0.26783
2020-02-11 | 13:04 | 6.317 | 6.325 | 6.321 | 674.0033 | 1.658 | 1.659 | 1.6585 | 180.5999 | 0.26795
2020-02-11 | 13:04 | 6.306 | 6.315 | 6.3105 | 672.8837 | 1.655 | 1.658 | 1.6565 | 180.3821 | 0.26807
2020-02-11 | 13:04 | 6.299 | 6.304 | 6.3015 | 671.924 | 1.652 | 1.654 | 1.653 | 180.001 | 0.26789
2020-02-11 | 13:04 | 6.301 | 6.309 | 6.305 | 672.2973 | 1.653 | 1.657 | 1.655 | 180.2187 | 0.26806
2020-02-11 | 13:04 | 6.31 | 6.311 | 6.3105 | 672.8837 | 1.655 | 1.657 | 1.656 | 180.3276 | 0.26799
2020-02-11 | 13:04 | 6.3 | 6.309 | 6.3045 | 672.2439 | 1.654 | 1.656 | 1.655 | 180.2187 | 0.26809
2020-02-11 | 13:04 | 6.3 | 6.303 | 6.3015 | 671.924 | 1.655 | 1.656 | 1.6555 | 180.2732 | 0.26829
2020-02-11 | 13:05 | 6.301 | 6.302 | 6.3015 | 671.924 | 1.654 | 1.654 | 1.654 | 180.1099 | 0.26805
2020-02-11 | 13:05 | 6.3 | 6.302 | 6.301 | 671.8707 | 1.653 | 1.655 | 1.654 | 180.1099 | 0.26807
2020-02-11 | 13:05 | 6.3 | 6.302 | 6.301 | 671.8707 | 1.654 | 1.656 | 1.655 | 180.2187 | 0.26823
2020-02-11 | 13:05 | 6.301 | 6.302 | 6.3015 | 671.924 | 1.655 | 1.656 | 1.6555 | 180.2732 | 0.26829
2020-02-11 | 13:05 | 6.304 | 6.307 | 6.3055 | 672.3506 | 1.656 | 1.657 | 1.6565 | 180.3821 | 0.26829
2020-02-11 | 13:05 | 6.307 | 6.312 | 6.3095 | 672.7771 | 1.656 | 1.658 | 1.657 | 180.4365 | 0.2682
2020-02-11 | 13:05 | 6.313 | 6.315 | 6.314 | 673.2569 | 1.657 | 1.659 | 1.658 | 180.5454 | 0.26817
2020-02-11 | 13:05 | 6.313 | 6.316 | 6.3145 | 673.3102 | 1.655 | 1.658 | 1.6565 | 180.3821 | 0.2679
2020-02-11 | 13:05 | 6.31 | 6.312 | 6.311 | 672.937 | 1.656 | 1.658 | 1.657 | 180.4365 | 0.26813
2020-02-11 | 13:05 | 6.314 | 6.316 | 6.315 | 673.3635 | 1.658 | 1.659 | 1.6585 | 180.5999 | 0.26821
2020-02-11 | 13:05 | 6.313 | 6.316 | 6.3145 | 673.3102 | 1.657 | 1.66 | 1.6585 | 180.5999 | 0.26823
2020-02-11 | 13:06 | 6.319 | 6.325 | 6.322 | 674.1099 | 1.66 | 1.662 | 1.661 | 180.8721 | 0.26831
2020-02-11 | 13:06 | 6.316 | 6.32 | 6.318 | 673.6834 | 1.66 | 1.661 | 1.6605 | 180.8177 | 0.2684
2020-02-11 | 13:06 | 6.319 | 6.324 | 6.3215 | 674.0566 | 1.661 | 1.662 | 1.6615 | 180.9266 | 0.26841
2020-02-11 | 13:06 | 6.323 | 6.325 | 6.324 | 674.3232 | 1.662 | 1.664 | 1.663 | 181.0899 | 0.26855
2020-02-11 | 13:06 | 6.317 | 6.323 | 6.32 | 673.8967 | 1.66 | 1.663 | 1.6615 | 180.9266 | 0.26848
2020-02-11 | 13:06 | 6.315 | 6.316 | 6.3155 | 673.4169 | 1.66 | 1.661 | 1.6605 | 180.8177 | 0.26851
2020-02-11 | 13:06 | 6.313 | 6.316 | 6.3145 | 673.3102 | 1.659 | 1.661 | 1.66 | 180.7632 | 0.26847
2020-02-11 | 13:06 | 6.309 | 6.313 | 6.311 | 672.937 | 1.658 | 1.661 | 1.6595 | 180.7088 | 0.26854
2020-02-11 | 13:06 | 6.305 | 6.309 | 6.307 | 672.5105 | 1.658 | 1.659 | 1.6585 | 180.5999 | 0.26855
2020-02-11 | 13:06 | 6.305 | 6.307 | 6.306 | 672.4039 | 1.658 | 1.66 | 1.659 | 180.6543 | 0.26867
2020-02-11 | 13:06 | 6.3 | 6.305 | 6.3025 | 672.0307 | 1.658 | 1.658 | 1.658 | 180.5454 | 0.26866
2020-02-11 | 13:06 | 6.289 | 6.299 | 6.294 | 671.1243 | 1.653 | 1.656 | 1.6545 | 180.1643 | 0.26845
2020-02-11 | 13:06 | 6.28 | 6.285 | 6.2825 | 669.8981 | 1.651 | 1.652 | 1.6515 | 179.8376 | 0.26846
2020-02-11 | 13:07 | 6.277 | 6.28 | 6.2785 | 669.4716 | 1.649 | 1.65 | 1.6495 | 179.6198 | 0.2683
2020-02-11 | 13:07 | 6.275 | 6.277 | 6.276 | 669.205 | 1.648 | 1.65 | 1.649 | 179.5654 | 0.26833
2020-02-11 | 13:07 | 6.271 | 6.275 | 6.273 | 668.8851 | 1.647 | 1.65 | 1.6485 | 179.5109 | 0.26837
2020-02-11 | 13:07 | 6.267 | 6.273 | 6.27 | 668.5652 | 1.648 | 1.65 | 1.649 | 179.5654 | 0.26858
2020-02-11 | 13:07 | 6.263 | 6.268 | 6.2655 | 668.0854 | 1.647 | 1.648 | 1.6475 | 179.402 | 0.26853
2020-02-11 | 13:07 | 6.26 | 6.266 | 6.263 | 667.8188 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26856
2020-02-11 | 13:07 | 6.269 | 6.273 | 6.271 | 668.6719 | 1.649 | 1.651 | 1.65 | 179.6743 | 0.2687
2020-02-11 | 13:07 | 6.268 | 6.273 | 6.2705 | 668.6185 | 1.648 | 1.651 | 1.6495 | 179.6198 | 0.26864
2020-02-11 | 13:07 | 6.273 | 6.285 | 6.279 | 669.5249 | 1.65 | 1.653 | 1.6515 | 179.8376 | 0.2686
2020-02-11 | 13:07 | 6.277 | 6.286 | 6.2815 | 669.7915 | 1.652 | 1.653 | 1.6525 | 179.9465 | 0.26866
2020-02-11 | 13:07 | 6.267 | 6.273 | 6.27 | 668.5652 | 1.649 | 1.65 | 1.6495 | 179.6198 | 0.26866
2020-02-11 | 13:08 | 6.271 | 6.275 | 6.273 | 668.8851 | 1.649 | 1.652 | 1.6505 | 179.7287 | 0.2687
2020-02-11 | 13:08 | 6.271 | 6.276 | 6.2735 | 668.9384 | 1.651 | 1.654 | 1.6525 | 179.9465 | 0.269
2020-02-11 | 13:08 | 6.265 | 6.271 | 6.268 | 668.352 | 1.65 | 1.652 | 1.651 | 179.7832 | 0.26899
2020-02-11 | 13:08 | 6.257 | 6.263 | 6.26 | 667.4989 | 1.647 | 1.651 | 1.649 | 179.5654 | 0.26901
2020-02-11 | 13:08 | 6.252 | 6.254 | 6.253 | 666.7525 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26899
2020-02-11 | 13:08 | 6.247 | 6.251 | 6.249 | 666.326 | 1.644 | 1.646 | 1.645 | 179.1298 | 0.26883
2020-02-11 | 13:08 | 6.242 | 6.247 | 6.2445 | 665.8462 | 1.642 | 1.645 | 1.6435 | 178.9665 | 0.26878
2020-02-11 | 13:08 | 6.234 | 6.241 | 6.2375 | 665.0998 | 1.64 | 1.643 | 1.6415 | 178.7487 | 0.26875
2020-02-11 | 13:08 | 6.226 | 6.231 | 6.2285 | 664.1401 | 1.638 | 1.639 | 1.6385 | 178.422 | 0.26865
2020-02-11 | 13:08 | 6.217 | 6.221 | 6.219 | 663.1271 | 1.635 | 1.636 | 1.6355 | 178.0953 | 0.26857
2020-02-11 | 13:08 | 6.216 | 6.218 | 6.217 | 662.9139 | 1.634 | 1.635 | 1.6345 | 177.9864 | 0.26849
2020-02-11 | 13:08 | 6.212 | 6.215 | 6.2135 | 662.5407 | 1.633 | 1.634 | 1.6335 | 177.8775 | 0.26848
2020-02-11 | 13:08 | 6.213 | 6.218 | 6.2155 | 662.7539 | 1.634 | 1.636 | 1.635 | 178.0409 | 0.26864
2020-02-11 | 13:09 | 6.22 | 6.222 | 6.221 | 663.3404 | 1.636 | 1.637 | 1.6365 | 178.2042 | 0.26865
2020-02-11 | 13:09 | 6.213 | 6.218 | 6.2155 | 662.7539 | 1.634 | 1.635 | 1.6345 | 177.9864 | 0.26856
2020-02-11 | 13:09 | 6.204 | 6.212 | 6.208 | 661.9542 | 1.633 | 1.634 | 1.6335 | 177.8775 | 0.26872
2020-02-11 | 13:09 | 6.202 | 6.209 | 6.2055 | 661.6876 | 1.632 | 1.635 | 1.6335 | 177.8775 | 0.26882
2020-02-11 | 13:09 | 6.21 | 6.217 | 6.2135 | 662.5407 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.26897
2020-02-11 | 13:09 | 6.219 | 6.232 | 6.2255 | 663.8202 | 1.638 | 1.641 | 1.6395 | 178.5309 | 0.26894
2020-02-11 | 13:09 | 6.233 | 6.24 | 6.2365 | 664.9931 | 1.642 | 1.645 | 1.6435 | 178.9665 | 0.26913
2020-02-11 | 13:09 | 6.241 | 6.242 | 6.2415 | 665.5263 | 1.644 | 1.646 | 1.645 | 179.1298 | 0.26916
2020-02-11 | 13:09 | 6.244 | 6.25 | 6.247 | 666.1128 | 1.645 | 1.647 | 1.646 | 179.2387 | 0.26908
2020-02-11 | 13:09 | 6.251 | 6.254 | 6.2525 | 666.6992 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26917
2020-02-11 | 13:09 | 6.254 | 6.256 | 6.255 | 666.9658 | 1.648 | 1.649 | 1.6485 | 179.5109 | 0.26915
2020-02-11 | 13:10 | 6.253 | 6.255 | 6.254 | 666.8592 | 1.649 | 1.651 | 1.65 | 179.6743 | 0.26943
2020-02-11 | 13:10 | 6.246 | 6.254 | 6.25 | 666.4326 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.26928
2020-02-11 | 13:10 | 6.243 | 6.246 | 6.2445 | 665.8462 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26935
2020-02-11 | 13:10 | 6.241 | 6.245 | 6.243 | 665.6862 | 1.646 | 1.649 | 1.6475 | 179.402 | 0.2695
2020-02-11 | 13:10 | 6.247 | 6.249 | 6.248 | 666.2194 | 1.648 | 1.65 | 1.649 | 179.5654 | 0.26953
2020-02-11 | 13:10 | 6.248 | 6.251 | 6.2495 | 666.3793 | 1.648 | 1.649 | 1.6485 | 179.5109 | 0.26938
2020-02-11 | 13:10 | 6.248 | 6.251 | 6.2495 | 666.3793 | 1.649 | 1.65 | 1.6495 | 179.6198 | 0.26955
2020-02-11 | 13:10 | 6.243 | 6.248 | 6.2455 | 665.9528 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.26931
2020-02-11 | 13:10 | 6.239 | 6.243 | 6.241 | 665.473 | 1.646 | 1.648 | 1.647 | 179.3476 | 0.2695
2020-02-11 | 13:10 | 6.236 | 6.239 | 6.2375 | 665.0998 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.26957
2020-02-11 | 13:10 | 6.24 | 6.244 | 6.242 | 665.5796 | 1.646 | 1.649 | 1.6475 | 179.402 | 0.26954
2020-02-11 | 13:10 | 6.233 | 6.242 | 6.2375 | 665.0998 | 1.644 | 1.646 | 1.645 | 179.1298 | 0.26933
2020-02-11 | 13:10 | 6.223 | 6.229 | 6.226 | 663.8735 | 1.641 | 1.644 | 1.6425 | 178.8576 | 0.26942
2020-02-11 | 13:11 | 6.219 | 6.222 | 6.2205 | 663.2871 | 1.64 | 1.642 | 1.641 | 178.6942 | 0.26941
2020-02-11 | 13:11 | 6.219 | 6.222 | 6.2205 | 663.2871 | 1.64 | 1.641 | 1.6405 | 178.6398 | 0.26932
2020-02-11 | 13:11 | 6.223 | 6.228 | 6.2255 | 663.8202 | 1.641 | 1.644 | 1.6425 | 178.8576 | 0.26944
2020-02-11 | 13:11 | 6.223 | 6.227 | 6.225 | 663.7669 | 1.64 | 1.642 | 1.641 | 178.6942 | 0.26921
2020-02-11 | 13:11 | 6.215 | 6.224 | 6.2195 | 663.1805 | 1.638 | 1.64 | 1.639 | 178.4764 | 0.26912
2020-02-11 | 13:11 | 6.204 | 6.212 | 6.208 | 661.9542 | 1.634 | 1.637 | 1.6355 | 178.0953 | 0.26904
2020-02-11 | 13:11 | 6.202 | 6.204 | 6.203 | 661.4211 | 1.634 | 1.635 | 1.6345 | 177.9864 | 0.2691
2020-02-11 | 13:11 | 6.192 | 6.199 | 6.1955 | 660.6214 | 1.632 | 1.633 | 1.6325 | 177.7686 | 0.26909
2020-02-11 | 13:11 | 6.186 | 6.189 | 6.1875 | 659.7683 | 1.629 | 1.631 | 1.63 | 177.4964 | 0.26903
2020-02-11 | 13:11 | 6.187 | 6.19 | 6.1885 | 659.8749 | 1.629 | 1.632 | 1.6305 | 177.5509 | 0.26907
2020-02-11 | 13:11 | 6.188 | 6.191 | 6.1895 | 659.9816 | 1.629 | 1.63 | 1.6295 | 177.442 | 0.26886
2020-02-11 | 13:12 | 6.177 | 6.186 | 6.1815 | 659.1285 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26896
2020-02-11 | 13:12 | 6.177 | 6.177 | 6.177 | 658.6487 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.26907
2020-02-11 | 13:12 | 6.176 | 6.179 | 6.1775 | 658.702 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26889
2020-02-11 | 13:12 | 6.175 | 6.178 | 6.1765 | 658.5954 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26885
2020-02-11 | 13:12 | 6.173 | 6.178 | 6.1755 | 658.4888 | 1.624 | 1.627 | 1.6255 | 177.0064 | 0.26881
2020-02-11 | 13:12 | 6.18 | 6.185 | 6.1825 | 659.2352 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26867
2020-02-11 | 13:12 | 6.185 | 6.189 | 6.187 | 659.715 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.26864
2020-02-11 | 13:12 | 6.19 | 6.195 | 6.1925 | 660.3015 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.2684
2020-02-11 | 13:12 | 6.195 | 6.203 | 6.199 | 660.9946 | 1.628 | 1.63 | 1.629 | 177.3875 | 0.26836
2020-02-11 | 13:12 | 6.202 | 6.203 | 6.2025 | 661.3678 | 1.63 | 1.632 | 1.631 | 177.6053 | 0.26854
2020-02-11 | 13:12 | 6.204 | 6.22 | 6.212 | 662.3807 | 1.632 | 1.636 | 1.634 | 177.932 | 0.26862
2020-02-11 | 13:12 | 6.223 | 6.232 | 6.2275 | 664.0335 | 1.636 | 1.638 | 1.637 | 178.2587 | 0.26845
2020-02-11 | 13:12 | 6.224 | 6.231 | 6.2275 | 664.0335 | 1.635 | 1.638 | 1.6365 | 178.2042 | 0.26837
2020-02-11 | 13:13 | 6.217 | 6.222 | 6.2195 | 663.1805 | 1.635 | 1.636 | 1.6355 | 178.0953 | 0.26855
2020-02-11 | 13:13 | 6.216 | 6.22 | 6.218 | 663.0205 | 1.635 | 1.636 | 1.6355 | 178.0953 | 0.26861
2020-02-11 | 13:13 | 6.216 | 6.218 | 6.217 | 662.9139 | 1.635 | 1.636 | 1.6355 | 178.0953 | 0.26866
2020-02-11 | 13:13 | 6.209 | 6.215 | 6.212 | 662.3807 | 1.634 | 1.637 | 1.6355 | 178.0953 | 0.26887
2020-02-11 | 13:13 | 6.205 | 6.209 | 6.207 | 661.8476 | 1.631 | 1.633 | 1.632 | 177.7142 | 0.26851
2020-02-11 | 13:13 | 6.199 | 6.206 | 6.2025 | 661.3678 | 1.63 | 1.633 | 1.6315 | 177.6597 | 0.26862
2020-02-11 | 13:13 | 6.19 | 6.198 | 6.194 | 660.4614 | 1.628 | 1.631 | 1.6295 | 177.442 | 0.26866
2020-02-11 | 13:13 | 6.187 | 6.189 | 6.188 | 659.8216 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26868
2020-02-11 | 13:13 | 6.19 | 6.193 | 6.1915 | 660.1948 | 1.628 | 1.63 | 1.629 | 177.3875 | 0.26869
2020-02-11 | 13:13 | 6.192 | 6.194 | 6.193 | 660.3548 | 1.629 | 1.63 | 1.6295 | 177.442 | 0.26871
2020-02-11 | 13:13 | 6.181 | 6.19 | 6.1855 | 659.5551 | 1.625 | 1.629 | 1.627 | 177.1697 | 0.26862
2020-02-11 | 13:14 | 6.171 | 6.177 | 6.174 | 658.3288 | 1.624 | 1.624 | 1.624 | 176.843 | 0.26862
2020-02-11 | 13:14 | 6.17 | 6.172 | 6.171 | 658.0089 | 1.622 | 1.624 | 1.623 | 176.7342 | 0.26859
2020-02-11 | 13:14 | 6.169 | 6.174 | 6.1715 | 658.0622 | 1.622 | 1.624 | 1.623 | 176.7342 | 0.26857
2020-02-11 | 13:14 | 6.161 | 6.167 | 6.164 | 657.2625 | 1.62 | 1.621 | 1.6205 | 176.4619 | 0.26848
2020-02-11 | 13:14 | 6.163 | 6.168 | 6.1655 | 657.4225 | 1.621 | 1.622 | 1.6215 | 176.5708 | 0.26858
2020-02-11 | 13:14 | 6.163 | 6.165 | 6.164 | 657.2625 | 1.621 | 1.622 | 1.6215 | 176.5708 | 0.26865
2020-02-11 | 13:14 | 6.168 | 6.174 | 6.171 | 658.0089 | 1.622 | 1.623 | 1.6225 | 176.6797 | 0.26851
2020-02-11 | 13:14 | 6.176 | 6.183 | 6.1795 | 658.9153 | 1.624 | 1.628 | 1.626 | 177.0608 | 0.26872
2020-02-11 | 13:14 | 6.182 | 6.184 | 6.183 | 659.2885 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26856
2020-02-11 | 13:14 | 6.183 | 6.19 | 6.1865 | 659.6617 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26849
2020-02-11 | 13:14 | 6.19 | 6.192 | 6.191 | 660.1415 | 1.628 | 1.629 | 1.6285 | 177.3331 | 0.26863
2020-02-11 | 13:14 | 6.192 | 6.194 | 6.193 | 660.3548 | 1.629 | 1.629 | 1.629 | 177.3875 | 0.26862
2020-02-11 | 13:14 | 6.188 | 6.191 | 6.1895 | 659.9816 | 1.628 | 1.629 | 1.6285 | 177.3331 | 0.26869
2020-02-11 | 13:15 | 6.189 | 6.2 | 6.1945 | 660.5147 | 1.628 | 1.631 | 1.6295 | 177.442 | 0.26864
2020-02-11 | 13:15 | 6.192 | 6.2 | 6.196 | 660.6747 | 1.629 | 1.63 | 1.6295 | 177.442 | 0.26858
2020-02-11 | 13:15 | 6.187 | 6.19 | 6.1885 | 659.8749 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.26857
2020-02-11 | 13:15 | 6.187 | 6.19 | 6.1885 | 659.8749 | 1.627 | 1.63 | 1.6285 | 177.3331 | 0.26874
2020-02-11 | 13:15 | 6.191 | 6.198 | 6.1945 | 660.5147 | 1.629 | 1.632 | 1.6305 | 177.5509 | 0.26881
2020-02-11 | 13:15 | 6.195 | 6.199 | 6.197 | 660.7813 | 1.628 | 1.63 | 1.629 | 177.3875 | 0.26845
2020-02-11 | 13:15 | 6.19 | 6.194 | 6.192 | 660.2481 | 1.628 | 1.629 | 1.6285 | 177.3331 | 0.26859
2020-02-11 | 13:15 | 6.185 | 6.191 | 6.188 | 659.8216 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26868
2020-02-11 | 13:15 | 6.184 | 6.189 | 6.1865 | 659.6617 | 1.626 | 1.628 | 1.627 | 177.1697 | 0.26858
2020-02-11 | 13:15 | 6.188 | 6.191 | 6.1895 | 659.9816 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.26853
2020-02-11 | 13:15 | 6.189 | 6.19 | 6.1895 | 659.9816 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26861
2020-02-11 | 13:16 | 6.191 | 6.193 | 6.192 | 660.2481 | 1.628 | 1.629 | 1.6285 | 177.3331 | 0.26859
2020-02-11 | 13:16 | 6.181 | 6.191 | 6.186 | 659.6084 | 1.625 | 1.628 | 1.6265 | 177.1153 | 0.26852
2020-02-11 | 13:16 | 6.179 | 6.186 | 6.1825 | 659.2352 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26859
2020-02-11 | 13:16 | 6.183 | 6.186 | 6.1845 | 659.4484 | 1.625 | 1.628 | 1.6265 | 177.1153 | 0.26858
2020-02-11 | 13:16 | 6.18 | 6.183 | 6.1815 | 659.1285 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26863
2020-02-11 | 13:16 | 6.182 | 6.196 | 6.189 | 659.9283 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26863
2020-02-11 | 13:16 | 6.191 | 6.198 | 6.1945 | 660.5147 | 1.628 | 1.63 | 1.629 | 177.3875 | 0.26856
2020-02-11 | 13:16 | 6.186 | 6.192 | 6.189 | 659.9283 | 1.626 | 1.628 | 1.627 | 177.1697 | 0.26847
2020-02-11 | 13:16 | 6.192 | 6.192 | 6.192 | 660.2481 | 1.628 | 1.629 | 1.6285 | 177.3331 | 0.26859
2020-02-11 | 13:16 | 6.192 | 6.198 | 6.195 | 660.568 | 1.627 | 1.629 | 1.628 | 177.2786 | 0.26837
2020-02-11 | 13:16 | 6.196 | 6.197 | 6.1965 | 660.728 | 1.628 | 1.63 | 1.629 | 177.3875 | 0.26847
2020-02-11 | 13:16 | 6.189 | 6.196 | 6.1925 | 660.3015 | 1.628 | 1.629 | 1.6285 | 177.3331 | 0.26856
2020-02-11 | 13:16 | 6.183 | 6.188 | 6.1855 | 659.5551 | 1.624 | 1.627 | 1.6255 | 177.0064 | 0.26837
2020-02-11 | 13:17 | 6.183 | 6.184 | 6.1835 | 659.3418 | 1.623 | 1.625 | 1.624 | 176.843 | 0.26821
2020-02-11 | 13:17 | 6.175 | 6.183 | 6.179 | 658.862 | 1.62 | 1.624 | 1.622 | 176.6253 | 0.26808
2020-02-11 | 13:17 | 6.163 | 6.172 | 6.1675 | 657.6357 | 1.617 | 1.62 | 1.6185 | 176.2441 | 0.268
2020-02-11 | 13:17 | 6.154 | 6.161 | 6.1575 | 656.5694 | 1.615 | 1.618 | 1.6165 | 176.0263 | 0.2681
2020-02-11 | 13:17 | 6.15 | 6.153 | 6.1515 | 655.9297 | 1.61 | 1.613 | 1.6115 | 175.4819 | 0.26753
2020-02-11 | 13:17 | 6.151 | 6.153 | 6.152 | 655.983 | 1.607 | 1.611 | 1.609 | 175.2096 | 0.26709
2020-02-11 | 13:17 | 6.155 | 6.166 | 6.1605 | 656.8893 | 1.607 | 1.608 | 1.6075 | 175.0463 | 0.26648
2020-02-11 | 13:17 | 6.167 | 6.173 | 6.17 | 657.9023 | 1.61 | 1.613 | 1.6115 | 175.4819 | 0.26673
2020-02-11 | 13:17 | 6.174 | 6.178 | 6.176 | 658.5421 | 1.614 | 1.615 | 1.6145 | 175.8086 | 0.26697
2020-02-11 | 13:17 | 6.179 | 6.182 | 6.1805 | 659.0219 | 1.614 | 1.615 | 1.6145 | 175.8086 | 0.26677
2020-02-11 | 13:17 | 6.184 | 6.192 | 6.188 | 659.8216 | 1.615 | 1.62 | 1.6175 | 176.1352 | 0.26694
2020-02-11 | 13:18 | 6.194 | 6.204 | 6.199 | 660.9946 | 1.623 | 1.625 | 1.624 | 176.843 | 0.26754
2020-02-11 | 13:18 | 6.206 | 6.208 | 6.207 | 661.8476 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26753
2020-02-11 | 13:18 | 6.207 | 6.21 | 6.2085 | 662.0075 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26754
2020-02-11 | 13:18 | 6.209 | 6.211 | 6.21 | 662.1675 | 1.627 | 1.627 | 1.627 | 177.1697 | 0.26756
2020-02-11 | 13:18 | 6.213 | 6.218 | 6.2155 | 662.7539 | 1.627 | 1.63 | 1.6285 | 177.3331 | 0.26757
2020-02-11 | 13:18 | 6.213 | 6.219 | 6.216 | 662.8072 | 1.628 | 1.631 | 1.6295 | 177.442 | 0.26771
2020-02-11 | 13:18 | 6.207 | 6.211 | 6.209 | 662.0608 | 1.627 | 1.628 | 1.6275 | 177.2242 | 0.26769
2020-02-11 | 13:18 | 6.205 | 6.206 | 6.2055 | 661.6876 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26767
2020-02-11 | 13:18 | 6.2 | 6.204 | 6.202 | 661.3144 | 1.626 | 1.627 | 1.6265 | 177.1153 | 0.26782
2020-02-11 | 13:18 | 6.199 | 6.202 | 6.2005 | 661.1545 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26781
2020-02-11 | 13:18 | 6.194 | 6.198 | 6.196 | 660.6747 | 1.622 | 1.626 | 1.624 | 176.843 | 0.26767
2020-02-11 | 13:18 | 6.193 | 6.195 | 6.194 | 660.4614 | 1.623 | 1.624 | 1.6235 | 176.7886 | 0.26767
2020-02-11 | 13:18 | 6.195 | 6.2 | 6.1975 | 660.8346 | 1.624 | 1.626 | 1.625 | 176.9519 | 0.26777
2020-02-11 | 13:19 | 6.201 | 6.206 | 6.2035 | 661.4744 | 1.624 | 1.628 | 1.626 | 177.0608 | 0.26768
2020-02-11 | 13:19 | 6.2 | 6.206 | 6.203 | 661.4211 | 1.626 | 1.628 | 1.627 | 177.1697 | 0.26786
2020-02-11 | 13:19 | 6.195 | 6.201 | 6.198 | 660.8879 | 1.625 | 1.627 | 1.626 | 177.0608 | 0.26791
2020-02-11 | 13:19 | 6.187 | 6.195 | 6.191 | 660.1415 | 1.623 | 1.626 | 1.6245 | 176.8975 | 0.26797
2020-02-11 | 13:19 | 6.18 | 6.184 | 6.182 | 659.1819 | 1.621 | 1.624 | 1.6225 | 176.6797 | 0.26803
2020-02-11 | 13:19 | 6.174 | 6.18 | 6.177 | 658.6487 | 1.622 | 1.623 | 1.6225 | 176.6797 | 0.26825
2020-02-11 | 13:19 | 6.169 | 6.172 | 6.1705 | 657.9556 | 1.62 | 1.622 | 1.621 | 176.5164 | 0.26828
2020-02-11 | 13:19 | 6.169 | 6.175 | 6.172 | 658.1156 | 1.618 | 1.621 | 1.6195 | 176.353 | 0.26797
2020-02-11 | 13:19 | 6.161 | 6.166 | 6.1635 | 657.2092 | 1.615 | 1.618 | 1.6165 | 176.0263 | 0.26784
2020-02-11 | 13:19 | 6.154 | 6.158 | 6.156 | 656.4095 | 1.613 | 1.615 | 1.614 | 175.7541 | 0.26775
2020-02-11 | 13:19 | 6.149 | 6.154 | 6.1515 | 655.9297 | 1.613 | 1.614 | 1.6135 | 175.6997 | 0.26786
2020-02-11 | 13:20 | 6.139 | 6.145 | 6.142 | 654.9167 | 1.61 | 1.613 | 1.6115 | 175.4819 | 0.26795
2020-02-11 | 13:20 | 6.133 | 6.137 | 6.135 | 654.1703 | 1.607 | 1.611 | 1.609 | 175.2096 | 0.26783
2020-02-11 | 13:20 | 6.127 | 6.13 | 6.1285 | 653.4772 | 1.607 | 1.609 | 1.608 | 175.1008 | 0.26795
2020-02-11 | 13:20 | 6.126 | 6.127 | 6.1265 | 653.2639 | 1.607 | 1.609 | 1.608 | 175.1008 | 0.26804
2020-02-11 | 13:20 | 6.127 | 6.13 | 6.1285 | 653.4772 | 1.607 | 1.609 | 1.608 | 175.1008 | 0.26795
2020-02-11 | 13:20 | 6.127 | 6.129 | 6.128 | 653.4239 | 1.606 | 1.608 | 1.607 | 174.9919 | 0.26781
2020-02-11 | 13:20 | 6.126 | 6.128 | 6.127 | 653.3172 | 1.607 | 1.609 | 1.608 | 175.1008 | 0.26802
2020-02-11 | 13:20 | 6.117 | 6.126 | 6.1215 | 652.7308 | 1.604 | 1.607 | 1.6055 | 174.8285 | 0.26784
2020-02-11 | 13:20 | 6.116 | 6.123 | 6.1195 | 652.5175 | 1.604 | 1.607 | 1.6055 | 174.8285 | 0.26793
2020-02-11 | 13:20 | 6.127 | 6.13 | 6.1285 | 653.4772 | 1.608 | 1.609 | 1.6085 | 175.1552 | 0.26804
2020-02-11 | 13:20 | 6.129 | 6.133 | 6.131 | 653.7438 | 1.609 | 1.61 | 1.6095 | 175.2641 | 0.26809
2020-02-11 | 13:20 | 6.132 | 6.135 | 6.1335 | 654.0103 | 1.609 | 1.61 | 1.6095 | 175.2641 | 0.26798
2020-02-11 | 13:20 | 6.136 | 6.138 | 6.137 | 654.3835 | 1.61 | 1.611 | 1.6105 | 175.373 | 0.268
2020-02-11 | 13:21 | 6.133 | 6.135 | 6.134 | 654.0637 | 1.609 | 1.611 | 1.61 | 175.3185 | 0.26805
2020-02-11 | 13:21 | 6.133 | 6.136 | 6.1345 | 654.117 | 1.609 | 1.61 | 1.6095 | 175.2641 | 0.26794
2020-02-11 | 13:21 | 6.127 | 6.133 | 6.13 | 653.6371 | 1.607 | 1.61 | 1.6085 | 175.1552 | 0.26797
2020-02-11 | 13:21 | 6.124 | 6.127 | 6.1255 | 653.1573 | 1.607 | 1.609 | 1.608 | 175.1008 | 0.26808
2020-02-11 | 13:21 | 6.123 | 6.125 | 6.124 | 652.9974 | 1.606 | 1.607 | 1.6065 | 174.9374 | 0.2679
2020-02-11 | 13:21 | 6.12 | 6.124 | 6.122 | 652.7841 | 1.605 | 1.607 | 1.606 | 174.883 | 0.2679
2020-02-11 | 13:21 | 6.109 | 6.116 | 6.1125 | 651.7711 | 1.602 | 1.605 | 1.6035 | 174.6107 | 0.2679
2020-02-11 | 13:21 | 6.106 | 6.108 | 6.107 | 651.1847 | 1.603 | 1.604 | 1.6035 | 174.6107 | 0.26814
2020-02-11 | 13:21 | 6.106 | 6.108 | 6.107 | 651.1847 | 1.602 | 1.604 | 1.603 | 174.5563 | 0.26806
2020-02-11 | 13:21 | 6.109 | 6.112 | 6.1105 | 651.5579 | 1.604 | 1.604 | 1.604 | 174.6652 | 0.26807
2020-02-11 | 13:21 | 6.102 | 6.108 | 6.105 | 650.9714 | 1.602 | 1.603 | 1.6025 | 174.5018 | 0.26806
2020-02-11 | 13:22 | 6.103 | 6.105 | 6.104 | 650.8648 | 1.601 | 1.603 | 1.602 | 174.4474 | 0.26802
2020-02-11 | 13:22 | 6.1 | 6.104 | 6.102 | 650.6515 | 1.601 | 1.603 | 1.602 | 174.4474 | 0.26811
2020-02-11 | 13:22 | 6.106 | 6.109 | 6.1075 | 651.238 | 1.603 | 1.606 | 1.6045 | 174.7196 | 0.26829
2020-02-11 | 13:22 | 6.099 | 6.11 | 6.1045 | 650.9181 | 1.603 | 1.605 | 1.604 | 174.6652 | 0.26834
2020-02-11 | 13:22 | 6.089 | 6.098 | 6.0935 | 649.7452 | 1.601 | 1.602 | 1.6015 | 174.3929 | 0.2684
2020-02-11 | 13:22 | 6.082 | 6.088 | 6.085 | 648.8388 | 1.598 | 1.6 | 1.599 | 174.1207 | 0.26836
2020-02-11 | 13:22 | 6.079 | 6.083 | 6.081 | 648.4123 | 1.597 | 1.598 | 1.5975 | 173.9574 | 0.26828
2020-02-11 | 13:22 | 6.07 | 6.076 | 6.073 | 647.5593 | 1.596 | 1.598 | 1.597 | 173.9029 | 0.26855
2020-02-11 | 13:22 | 6.065 | 6.069 | 6.067 | 646.9195 | 1.594 | 1.596 | 1.595 | 173.6851 | 0.26848
2020-02-11 | 13:22 | 6.061 | 6.065 | 6.063 | 646.493 | 1.593 | 1.596 | 1.5945 | 173.6307 | 0.26857
2020-02-11 | 13:22 | 6.056 | 6.059 | 6.0575 | 645.9065 | 1.593 | 1.593 | 1.593 | 173.4673 | 0.26856
2020-02-11 | 13:22 | 6.046 | 6.057 | 6.0515 | 645.2667 | 1.59 | 1.593 | 1.5915 | 173.304 | 0.26858
2020-02-11 | 13:22 | 6.04 | 6.044 | 6.042 | 644.2538 | 1.589 | 1.59 | 1.5895 | 173.0862 | 0.26866
2020-02-11 | 13:23 | 6.037 | 6.041 | 6.039 | 643.9339 | 1.588 | 1.589 | 1.5885 | 172.9773 | 0.26863
2020-02-11 | 13:23 | 6.039 | 6.04 | 6.0395 | 643.9872 | 1.587 | 1.589 | 1.588 | 172.9229 | 0.26852
2020-02-11 | 13:23 | 6.04 | 6.043 | 6.0415 | 644.2005 | 1.588 | 1.59 | 1.589 | 173.0318 | 0.2686
2020-02-11 | 13:23 | 6.031 | 6.038 | 6.0345 | 643.454 | 1.586 | 1.588 | 1.587 | 172.814 | 0.26857
2020-02-11 | 13:23 | 6.031 | 6.033 | 6.032 | 643.1875 | 1.585 | 1.586 | 1.5855 | 172.6506 | 0.26843
2020-02-11 | 13:23 | 6.024 | 6.03 | 6.027 | 642.6543 | 1.585 | 1.586 | 1.5855 | 172.6506 | 0.26865
2020-02-11 | 13:23 | 6.025 | 6.028 | 6.0265 | 642.601 | 1.584 | 1.586 | 1.585 | 172.5962 | 0.26859
2020-02-11 | 13:23 | 6.021 | 6.023 | 6.022 | 642.1212 | 1.583 | 1.584 | 1.5835 | 172.4329 | 0.26854
2020-02-11 | 13:23 | 6.016 | 6.02 | 6.018 | 641.6947 | 1.582 | 1.585 | 1.5835 | 172.4329 | 0.26871
2020-02-11 | 13:23 | 6.009 | 6.013 | 6.011 | 640.9483 | 1.579 | 1.581 | 1.58 | 172.0517 | 0.26843
2020-02-11 | 13:23 | 6.011 | 6.019 | 6.015 | 641.3748 | 1.581 | 1.582 | 1.5815 | 172.2151 | 0.26851
2020-02-11 | 13:24 | 6.023 | 6.032 | 6.0275 | 642.7076 | 1.583 | 1.586 | 1.5845 | 172.5418 | 0.26846
2020-02-11 | 13:24 | 6.034 | 6.039 | 6.0365 | 643.6673 | 1.587 | 1.588 | 1.5875 | 172.8684 | 0.26857
2020-02-11 | 13:24 | 6.04 | 6.042 | 6.041 | 644.1471 | 1.588 | 1.591 | 1.5895 | 173.0862 | 0.26871
2020-02-11 | 13:24 | 6.039 | 6.041 | 6.04 | 644.0405 | 1.588 | 1.59 | 1.589 | 173.0318 | 0.26867
2020-02-11 | 13:24 | 6.042 | 6.047 | 6.0445 | 644.5203 | 1.59 | 1.592 | 1.591 | 173.2496 | 0.2688
2020-02-11 | 13:24 | 6.046 | 6.047 | 6.0465 | 644.7336 | 1.59 | 1.592 | 1.591 | 173.2496 | 0.26871
2020-02-11 | 13:24 | 6.047 | 6.05 | 6.0485 | 644.9469 | 1.592 | 1.593 | 1.5925 | 173.4129 | 0.26888
2020-02-11 | 13:24 | 6.051 | 6.058 | 6.0545 | 645.5866 | 1.593 | 1.596 | 1.5945 | 173.6307 | 0.26895
2020-02-11 | 13:24 | 6.062 | 6.066 | 6.064 | 646.5996 | 1.595 | 1.597 | 1.596 | 173.794 | 0.26878
2020-02-11 | 13:24 | 6.067 | 6.071 | 6.069 | 647.1328 | 1.597 | 1.599 | 1.598 | 174.0118 | 0.2689
2020-02-11 | 13:24 | 6.07 | 6.074 | 6.072 | 647.4526 | 1.599 | 1.6 | 1.5995 | 174.1752 | 0.26902
2020-02-11 | 13:24 | 6.068 | 6.074 | 6.071 | 647.346 | 1.598 | 1.6 | 1.599 | 174.1207 | 0.26898
2020-02-11 | 13:24 | 6.065 | 6.066 | 6.0655 | 646.7596 | 1.598 | 1.599 | 1.5985 | 174.0663 | 0.26914
2020-02-11 | 13:25 | 6.066 | 6.07 | 6.068 | 647.0261 | 1.598 | 1.6 | 1.599 | 174.1207 | 0.26911
2020-02-11 | 13:25 | 6.067 | 6.071 | 6.069 | 647.1328 | 1.598 | 1.601 | 1.5995 | 174.1752 | 0.26915
2020-02-11 | 13:25 | 6.062 | 6.067 | 6.0645 | 646.6529 | 1.596 | 1.6 | 1.598 | 174.0118 | 0.2691
2020-02-11 | 13:25 | 6.053 | 6.063 | 6.058 | 645.9598 | 1.595 | 1.597 | 1.596 | 173.794 | 0.26905
2020-02-11 | 13:25 | 6.046 | 6.048 | 6.047 | 644.7869 | 1.593 | 1.594 | 1.5935 | 173.5218 | 0.26911
2020-02-11 | 13:25 | 6.043 | 6.046 | 6.0445 | 644.5203 | 1.593 | 1.594 | 1.5935 | 173.5218 | 0.26923
2020-02-11 | 13:25 | 6.035 | 6.04 | 6.0375 | 643.7739 | 1.592 | 1.593 | 1.5925 | 173.4129 | 0.26937
2020-02-11 | 13:25 | 6.033 | 6.035 | 6.034 | 643.4007 | 1.59 | 1.593 | 1.5915 | 173.304 | 0.26936
2020-02-11 | 13:25 | 6.028 | 6.034 | 6.031 | 643.0808 | 1.59 | 1.592 | 1.591 | 173.2496 | 0.26941
2020-02-11 | 13:25 | 6.028 | 6.031 | 6.0295 | 642.9209 | 1.59 | 1.591 | 1.5905 | 173.1951 | 0.26939
2020-02-11 | 13:25 | 6.028 | 6.03 | 6.029 | 642.8676 | 1.59 | 1.591 | 1.5905 | 173.1951 | 0.26941
2020-02-11 | 13:26 | 6.028 | 6.031 | 6.0295 | 642.9209 | 1.59 | 1.593 | 1.5915 | 173.304 | 0.26956
2020-02-11 | 13:26 | 6.033 | 6.037 | 6.035 | 643.5074 | 1.592 | 1.594 | 1.593 | 173.4673 | 0.26957
2020-02-11 | 13:26 | 6.04 | 6.046 | 6.043 | 644.3604 | 1.594 | 1.596 | 1.595 | 173.6851 | 0.26955
2020-02-11 | 13:26 | 6.048 | 6.055 | 6.0515 | 645.2667 | 1.595 | 1.598 | 1.5965 | 173.8485 | 0.26942
2020-02-11 | 13:26 | 6.052 | 6.055 | 6.0535 | 645.48 | 1.597 | 1.599 | 1.598 | 174.0118 | 0.26959
2020-02-11 | 13:26 | 6.051 | 6.053 | 6.052 | 645.3201 | 1.596 | 1.598 | 1.597 | 173.9029 | 0.26948
2020-02-11 | 13:26 | 6.045 | 6.049 | 6.047 | 644.7869 | 1.594 | 1.599 | 1.5965 | 173.8485 | 0.26962
2020-02-11 | 13:26 | 6.046 | 6.051 | 6.0485 | 644.9469 | 1.597 | 1.598 | 1.5975 | 173.9574 | 0.26972
2020-02-11 | 13:26 | 6.033 | 6.045 | 6.039 | 643.9339 | 1.593 | 1.596 | 1.5945 | 173.6307 | 0.26964
2020-02-11 | 13:26 | 6.029 | 6.032 | 6.0305 | 643.0275 | 1.592 | 1.593 | 1.5925 | 173.4129 | 0.26968
2020-02-11 | 13:26 | 6.028 | 6.031 | 6.0295 | 642.9209 | 1.592 | 1.593 | 1.5925 | 173.4129 | 0.26973
2020-02-11 | 13:26 | 6.017 | 6.026 | 6.0215 | 642.0679 | 1.589 | 1.591 | 1.59 | 173.1407 | 0.26966
2020-02-11 | 13:26 | 6.018 | 6.025 | 6.0215 | 642.0679 | 1.588 | 1.591 | 1.5895 | 173.0862 | 0.26958
2020-02-11 | 13:27 | 6.027 | 6.028 | 6.0275 | 642.7076 | 1.591 | 1.593 | 1.592 | 173.3585 | 0.26973
2020-02-11 | 13:27 | 6.022 | 6.029 | 6.0255 | 642.4944 | 1.589 | 1.592 | 1.5905 | 173.1951 | 0.26957
2020-02-11 | 13:27 | 6.016 | 6.02 | 6.018 | 641.6947 | 1.584 | 1.588 | 1.586 | 172.7051 | 0.26914
2020-02-11 | 13:27 | 6.011 | 6.017 | 6.014 | 641.2681 | 1.512 | 1.577 | 1.5445 | 168.186 | 0.26227
2020-02-11 | 13:27 | 6.011 | 6.014 | 6.0125 | 641.1082 | 1.583 | 1.587 | 1.585 | 172.5962 | 0.26922
2020-02-11 | 13:27 | 5.998 | 6.009 | 6.0035 | 640.1485 | 1.583 | 1.586 | 1.5845 | 172.5418 | 0.26953
2020-02-11 | 13:27 | 5.996 | 5.997 | 5.9965 | 639.4021 | 1.582 | 1.584 | 1.583 | 172.3784 | 0.26959
2020-02-11 | 13:27 | 5.997 | 6 | 5.9985 | 639.6154 | 1.583 | 1.585 | 1.584 | 172.4873 | 0.26967
2020-02-11 | 13:27 | 6.003 | 6.007 | 6.005 | 640.3085 | 1.586 | 1.587 | 1.5865 | 172.7595 | 0.26981
2020-02-11 | 13:27 | 6 | 6.006 | 6.003 | 640.0952 | 1.586 | 1.587 | 1.5865 | 172.7595 | 0.2699
2020-02-11 | 13:27 | 6.001 | 6.006 | 6.0035 | 640.1485 | 1.585 | 1.587 | 1.586 | 172.7051 | 0.26979
2020-02-11 | 13:28 | 6.007 | 6.015 | 6.011 | 640.9483 | 1.588 | 1.59 | 1.589 | 173.0318 | 0.26996
2020-02-11 | 13:28 | 6.017 | 6.019 | 6.018 | 641.6947 | 1.589 | 1.591 | 1.59 | 173.1407 | 0.26982
2020-02-11 | 13:28 | 6.018 | 6.021 | 6.0195 | 641.8546 | 1.59 | 1.592 | 1.591 | 173.2496 | 0.26992
2020-02-11 | 13:28 | 6.021 | 6.028 | 6.0245 | 642.3878 | 1.59 | 1.592 | 1.591 | 173.2496 | 0.2697
2020-02-11 | 13:28 | 6.028 | 6.04 | 6.034 | 643.4007 | 1.593 | 1.596 | 1.5945 | 173.6307 | 0.26986
2020-02-11 | 13:28 | 6.042 | 6.045 | 6.0435 | 644.4137 | 1.596 | 1.598 | 1.597 | 173.9029 | 0.26986
2020-02-11 | 13:28 | 6.034 | 6.042 | 6.038 | 643.8272 | 1.593 | 1.596 | 1.5945 | 173.6307 | 0.26969
2020-02-11 | 13:28 | 6.033 | 6.034 | 6.0335 | 643.3474 | 1.594 | 1.595 | 1.5945 | 173.6307 | 0.26989
2020-02-11 | 13:28 | 6.034 | 6.046 | 6.04 | 644.0405 | 1.593 | 1.596 | 1.5945 | 173.6307 | 0.2696
2020-02-11 | 13:28 | 6.048 | 6.05 | 6.049 | 645.0002 | 1.596 | 1.598 | 1.597 | 173.9029 | 0.26962
2020-02-11 | 13:28 | 6.042 | 6.048 | 6.045 | 644.5737 | 1.596 | 1.597 | 1.5965 | 173.8485 | 0.26971
2020-02-11 | 13:28 | 6.032 | 6.038 | 6.035 | 643.5074 | 1.593 | 1.594 | 1.5935 | 173.5218 | 0.26965
2020-02-11 | 13:28 | 6.032 | 6.034 | 6.033 | 643.2941 | 1.592 | 1.595 | 1.5935 | 173.5218 | 0.26974
2020-02-11 | 13:29 | 6.037 | 6.04 | 6.0385 | 643.8806 | 1.594 | 1.596 | 1.595 | 173.6851 | 0.26975
2020-02-11 | 13:29 | 6.041 | 6.043 | 6.042 | 644.2538 | 1.595 | 1.597 | 1.596 | 173.794 | 0.26976
2020-02-11 | 13:29 | 6.036 | 6.04 | 6.038 | 643.8272 | 1.594 | 1.595 | 1.5945 | 173.6307 | 0.26969
2020-02-11 | 13:29 | 6.034 | 6.038 | 6.036 | 643.614 | 1.594 | 1.594 | 1.594 | 173.5762 | 0.26969
2020-02-11 | 13:29 | 6.038 | 6.041 | 6.0395 | 643.9872 | 1.595 | 1.596 | 1.5955 | 173.7396 | 0.26979
2020-02-11 | 13:29 | 6.043 | 6.046 | 6.0445 | 644.5203 | 1.594 | 1.597 | 1.5955 | 173.7396 | 0.26956
2020-02-11 | 13:29 | 6.037 | 6.045 | 6.041 | 644.1471 | 1.592 | 1.594 | 1.593 | 173.4673 | 0.2693
2020-02-11 | 13:29 | 6.027 | 6.034 | 6.0305 | 643.0275 | 1.591 | 1.592 | 1.5915 | 173.304 | 0.26951
2020-02-11 | 13:29 | 6.025 | 6.027 | 6.026 | 642.5477 | 1.59 | 1.592 | 1.591 | 173.2496 | 0.26963
2020-02-11 | 13:29 | 6.009 | 6.021 | 6.015 | 641.3748 | 1.586 | 1.59 | 1.588 | 172.9229 | 0.26961
2020-02-11 | 13:29 | 6.006 | 6.009 | 6.0075 | 640.5751 | 1.582 | 1.584 | 1.583 | 172.3784 | 0.2691
2020-02-11 | 13:30 | 6.006 | 6.01 | 6.008 | 640.6284 | 1.584 | 1.586 | 1.585 | 172.5962 | 0.26942
2020-02-11 | 13:30 | 6.011 | 6.014 | 6.0125 | 641.1082 | 1.587 | 1.587 | 1.587 | 172.814 | 0.26956
2020-02-11 | 13:30 | 6.012 | 6.017 | 6.0145 | 641.3215 | 1.587 | 1.588 | 1.5875 | 172.8684 | 0.26955
2020-02-11 | 13:30 | 6.018 | 6.02 | 6.019 | 641.8013 | 1.588 | 1.59 | 1.589 | 173.0318 | 0.2696
2020-02-11 | 13:30 | 6.019 | 6.022 | 6.0205 | 641.9612 | 1.588 | 1.59 | 1.589 | 173.0318 | 0.26954
2020-02-11 | 13:30 | 6.018 | 6.02 | 6.019 | 641.8013 | 1.587 | 1.589 | 1.588 | 172.9229 | 0.26943
2020-02-11 | 13:30 | 6.011 | 6.015 | 6.013 | 641.1615 | 1.586 | 1.589 | 1.5875 | 172.8684 | 0.26962
2020-02-11 | 13:30 | 6.009 | 6.013 | 6.011 | 640.9483 | 1.586 | 1.589 | 1.5875 | 172.8684 | 0.26971
2020-02-11 | 13:30 | 6.007 | 6.014 | 6.0105 | 640.8949 | 1.587 | 1.588 | 1.5875 | 172.8684 | 0.26973
2020-02-11 | 13:30 | 6.005 | 6.011 | 6.008 | 640.6284 | 1.586 | 1.589 | 1.5875 | 172.8684 | 0.26984
2020-02-11 | 13:30 | 6.002 | 6.004 | 6.003 | 640.0952 | 1.585 | 1.587 | 1.586 | 172.7051 | 0.26981
2020-02-11 | 13:30 | 6.003 | 6.005 | 6.004 | 640.2019 | 1.584 | 1.586 | 1.585 | 172.5962 | 0.2696
2020-02-11 | 13:30 | 5.998 | 6.003 | 6.0005 | 639.8287 | 1.584 | 1.586 | 1.585 | 172.5962 | 0.26975
2020-02-11 | 13:31 | 5.993 | 5.997 | 5.995 | 639.2422 | 1.582 | 1.583 | 1.5825 | 172.324 | 0.26958
2020-02-11 | 13:31 | 5.979 | 5.991 | 5.985 | 638.1759 | 1.579 | 1.582 | 1.5805 | 172.1062 | 0.26968
2020-02-11 | 13:31 | 5.972 | 5.978 | 5.975 | 637.1096 | 1.578 | 1.579 | 1.5785 | 171.8884 | 0.26979
2020-02-11 | 13:31 | 5.966 | 5.969 | 5.9675 | 636.3099 | 1.576 | 1.578 | 1.577 | 171.7251 | 0.26988
2020-02-11 | 13:31 | 5.959 | 5.964 | 5.9615 | 635.6701 | 1.576 | 1.578 | 1.577 | 171.7251 | 0.27015
2020-02-11 | 13:31 | 5.954 | 5.958 | 5.956 | 635.0837 | 1.575 | 1.576 | 1.5755 | 171.5617 | 0.27014
2020-02-11 | 13:31 | 5.945 | 5.952 | 5.9485 | 634.2839 | 1.573 | 1.575 | 1.574 | 171.3984 | 0.27022
2020-02-11 | 13:31 | 5.943 | 5.945 | 5.944 | 633.8041 | 1.571 | 1.573 | 1.572 | 171.1806 | 0.27008
2020-02-11 | 13:31 | 5.941 | 5.945 | 5.943 | 633.6975 | 1.57 | 1.572 | 1.571 | 171.0717 | 0.26996
2020-02-11 | 13:31 | 5.931 | 5.938 | 5.9345 | 632.7911 | 1.568 | 1.571 | 1.5695 | 170.9084 | 0.27009
2020-02-11 | 13:31 | 5.924 | 5.929 | 5.9265 | 631.9381 | 1.566 | 1.568 | 1.567 | 170.6361 | 0.27002
2020-02-11 | 13:32 | 5.92 | 5.922 | 5.921 | 631.3516 | 1.565 | 1.567 | 1.566 | 170.5272 | 0.2701
2020-02-11 | 13:32 | 5.921 | 5.923 | 5.922 | 631.4583 | 1.566 | 1.567 | 1.5665 | 170.5817 | 0.27014
2020-02-11 | 13:32 | 5.92 | 5.927 | 5.9235 | 631.6182 | 1.565 | 1.568 | 1.5665 | 170.5817 | 0.27007
2020-02-11 | 13:32 | 5.929 | 5.934 | 5.9315 | 632.4712 | 1.568 | 1.57 | 1.569 | 170.8539 | 0.27014
2020-02-11 | 13:32 | 5.934 | 5.937 | 5.9355 | 632.8978 | 1.569 | 1.57 | 1.5695 | 170.9084 | 0.27004
2020-02-11 | 13:32 | 5.935 | 5.939 | 5.937 | 633.0577 | 1.569 | 1.571 | 1.57 | 170.9628 | 0.27006
2020-02-11 | 13:32 | 5.938 | 5.947 | 5.9425 | 633.6442 | 1.57 | 1.573 | 1.5715 | 171.1261 | 0.27007
2020-02-11 | 13:32 | 5.946 | 5.955 | 5.9505 | 634.4972 | 1.572 | 1.575 | 1.5735 | 171.3439 | 0.27005
2020-02-11 | 13:32 | 5.956 | 5.959 | 5.9575 | 635.2436 | 1.575 | 1.576 | 1.5755 | 171.5617 | 0.27007
2020-02-11 | 13:32 | 5.959 | 5.963 | 5.961 | 635.6168 | 1.575 | 1.577 | 1.576 | 171.6162 | 0.27
2020-02-11 | 13:32 | 5.963 | 5.966 | 5.9645 | 635.99 | 1.576 | 1.578 | 1.577 | 171.7251 | 0.27001
2020-02-11 | 13:32 | 5.955 | 5.961 | 5.958 | 635.2969 | 1.575 | 1.577 | 1.576 | 171.6162 | 0.27014
2020-02-11 | 13:32 | 5.954 | 5.967 | 5.9605 | 635.5635 | 1.575 | 1.578 | 1.5765 | 171.6706 | 0.27011
2020-02-11 | 13:33 | 5.971 | 5.975 | 5.973 | 636.8963 | 1.579 | 1.58 | 1.5795 | 171.9973 | 0.27006
2020-02-11 | 13:33 | 5.975 | 5.98 | 5.9775 | 637.3762 | 1.579 | 1.582 | 1.5805 | 172.1062 | 0.27002
2020-02-11 | 13:33 | 5.983 | 5.988 | 5.9855 | 638.2292 | 1.582 | 1.584 | 1.583 | 172.3784 | 0.27009
2020-02-11 | 13:33 | 5.98 | 5.987 | 5.9835 | 638.016 | 1.582 | 1.584 | 1.583 | 172.3784 | 0.27018
2020-02-11 | 13:33 | 5.981 | 5.985 | 5.983 | 637.9626 | 1.581 | 1.582 | 1.5815 | 172.2151 | 0.26995
2020-02-11 | 13:33 | 5.974 | 5.979 | 5.9765 | 637.2696 | 1.579 | 1.581 | 1.58 | 172.0517 | 0.26998
2020-02-11 | 13:33 | 5.966 | 5.971 | 5.9685 | 636.4165 | 1.577 | 1.579 | 1.578 | 171.8339 | 0.27
2020-02-11 | 13:33 | 5.963 | 5.965 | 5.964 | 635.9367 | 1.576 | 1.578 | 1.577 | 171.7251 | 0.27003
2020-02-11 | 13:33 | 5.96 | 5.964 | 5.962 | 635.7234 | 1.576 | 1.577 | 1.5765 | 171.6706 | 0.27004
2020-02-11 | 13:33 | 5.966 | 5.969 | 5.9675 | 636.3099 | 1.578 | 1.579 | 1.5785 | 171.8884 | 0.27013
2020-02-11 | 13:33 | 5.971 | 5.977 | 5.974 | 637.003 | 1.579 | 1.581 | 1.58 | 172.0517 | 0.2701
2020-02-11 | 13:34 | 5.978 | 5.987 | 5.9825 | 637.9093 | 1.579 | 1.582 | 1.5805 | 172.1062 | 0.2698
2020-02-11 | 13:34 | 5.987 | 5.992 | 5.9895 | 638.6557 | 1.582 | 1.583 | 1.5825 | 172.324 | 0.26982
2020-02-11 | 13:34 | 5.987 | 5.994 | 5.9905 | 638.7624 | 1.583 | 1.584 | 1.5835 | 172.4329 | 0.26995
2020-02-11 | 13:34 | 5.986 | 5.994 | 5.99 | 638.709 | 1.583 | 1.585 | 1.584 | 172.4873 | 0.27006
2020-02-11 | 13:34 | 5.986 | 5.994 | 5.99 | 638.709 | 1.582 | 1.584 | 1.583 | 172.3784 | 0.26989
2020-02-11 | 13:34 | 5.961 | 5.978 | 5.9695 | 636.5231 | 1.577 | 1.58 | 1.5785 | 171.8884 | 0.27004
2020-02-11 | 13:34 | 5.955 | 5.958 | 5.9565 | 635.137 | 1.575 | 1.576 | 1.5755 | 171.5617 | 0.27012
2020-02-11 | 13:34 | 5.949 | 5.959 | 5.954 | 634.8704 | 1.574 | 1.577 | 1.5755 | 171.5617 | 0.27023
2020-02-11 | 13:34 | 5.939 | 5.944 | 5.9415 | 633.5375 | 1.572 | 1.573 | 1.5725 | 171.235 | 0.27028
2020-02-11 | 13:34 | 5.942 | 5.944 | 5.943 | 633.6975 | 1.573 | 1.575 | 1.574 | 171.3984 | 0.27047
2020-02-11 | 13:34 | 5.943 | 5.945 | 5.944 | 633.8041 | 1.573 | 1.575 | 1.574 | 171.3984 | 0.27043
2020-02-11 | 13:34 | 5.948 | 5.951 | 5.9495 | 634.3906 | 1.574 | 1.576 | 1.575 | 171.5073 | 0.27035
2020-02-11 | 13:34 | 5.951 | 5.953 | 5.952 | 634.6571 | 1.575 | 1.576 | 1.5755 | 171.5617 | 0.27032
2020-02-11 | 13:35 | 5.952 | 5.958 | 5.955 | 634.977 | 1.576 | 1.577 | 1.5765 | 171.6706 | 0.27036
2020-02-11 | 13:35 | 5.959 | 5.961 | 5.96 | 635.5102 | 1.577 | 1.579 | 1.578 | 171.8339 | 0.27039
2020-02-11 | 13:35 | 5.961 | 5.967 | 5.964 | 635.9367 | 1.578 | 1.58 | 1.579 | 171.9428 | 0.27038
2020-02-11 | 13:35 | 5.962 | 5.967 | 5.9645 | 635.99 | 1.579 | 1.58 | 1.5795 | 171.9973 | 0.27044
2020-02-11 | 13:35 | 5.959 | 5.962 | 5.9605 | 635.5635 | 1.577 | 1.579 | 1.578 | 171.8339 | 0.27036
2020-02-11 | 13:35 | 5.959 | 5.964 | 5.9615 | 635.6701 | 1.575 | 1.577 | 1.576 | 171.6162 | 0.26998
2020-02-11 | 13:35 | 5.961 | 5.964 | 5.9625 | 635.7767 | 1.575 | 1.577 | 1.576 | 171.6162 | 0.26993
2020-02-11 | 13:35 | 5.954 | 5.959 | 5.9565 | 635.137 | 1.573 | 1.574 | 1.5735 | 171.3439 | 0.26977
2020-02-11 | 13:35 | 5.951 | 5.955 | 5.953 | 634.7638 | 1.573 | 1.575 | 1.574 | 171.3984 | 0.27002
2020-02-11 | 13:35 | 5.948 | 5.952 | 5.95 | 634.4439 | 1.573 | 1.574 | 1.5735 | 171.3439 | 0.27007
2020-02-11 | 13:35 | 5.947 | 5.952 | 5.9495 | 634.3906 | 1.572 | 1.573 | 1.5725 | 171.235 | 0.26992
2020-02-11 | 13:36 | 5.941 | 5.946 | 5.9435 | 633.7508 | 1.57 | 1.573 | 1.5715 | 171.1261 | 0.27002
2020-02-11 | 13:36 | 5.936 | 5.94 | 5.938 | 633.1643 | 1.57 | 1.571 | 1.5705 | 171.0172 | 0.2701
2020-02-11 | 13:36 | 5.932 | 5.934 | 5.933 | 632.6312 | 1.568 | 1.57 | 1.569 | 170.8539 | 0.27007
2020-02-11 | 13:36 | 5.93 | 5.932 | 5.931 | 632.4179 | 1.567 | 1.569 | 1.568 | 170.745 | 0.26999
2020-02-11 | 13:36 | 5.935 | 5.937 | 5.936 | 632.9511 | 1.569 | 1.57 | 1.5695 | 170.9084 | 0.27002
2020-02-11 | 13:36 | 5.929 | 5.935 | 5.932 | 632.5246 | 1.569 | 1.57 | 1.5695 | 170.9084 | 0.2702
2020-02-11 | 13:36 | 5.932 | 5.936 | 5.934 | 632.7378 | 1.569 | 1.572 | 1.5705 | 171.0172 | 0.27028
2020-02-11 | 13:36 | 5.935 | 5.937 | 5.936 | 632.9511 | 1.57 | 1.572 | 1.571 | 171.0717 | 0.27028
2020-02-11 | 13:36 | 5.934 | 5.938 | 5.936 | 632.9511 | 1.569 | 1.57 | 1.5695 | 170.9084 | 0.27002
2020-02-11 | 13:36 | 5.931 | 5.938 | 5.9345 | 632.7911 | 1.568 | 1.57 | 1.569 | 170.8539 | 0.27
2020-02-11 | 13:36 | 5.928 | 5.935 | 5.9315 | 632.4712 | 1.567 | 1.568 | 1.5675 | 170.6906 | 0.26988
2020-02-11 | 13:36 | 5.934 | 5.936 | 5.935 | 632.8444 | 1.567 | 1.57 | 1.5685 | 170.7995 | 0.26989
2020-02-11 | 13:36 | 5.931 | 5.938 | 5.9345 | 632.7911 | 1.567 | 1.57 | 1.5685 | 170.7995 | 0.26991
2020-02-11 | 13:37 | 5.941 | 5.942 | 5.9415 | 633.5375 | 1.568 | 1.57 | 1.569 | 170.8539 | 0.26968
2020-02-11 | 13:37 | 5.932 | 5.942 | 5.937 | 633.0577 | 1.568 | 1.57 | 1.569 | 170.8539 | 0.26989
2020-02-11 | 13:37 | 5.926 | 5.932 | 5.929 | 632.2047 | 1.565 | 1.567 | 1.566 | 170.5272 | 0.26973
2020-02-11 | 13:37 | 5.93 | 5.941 | 5.9355 | 632.8978 | 1.567 | 1.569 | 1.568 | 170.745 | 0.26978
2020-02-11 | 13:37 | 5.943 | 5.948 | 5.9455 | 633.964 | 1.57 | 1.571 | 1.5705 | 171.0172 | 0.26976
2020-02-11 | 13:37 | 5.944 | 5.952 | 5.948 | 634.2306 | 1.57 | 1.573 | 1.5715 | 171.1261 | 0.26982
2020-02-11 | 13:37 | 5.928 | 5.949 | 5.9385 | 633.2176 | 1.567 | 1.573 | 1.57 | 170.9628 | 0.26999
2020-02-11 | 13:37 | 5.922 | 5.926 | 5.924 | 631.6715 | 1.564 | 1.566 | 1.565 | 170.4183 | 0.26979
2020-02-11 | 13:37 | 5.917 | 5.922 | 5.9195 | 631.1917 | 1.563 | 1.565 | 1.564 | 170.3094 | 0.26982
2020-02-11 | 13:37 | 5.922 | 5.933 | 5.9275 | 632.0447 | 1.564 | 1.567 | 1.5655 | 170.4728 | 0.26972
2020-02-11 | 13:37 | 5.928 | 5.938 | 5.933 | 632.6312 | 1.566 | 1.568 | 1.567 | 170.6361 | 0.26972
2020-02-11 | 13:38 | 5.935 | 5.938 | 5.9365 | 633.0044 | 1.567 | 1.569 | 1.568 | 170.745 | 0.26974
2020-02-11 | 13:38 | 5.928 | 5.935 | 5.9315 | 632.4712 | 1.564 | 1.568 | 1.566 | 170.5272 | 0.26962
2020-02-11 | 13:38 | 5.916 | 5.925 | 5.9205 | 631.2983 | 1.563 | 1.565 | 1.564 | 170.3094 | 0.26978
2020-02-11 | 13:38 | 5.905 | 5.914 | 5.9095 | 630.1254 | 1.56 | 1.564 | 1.562 | 170.0916 | 0.26993
2020-02-11 | 13:38 | 5.894 | 5.902 | 5.898 | 628.8992 | 1.557 | 1.559 | 1.558 | 169.6561 | 0.26977
2020-02-11 | 13:38 | 5.886 | 5.894 | 5.89 | 628.0461 | 1.554 | 1.558 | 1.556 | 169.4383 | 0.26979
2020-02-11 | 13:38 | 5.873 | 5.882 | 5.8775 | 626.7133 | 1.552 | 1.554 | 1.553 | 169.1116 | 0.26984
2020-02-11 | 13:38 | 5.863 | 5.87 | 5.8665 | 625.5403 | 1.551 | 1.553 | 1.552 | 169.0027 | 0.27017
2020-02-11 | 13:38 | 5.844 | 5.86 | 5.852 | 623.9942 | 1.544 | 1.549 | 1.5465 | 168.4038 | 0.26988
2020-02-11 | 13:38 | 5.839 | 5.843 | 5.841 | 622.8213 | 1.542 | 1.544 | 1.543 | 168.0227 | 0.26978
2020-02-11 | 13:38 | 5.844 | 5.849 | 5.8465 | 623.4078 | 1.544 | 1.545 | 1.5445 | 168.186 | 0.26978
2020-02-11 | 13:38 | 5.845 | 5.85 | 5.8475 | 623.5144 | 1.543 | 1.545 | 1.544 | 168.1316 | 0.26965
2020-02-11 | 13:38 | 5.841 | 5.849 | 5.845 | 623.2478 | 1.543 | 1.546 | 1.5445 | 168.186 | 0.26985
2020-02-11 | 13:39 | 5.848 | 5.85 | 5.849 | 623.6743 | 1.546 | 1.546 | 1.546 | 168.3494 | 0.26993
2020-02-11 | 13:39 | 5.84 | 5.848 | 5.844 | 623.1412 | 1.543 | 1.545 | 1.544 | 168.1316 | 0.26981
2020-02-11 | 13:39 | 5.832 | 5.839 | 5.8355 | 622.2348 | 1.542 | 1.543 | 1.5425 | 167.9682 | 0.26994
2020-02-11 | 13:39 | 5.83 | 5.834 | 5.832 | 621.8616 | 1.543 | 1.545 | 1.544 | 168.1316 | 0.27037
2020-02-11 | 13:39 | 5.817 | 5.828 | 5.8225 | 620.8487 | 1.539 | 1.542 | 1.5405 | 167.7504 | 0.2702
2020-02-11 | 13:39 | 5.808 | 5.815 | 5.8115 | 619.6757 | 1.536 | 1.539 | 1.5375 | 167.4238 | 0.27018
2020-02-11 | 13:39 | 5.795 | 5.807 | 5.801 | 618.5561 | 1.532 | 1.537 | 1.5345 | 167.0971 | 0.27014
2020-02-11 | 13:39 | 5.775 | 5.789 | 5.782 | 616.5302 | 1.527 | 1.531 | 1.529 | 166.4982 | 0.27006
2020-02-11 | 13:39 | 5.772 | 5.773 | 5.7725 | 615.5172 | 1.525 | 1.528 | 1.5265 | 166.2259 | 0.27006
2020-02-11 | 13:39 | 5.76 | 5.771 | 5.7655 | 614.7708 | 1.522 | 1.526 | 1.524 | 165.9537 | 0.26994
2020-02-11 | 13:39 | 5.753 | 5.758 | 5.7555 | 613.7045 | 1.521 | 1.524 | 1.5225 | 165.7904 | 0.27015
2020-02-11 | 13:40 | 5.753 | 5.759 | 5.756 | 613.7578 | 1.522 | 1.522 | 1.522 | 165.7359 | 0.27003
2020-02-11 | 13:40 | 5.75 | 5.759 | 5.7545 | 613.5979 | 1.522 | 1.524 | 1.523 | 165.8448 | 0.27028
2020-02-11 | 13:40 | 5.741 | 5.746 | 5.7435 | 612.4249 | 1.518 | 1.52 | 1.519 | 165.4092 | 0.27009
2020-02-11 | 13:40 | 5.745 | 5.76 | 5.7525 | 613.3846 | 1.52 | 1.524 | 1.522 | 165.7359 | 0.2702
2020-02-11 | 13:40 | 5.76 | 5.764 | 5.762 | 614.3976 | 1.525 | 1.526 | 1.5255 | 166.117 | 0.27037
2020-02-11 | 13:40 | 5.754 | 5.759 | 5.7565 | 613.8111 | 1.522 | 1.524 | 1.523 | 165.8448 | 0.27019
2020-02-11 | 13:40 | 5.738 | 5.753 | 5.7455 | 612.6382 | 1.519 | 1.522 | 1.5205 | 165.5726 | 0.27026
2020-02-11 | 13:40 | 5.726 | 5.733 | 5.7295 | 610.9321 | 1.514 | 1.517 | 1.5155 | 165.0281 | 0.27013
2020-02-11 | 13:40 | 5.725 | 5.727 | 5.726 | 610.5589 | 1.514 | 1.514 | 1.514 | 164.8648 | 0.27002
2020-02-11 | 13:40 | 5.716 | 5.723 | 5.7195 | 609.8658 | 1.512 | 1.514 | 1.513 | 164.7559 | 0.27015
2020-02-11 | 13:40 | 5.697 | 5.714 | 5.7055 | 608.373 | 1.506 | 1.511 | 1.5085 | 164.2658 | 0.27001
2020-02-11 | 13:40 | 5.67 | 5.691 | 5.6805 | 605.7073 | 1.497 | 1.505 | 1.501 | 163.4491 | 0.26985
2020-02-11 | 13:40 | 5.663 | 5.667 | 5.665 | 604.0545 | 1.494 | 1.496 | 1.495 | 162.7958 | 0.26951
2020-02-11 | 13:41 | 5.665 | 5.672 | 5.6685 | 604.4278 | 1.494 | 1.498 | 1.496 | 162.9047 | 0.26952
2020-02-11 | 13:41 | 5.676 | 5.688 | 5.682 | 605.8672 | 1.499 | 1.502 | 1.5005 | 163.3947 | 0.26969
2020-02-11 | 13:41 | 5.687 | 5.696 | 5.6915 | 606.8802 | 1.502 | 1.504 | 1.503 | 163.6669 | 0.26969
2020-02-11 | 13:41 | 5.699 | 5.709 | 5.704 | 608.2131 | 1.504 | 1.508 | 1.506 | 163.9936 | 0.26963
2020-02-11 | 13:41 | 5.714 | 5.743 | 5.7285 | 610.8255 | 1.51 | 1.519 | 1.5145 | 164.9192 | 0.26999
2020-02-11 | 13:41 | 5.747 | 5.763 | 5.755 | 613.6512 | 1.519 | 1.523 | 1.521 | 165.627 | 0.2699
2020-02-11 | 13:41 | 5.766 | 5.769 | 5.7675 | 614.984 | 1.524 | 1.527 | 1.5255 | 166.117 | 0.27012
2020-02-11 | 13:41 | 5.77 | 5.773 | 5.7715 | 615.4106 | 1.526 | 1.529 | 1.5275 | 166.3348 | 0.27028
2020-02-11 | 13:41 | 5.776 | 5.782 | 5.779 | 616.2103 | 1.528 | 1.531 | 1.5295 | 166.5526 | 0.27029
2020-02-11 | 13:41 | 5.785 | 5.799 | 5.792 | 617.5965 | 1.531 | 1.536 | 1.5335 | 166.9882 | 0.27038
2020-02-11 | 13:41 | 5.801 | 5.818 | 5.8095 | 619.4625 | 1.536 | 1.542 | 1.539 | 167.5871 | 0.27054
2020-02-11 | 13:42 | 5.821 | 5.824 | 5.8225 | 620.8487 | 1.542 | 1.543 | 1.5425 | 167.9682 | 0.27055
2020-02-11 | 13:42 | 5.826 | 5.834 | 5.83 | 621.6484 | 1.544 | 1.547 | 1.5455 | 168.2949 | 0.27072
2020-02-11 | 13:42 | 5.834 | 5.837 | 5.8355 | 622.2348 | 1.546 | 1.548 | 1.547 | 168.4582 | 0.27073
2020-02-11 | 13:42 | 5.837 | 5.843 | 5.84 | 622.7147 | 1.548 | 1.55 | 1.549 | 168.676 | 0.27087
2020-02-11 | 13:42 | 5.842 | 5.845 | 5.8435 | 623.0879 | 1.548 | 1.55 | 1.549 | 168.676 | 0.27071
2020-02-11 | 13:42 | 5.844 | 5.848 | 5.846 | 623.3544 | 1.549 | 1.55 | 1.5495 | 168.7305 | 0.27068
2020-02-11 | 13:42 | 5.842 | 5.844 | 5.843 | 623.0345 | 1.548 | 1.549 | 1.5485 | 168.6216 | 0.27065
2020-02-11 | 13:42 | 5.84 | 5.843 | 5.8415 | 622.8746 | 1.547 | 1.549 | 1.548 | 168.5671 | 0.27063
2020-02-11 | 13:42 | 5.846 | 5.859 | 5.8525 | 624.0475 | 1.549 | 1.553 | 1.551 | 168.8938 | 0.27064
2020-02-11 | 13:42 | 5.861 | 5.866 | 5.8635 | 625.2204 | 1.553 | 1.556 | 1.5545 | 169.2749 | 0.27074
2020-02-11 | 13:42 | 5.867 | 5.872 | 5.8695 | 625.8602 | 1.555 | 1.558 | 1.5565 | 169.4927 | 0.27082
2020-02-11 | 13:42 | 5.871 | 5.873 | 5.872 | 626.1268 | 1.556 | 1.558 | 1.557 | 169.5472 | 0.27079
2020-02-11 | 13:42 | 5.873 | 5.882 | 5.8775 | 626.7133 | 1.558 | 1.56 | 1.559 | 169.765 | 0.27088
2020-02-11 | 13:43 | 5.881 | 5.884 | 5.8825 | 627.2464 | 1.56 | 1.562 | 1.561 | 169.9828 | 0.271
2020-02-11 | 13:43 | 5.877 | 5.883 | 5.88 | 626.9798 | 1.559 | 1.562 | 1.5605 | 169.9283 | 0.27103
2020-02-11 | 13:43 | 5.831 | 5.868 | 5.8495 | 623.7276 | 1.545 | 1.558 | 1.5515 | 168.9483 | 0.27087
2020-02-11 | 13:43 | 5.81 | 5.823 | 5.8165 | 620.2089 | 1.54 | 1.542 | 1.541 | 167.8049 | 0.27056
2020-02-11 | 13:43 | 5.809 | 5.811 | 5.81 | 619.5158 | 1.539 | 1.541 | 1.54 | 167.696 | 0.27069
2020-02-11 | 13:43 | 5.797 | 5.807 | 5.802 | 618.6628 | 1.536 | 1.538 | 1.537 | 167.3693 | 0.27053
2020-02-11 | 13:43 | 5.797 | 5.801 | 5.799 | 618.3429 | 1.535 | 1.538 | 1.5365 | 167.3149 | 0.27059
2020-02-11 | 13:43 | 5.802 | 5.804 | 5.803 | 618.7694 | 1.537 | 1.538 | 1.5375 | 167.4238 | 0.27058
2020-02-11 | 13:43 | 5.789 | 5.801 | 5.795 | 617.9163 | 1.531 | 1.535 | 1.533 | 166.9337 | 0.27016
2020-02-11 | 13:43 | 5.77 | 5.784 | 5.777 | 615.997 | 1.526 | 1.529 | 1.5275 | 166.3348 | 0.27003
2020-02-11 | 13:43 | 5.741 | 5.766 | 5.7535 | 613.4912 | 1.517 | 1.527 | 1.522 | 165.7359 | 0.27015
2020-02-11 | 13:44 | 5.724 | 5.735 | 5.7295 | 610.9321 | 1.515 | 1.517 | 1.516 | 165.0825 | 0.27021
2020-02-11 | 13:44 | 5.712 | 5.723 | 5.7175 | 609.6526 | 1.513 | 1.516 | 1.5145 | 164.9192 | 0.27051
2020-02-11 | 13:44 | 5.698 | 5.711 | 5.7045 | 608.2664 | 1.51 | 1.513 | 1.5115 | 164.5925 | 0.27059
2020-02-11 | 13:44 | 5.694 | 5.696 | 5.695 | 607.2534 | 1.508 | 1.509 | 1.5085 | 164.2658 | 0.27051
2020-02-11 | 13:44 | 5.696 | 5.7 | 5.698 | 607.5733 | 1.508 | 1.51 | 1.509 | 164.3203 | 0.27045
2020-02-11 | 13:44 | 5.687 | 5.698 | 5.6925 | 606.9869 | 1.506 | 1.51 | 1.508 | 164.2114 | 0.27054
2020-02-11 | 13:44 | 5.681 | 5.687 | 5.684 | 606.0805 | 1.505 | 1.507 | 1.506 | 163.9936 | 0.27058
2020-02-11 | 13:44 | 5.672 | 5.682 | 5.677 | 605.3341 | 1.503 | 1.506 | 1.5045 | 163.8303 | 0.27064
2020-02-11 | 13:44 | 5.648 | 5.668 | 5.658 | 603.3081 | 1.497 | 1.501 | 1.499 | 163.2314 | 0.27056
2020-02-11 | 13:44 | 5.633 | 5.643 | 5.638 | 601.1756 | 1.493 | 1.494 | 1.4935 | 162.6324 | 0.27052
2020-02-11 | 13:44 | 5.631 | 5.637 | 5.634 | 600.749 | 1.492 | 1.493 | 1.4925 | 162.5236 | 0.27053
2020-02-11 | 13:44 | 5.641 | 5.662 | 5.6515 | 602.6151 | 1.496 | 1.502 | 1.499 | 163.2314 | 0.27087
2020-02-11 | 13:44 | 5.666 | 5.672 | 5.669 | 604.4811 | 1.503 | 1.505 | 1.504 | 163.7758 | 0.27094
2020-02-11 | 13:45 | 5.664 | 5.669 | 5.6665 | 604.2145 | 1.501 | 1.503 | 1.502 | 163.558 | 0.2707
2020-02-11 | 13:45 | 5.661 | 5.663 | 5.662 | 603.7347 | 1.5 | 1.502 | 1.501 | 163.4491 | 0.27073
2020-02-11 | 13:45 | 5.654 | 5.661 | 5.6575 | 603.2548 | 1.5 | 1.502 | 1.501 | 163.4491 | 0.27095
2020-02-11 | 13:45 | 5.647 | 5.653 | 5.65 | 602.4551 | 1.498 | 1.499 | 1.4985 | 163.1769 | 0.27085
2020-02-11 | 13:45 | 5.636 | 5.647 | 5.6415 | 601.5488 | 1.494 | 1.496 | 1.495 | 162.7958 | 0.27063
2020-02-11 | 13:45 | 5.639 | 5.642 | 5.6405 | 601.4421 | 1.495 | 1.497 | 1.496 | 162.9047 | 0.27086
2020-02-11 | 13:45 | 5.639 | 5.644 | 5.6415 | 601.5488 | 1.494 | 1.497 | 1.4955 | 162.8502 | 0.27072
2020-02-11 | 13:45 | 5.645 | 5.654 | 5.6495 | 602.4018 | 1.496 | 1.5 | 1.498 | 163.1225 | 0.27079
2020-02-11 | 13:45 | 5.653 | 5.656 | 5.6545 | 602.9349 | 1.499 | 1.501 | 1.5 | 163.3403 | 0.27091
2020-02-11 | 13:45 | 5.652 | 5.662 | 5.657 | 603.2015 | 1.499 | 1.502 | 1.5005 | 163.3947 | 0.27088
2020-02-11 | 13:45 | 5.667 | 5.67 | 5.6685 | 604.4278 | 1.503 | 1.504 | 1.5035 | 163.7214 | 0.27087
2020-02-11 | 13:46 | 5.651 | 5.668 | 5.6595 | 603.4681 | 1.501 | 1.504 | 1.5025 | 163.6125 | 0.27112
2020-02-11 | 13:46 | 5.606 | 5.641 | 5.6235 | 599.6294 | 1.485 | 1.496 | 1.4905 | 162.3058 | 0.27068
2020-02-11 | 13:46 | 5.57 | 5.598 | 5.584 | 595.4176 | 1.475 | 1.484 | 1.4795 | 161.1079 | 0.27058
2020-02-11 | 13:46 | 5.54 | 5.563 | 5.5515 | 591.9521 | 1.469 | 1.474 | 1.4715 | 160.2368 | 0.27069
2020-02-11 | 13:46 | 5.54 | 5.549 | 5.5445 | 591.2057 | 1.469 | 1.473 | 1.471 | 160.1823 | 0.27094
2020-02-11 | 13:46 | 5.549 | 5.555 | 5.552 | 592.0054 | 1.473 | 1.475 | 1.474 | 160.509 | 0.27113
2020-02-11 | 13:46 | 5.54 | 5.554 | 5.547 | 591.4723 | 1.47 | 1.475 | 1.4725 | 160.3457 | 0.2711
2020-02-11 | 13:46 | 5.497 | 5.537 | 5.517 | 588.2734 | 1.458 | 1.468 | 1.463 | 159.3112 | 0.27081
2020-02-11 | 13:46 | 5.419 | 5.48 | 5.4495 | 581.076 | 1.436 | 1.453 | 1.4445 | 157.2967 | 0.2707
2020-02-11 | 13:46 | 5.372 | 5.408 | 5.39 | 574.7315 | 1.423 | 1.432 | 1.4275 | 155.4455 | 0.27047
2020-02-11 | 13:46 | 5.348 | 5.366 | 5.357 | 571.2127 | 1.416 | 1.421 | 1.4185 | 154.4654 | 0.27042
2020-02-11 | 13:46 | 5.326 | 5.345 | 5.3355 | 568.9202 | 1.409 | 1.415 | 1.412 | 153.7576 | 0.27026
2020-02-11 | 13:46 | 5.307 | 5.32 | 5.3135 | 566.5744 | 1.404 | 1.407 | 1.4055 | 153.0498 | 0.27013
2020-02-11 | 13:47 | 5.306 | 5.31 | 5.308 | 565.9879 | 1.405 | 1.407 | 1.406 | 153.1043 | 0.27051
2020-02-11 | 13:47 | 5.312 | 5.314 | 5.313 | 566.5211 | 1.406 | 1.408 | 1.407 | 153.2132 | 0.27045
2020-02-11 | 13:47 | 5.313 | 5.338 | 5.3255 | 567.8539 | 1.409 | 1.416 | 1.4125 | 153.8121 | 0.27087
2020-02-11 | 13:47 | 5.343 | 5.351 | 5.347 | 570.1465 | 1.417 | 1.421 | 1.419 | 154.5199 | 0.27102
2020-02-11 | 13:47 | 5.326 | 5.339 | 5.3325 | 568.6003 | 1.413 | 1.417 | 1.415 | 154.0843 | 0.27099
2020-02-11 | 13:47 | 5.322 | 5.324 | 5.323 | 567.5874 | 1.41 | 1.412 | 1.411 | 153.6487 | 0.2707
2020-02-11 | 13:47 | 5.308 | 5.32 | 5.314 | 566.6277 | 1.407 | 1.412 | 1.4095 | 153.4854 | 0.27088
2020-02-11 | 13:47 | 5.31 | 5.318 | 5.314 | 566.6277 | 1.408 | 1.411 | 1.4095 | 153.4854 | 0.27088
2020-02-11 | 13:47 | 5.321 | 5.334 | 5.3275 | 568.0672 | 1.412 | 1.415 | 1.4135 | 153.921 | 0.27096
2020-02-11 | 13:47 | 5.32 | 5.334 | 5.327 | 568.0139 | 1.411 | 1.415 | 1.413 | 153.8665 | 0.27089
2020-02-11 | 13:47 | 5.3 | 5.316 | 5.308 | 565.9879 | 1.406 | 1.411 | 1.4085 | 153.3765 | 0.27099
2020-02-11 | 13:48 | 5.275 | 5.295 | 5.285 | 563.5354 | 1.399 | 1.405 | 1.402 | 152.6687 | 0.27091
2020-02-11 | 13:48 | 5.27 | 5.273 | 5.2715 | 562.0959 | 1.397 | 1.399 | 1.398 | 152.2331 | 0.27083
2020-02-11 | 13:48 | 5.275 | 5.286 | 5.2805 | 563.0556 | 1.398 | 1.402 | 1.4 | 152.4509 | 0.27076
2020-02-11 | 13:48 | 5.291 | 5.298 | 5.2945 | 564.5484 | 1.404 | 1.407 | 1.4055 | 153.0498 | 0.2711
2020-02-11 | 13:48 | 5.294 | 5.297 | 5.2955 | 564.6551 | 1.404 | 1.406 | 1.405 | 152.9954 | 0.27095
2020-02-11 | 13:48 | 5.294 | 5.307 | 5.3005 | 565.1882 | 1.406 | 1.41 | 1.408 | 153.3221 | 0.27128
2020-02-11 | 13:48 | 5.31 | 5.326 | 5.318 | 567.0542 | 1.412 | 1.415 | 1.4135 | 153.921 | 0.27144
2020-02-11 | 13:48 | 5.329 | 5.331 | 5.33 | 568.3338 | 1.416 | 1.418 | 1.417 | 154.3021 | 0.2715
2020-02-11 | 13:48 | 5.33 | 5.34 | 5.335 | 568.8669 | 1.415 | 1.42 | 1.4175 | 154.3565 | 0.27134
2020-02-11 | 13:48 | 5.341 | 5.346 | 5.3435 | 569.7733 | 1.42 | 1.421 | 1.4205 | 154.6832 | 0.27148
2020-02-11 | 13:48 | 5.346 | 5.357 | 5.3515 | 570.6263 | 1.421 | 1.424 | 1.4225 | 154.901 | 0.27146
2020-02-11 | 13:48 | 5.362 | 5.393 | 5.3775 | 573.3986 | 1.425 | 1.435 | 1.43 | 155.7177 | 0.27157
2020-02-11 | 13:48 | 5.398 | 5.422 | 5.41 | 576.8641 | 1.436 | 1.442 | 1.439 | 156.6977 | 0.27164
2020-02-11 | 13:49 | 5.422 | 5.437 | 5.4295 | 578.9434 | 1.443 | 1.449 | 1.446 | 157.46 | 0.27198
2020-02-11 | 13:49 | 5.445 | 5.475 | 5.46 | 582.1956 | 1.45 | 1.458 | 1.454 | 158.3312 | 0.27196
2020-02-11 | 13:49 | 5.483 | 5.501 | 5.492 | 585.6077 | 1.462 | 1.467 | 1.4645 | 159.4745 | 0.27232
2020-02-11 | 13:49 | 5.505 | 5.511 | 5.508 | 587.3138 | 1.468 | 1.47 | 1.469 | 159.9646 | 0.27237
2020-02-11 | 13:49 | 5.512 | 5.518 | 5.515 | 588.0602 | 1.471 | 1.472 | 1.4715 | 160.2368 | 0.27248
2020-02-11 | 13:49 | 5.521 | 5.528 | 5.5245 | 589.0731 | 1.472 | 1.474 | 1.473 | 160.4001 | 0.27229
2020-02-11 | 13:49 | 5.53 | 5.534 | 5.532 | 589.8729 | 1.473 | 1.475 | 1.474 | 160.509 | 0.27211
2020-02-11 | 13:49 | 5.533 | 5.538 | 5.5355 | 590.2461 | 1.475 | 1.477 | 1.476 | 160.7268 | 0.2723
2020-02-11 | 13:49 | 5.534 | 5.537 | 5.5355 | 590.2461 | 1.475 | 1.477 | 1.476 | 160.7268 | 0.2723
2020-02-11 | 13:49 | 5.526 | 5.533 | 5.5295 | 589.6063 | 1.472 | 1.475 | 1.4735 | 160.4546 | 0.27214
2020-02-11 | 13:49 | 5.523 | 5.526 | 5.5245 | 589.0731 | 1.472 | 1.473 | 1.4725 | 160.3457 | 0.2722
2020-02-11 | 13:50 | 5.518 | 5.523 | 5.5205 | 588.6466 | 1.47 | 1.472 | 1.471 | 160.1823 | 0.27212
2020-02-11 | 13:50 | 5.517 | 5.518 | 5.5175 | 588.3267 | 1.47 | 1.472 | 1.471 | 160.1823 | 0.27227
2020-02-11 | 13:50 | 5.512 | 5.518 | 5.515 | 588.0602 | 1.469 | 1.471 | 1.47 | 160.0734 | 0.27221
2020-02-11 | 13:50 | 5.507 | 5.511 | 5.509 | 587.4204 | 1.468 | 1.469 | 1.4685 | 159.9101 | 0.27222
2020-02-11 | 13:50 | 5.498 | 5.507 | 5.5025 | 586.7273 | 1.465 | 1.469 | 1.467 | 159.7468 | 0.27227
2020-02-11 | 13:50 | 5.492 | 5.496 | 5.494 | 585.821 | 1.463 | 1.464 | 1.4635 | 159.3656 | 0.27204
2020-02-11 | 13:50 | 5.495 | 5.503 | 5.499 | 586.3541 | 1.465 | 1.467 | 1.466 | 159.6379 | 0.27226
2020-02-11 | 13:50 | 5.506 | 5.509 | 5.5075 | 587.2604 | 1.466 | 1.467 | 1.4665 | 159.6923 | 0.27193
2020-02-11 | 13:50 | 5.511 | 5.519 | 5.515 | 588.0602 | 1.467 | 1.47 | 1.4685 | 159.9101 | 0.27193
2020-02-11 | 13:50 | 5.519 | 5.521 | 5.52 | 588.5933 | 1.47 | 1.471 | 1.4705 | 160.1279 | 0.27205
2020-02-11 | 13:50 | 5.519 | 5.525 | 5.522 | 588.8066 | 1.469 | 1.471 | 1.47 | 160.0734 | 0.27186
2020-02-11 | 13:50 | 5.515 | 5.516 | 5.5155 | 588.1135 | 1.466 | 1.469 | 1.4675 | 159.8012 | 0.27172
2020-02-11 | 13:50 | 5.516 | 5.521 | 5.5185 | 588.4334 | 1.468 | 1.47 | 1.469 | 159.9646 | 0.27185
2020-02-11 | 13:51 | 5.52 | 5.522 | 5.521 | 588.6999 | 1.469 | 1.47 | 1.4695 | 160.019 | 0.27182
2020-02-11 | 13:51 | 5.525 | 5.53 | 5.5275 | 589.393 | 1.471 | 1.472 | 1.4715 | 160.2368 | 0.27187
2020-02-11 | 13:51 | 5.529 | 5.531 | 5.53 | 589.6596 | 1.472 | 1.473 | 1.4725 | 160.3457 | 0.27193
2020-02-11 | 13:51 | 5.532 | 5.538 | 5.535 | 590.1927 | 1.474 | 1.476 | 1.475 | 160.6179 | 0.27214
2020-02-11 | 13:51 | 5.54 | 5.544 | 5.542 | 590.9392 | 1.475 | 1.476 | 1.4755 | 160.6724 | 0.27189
2020-02-11 | 13:51 | 5.541 | 5.547 | 5.544 | 591.1524 | 1.476 | 1.478 | 1.477 | 160.8357 | 0.27207
2020-02-11 | 13:51 | 5.552 | 5.555 | 5.5535 | 592.1654 | 1.479 | 1.481 | 1.48 | 161.1624 | 0.27216
2020-02-11 | 13:51 | 5.557 | 5.563 | 5.56 | 592.8585 | 1.481 | 1.483 | 1.482 | 161.3802 | 0.27221
2020-02-11 | 13:51 | 5.564 | 5.581 | 5.5725 | 594.1913 | 1.484 | 1.488 | 1.486 | 161.8157 | 0.27233
2020-02-11 | 13:51 | 5.577 | 5.582 | 5.5795 | 594.9377 | 1.488 | 1.49 | 1.489 | 162.1424 | 0.27254
2020-02-11 | 13:51 | 5.577 | 5.588 | 5.5825 | 595.2576 | 1.488 | 1.49 | 1.489 | 162.1424 | 0.27239
2020-02-11 | 13:52 | 5.592 | 5.596 | 5.594 | 596.4839 | 1.491 | 1.494 | 1.4925 | 162.5236 | 0.27247
2020-02-11 | 13:52 | 5.596 | 5.599 | 5.5975 | 596.8571 | 1.493 | 1.494 | 1.4935 | 162.6324 | 0.27248
2020-02-11 | 13:52 | 5.6 | 5.609 | 5.6045 | 597.6035 | 1.493 | 1.496 | 1.4945 | 162.7413 | 0.27232
2020-02-11 | 13:52 | 5.613 | 5.619 | 5.616 | 598.8297 | 1.497 | 1.501 | 1.499 | 163.2314 | 0.27258
2020-02-11 | 13:52 | 5.62 | 5.622 | 5.621 | 599.3629 | 1.5 | 1.502 | 1.501 | 163.4491 | 0.2727
2020-02-11 | 13:52 | 5.621 | 5.623 | 5.622 | 599.4695 | 1.5 | 1.503 | 1.5015 | 163.5036 | 0.27275
2020-02-11 | 13:52 | 5.619 | 5.621 | 5.62 | 599.2562 | 1.5 | 1.501 | 1.5005 | 163.3947 | 0.27266
2020-02-11 | 13:52 | 5.623 | 5.631 | 5.627 | 600.0026 | 1.501 | 1.503 | 1.502 | 163.558 | 0.2726
2020-02-11 | 13:52 | 5.633 | 5.649 | 5.641 | 601.4954 | 1.503 | 1.507 | 1.505 | 163.8847 | 0.27246
2020-02-11 | 13:52 | 5.639 | 5.654 | 5.6465 | 602.0819 | 1.505 | 1.51 | 1.5075 | 164.157 | 0.27265
2020-02-11 | 13:52 | 5.638 | 5.645 | 5.6415 | 601.5488 | 1.504 | 1.507 | 1.5055 | 163.9392 | 0.27253
2020-02-11 | 13:52 | 5.65 | 5.655 | 5.6525 | 602.7217 | 1.506 | 1.51 | 1.508 | 164.2114 | 0.27245
2020-02-11 | 13:52 | 5.639 | 5.649 | 5.644 | 601.8153 | 1.505 | 1.508 | 1.5065 | 164.0481 | 0.27259
2020-02-11 | 13:53 | 5.638 | 5.641 | 5.6395 | 601.3355 | 1.504 | 1.505 | 1.5045 | 163.8303 | 0.27244
2020-02-11 | 13:53 | 5.636 | 5.641 | 5.6385 | 601.2289 | 1.503 | 1.506 | 1.5045 | 163.8303 | 0.27249
2020-02-11 | 13:53 | 5.624 | 5.631 | 5.6275 | 600.056 | 1.5 | 1.502 | 1.501 | 163.4491 | 0.27239
2020-02-11 | 13:53 | 5.622 | 5.626 | 5.624 | 599.6827 | 1.499 | 1.501 | 1.5 | 163.3403 | 0.27238
2020-02-11 | 13:53 | 5.628 | 5.632 | 5.63 | 600.3225 | 1.501 | 1.504 | 1.5025 | 163.6125 | 0.27254
2020-02-11 | 13:53 | 5.63 | 5.633 | 5.6315 | 600.4825 | 1.502 | 1.503 | 1.5025 | 163.6125 | 0.27247
2020-02-11 | 13:53 | 5.621 | 5.628 | 5.6245 | 599.7361 | 1.501 | 1.502 | 1.5015 | 163.5036 | 0.27263
2020-02-11 | 13:53 | 5.614 | 5.621 | 5.6175 | 598.9897 | 1.498 | 1.5 | 1.499 | 163.2314 | 0.27251
2020-02-11 | 13:53 | 5.615 | 5.622 | 5.6185 | 599.0963 | 1.499 | 1.501 | 1.5 | 163.3403 | 0.27264
2020-02-11 | 13:53 | 5.624 | 5.628 | 5.626 | 599.896 | 1.502 | 1.504 | 1.503 | 163.6669 | 0.27283
2020-02-11 | 13:53 | 5.623 | 5.63 | 5.6265 | 599.9493 | 1.502 | 1.504 | 1.503 | 163.6669 | 0.2728
2020-02-11 | 13:54 | 5.623 | 5.631 | 5.627 | 600.0026 | 1.503 | 1.504 | 1.5035 | 163.7214 | 0.27287
2020-02-11 | 13:54 | 5.621 | 5.63 | 5.6255 | 599.8427 | 1.501 | 1.504 | 1.5025 | 163.6125 | 0.27276
2020-02-11 | 13:54 | 5.614 | 5.616 | 5.615 | 598.7231 | 1.5 | 1.501 | 1.5005 | 163.3947 | 0.27291
2020-02-11 | 13:54 | 5.607 | 5.615 | 5.611 | 598.2966 | 1.498 | 1.5 | 1.499 | 163.2314 | 0.27283
2020-02-11 | 13:54 | 5.598 | 5.603 | 5.6005 | 597.177 | 1.495 | 1.497 | 1.496 | 162.9047 | 0.27279
2020-02-11 | 13:54 | 5.588 | 5.598 | 5.593 | 596.3772 | 1.493 | 1.496 | 1.4945 | 162.7413 | 0.27288
2020-02-11 | 13:54 | 5.584 | 5.587 | 5.5855 | 595.5775 | 1.491 | 1.493 | 1.492 | 162.4691 | 0.27279
2020-02-11 | 13:54 | 5.573 | 5.588 | 5.5805 | 595.0444 | 1.489 | 1.494 | 1.4915 | 162.4147 | 0.27295
2020-02-11 | 13:54 | 5.55 | 5.567 | 5.5585 | 592.6985 | 1.483 | 1.488 | 1.4855 | 161.7613 | 0.27292
2020-02-11 | 13:54 | 5.537 | 5.549 | 5.543 | 591.0458 | 1.48 | 1.483 | 1.4815 | 161.3257 | 0.27295
2020-02-11 | 13:54 | 5.521 | 5.533 | 5.527 | 589.3397 | 1.475 | 1.479 | 1.477 | 160.8357 | 0.27291
2020-02-11 | 13:54 | 5.519 | 5.522 | 5.5205 | 588.6466 | 1.474 | 1.477 | 1.4755 | 160.6724 | 0.27295
2020-02-11 | 13:54 | 5.516 | 5.518 | 5.517 | 588.2734 | 1.474 | 1.475 | 1.4745 | 160.5635 | 0.27294
2020-02-11 | 13:55 | 5.522 | 5.526 | 5.524 | 589.0198 | 1.476 | 1.477 | 1.4765 | 160.7813 | 0.27296
2020-02-11 | 13:55 | 5.521 | 5.527 | 5.524 | 589.0198 | 1.475 | 1.477 | 1.476 | 160.7268 | 0.27287
2020-02-11 | 13:55 | 5.521 | 5.526 | 5.5235 | 588.9665 | 1.476 | 1.477 | 1.4765 | 160.7813 | 0.27299
2020-02-11 | 13:55 | 5.518 | 5.527 | 5.5225 | 588.8599 | 1.474 | 1.476 | 1.475 | 160.6179 | 0.27276
2020-02-11 | 13:55 | 5.511 | 5.515 | 5.513 | 587.8469 | 1.471 | 1.474 | 1.4725 | 160.3457 | 0.27277
2020-02-11 | 13:55 | 5.512 | 5.517 | 5.5145 | 588.0069 | 1.473 | 1.475 | 1.474 | 160.509 | 0.27297
2020-02-11 | 13:55 | 5.521 | 5.535 | 5.528 | 589.4463 | 1.474 | 1.481 | 1.4775 | 160.8901 | 0.27295
2020-02-11 | 13:55 | 5.54 | 5.551 | 5.5455 | 591.3124 | 1.481 | 1.484 | 1.4825 | 161.4346 | 0.27301
2020-02-11 | 13:55 | 5.553 | 5.557 | 5.555 | 592.3253 | 1.485 | 1.488 | 1.4865 | 161.8702 | 0.27328
2020-02-11 | 13:55 | 5.541 | 5.548 | 5.5445 | 591.2057 | 1.481 | 1.486 | 1.4835 | 161.5435 | 0.27324
2020-02-11 | 13:55 | 5.543 | 5.547 | 5.545 | 591.259 | 1.481 | 1.483 | 1.482 | 161.3802 | 0.27294
2020-02-11 | 13:56 | 5.538 | 5.544 | 5.541 | 590.8325 | 1.481 | 1.482 | 1.4815 | 161.3257 | 0.27305
2020-02-11 | 13:56 | 5.536 | 5.54 | 5.538 | 590.5126 | 1.481 | 1.482 | 1.4815 | 161.3257 | 0.2732
2020-02-11 | 13:56 | 5.543 | 5.552 | 5.5475 | 591.5256 | 1.483 | 1.485 | 1.484 | 161.598 | 0.27319
2020-02-11 | 13:56 | 5.554 | 5.561 | 5.5575 | 592.5919 | 1.484 | 1.488 | 1.486 | 161.8157 | 0.27306
2020-02-11 | 13:56 | 5.56 | 5.562 | 5.561 | 592.9651 | 1.486 | 1.487 | 1.4865 | 161.8702 | 0.27298
2020-02-11 | 13:56 | 5.556 | 5.559 | 5.5575 | 592.5919 | 1.486 | 1.487 | 1.4865 | 161.8702 | 0.27316
2020-02-11 | 13:56 | 5.549 | 5.556 | 5.5525 | 592.0588 | 1.484 | 1.487 | 1.4855 | 161.7613 | 0.27322
2020-02-11 | 13:56 | 5.552 | 5.557 | 5.5545 | 592.272 | 1.485 | 1.487 | 1.486 | 161.8157 | 0.27321
2020-02-11 | 13:56 | 5.555 | 5.559 | 5.557 | 592.5386 | 1.485 | 1.486 | 1.4855 | 161.7613 | 0.273
2020-02-11 | 13:56 | 5.56 | 5.566 | 5.563 | 593.1784 | 1.487 | 1.488 | 1.4875 | 161.9791 | 0.27307
2020-02-11 | 13:56 | 5.568 | 5.576 | 5.572 | 594.138 | 1.489 | 1.491 | 1.49 | 162.2513 | 0.27309
2020-02-11 | 13:56 | 5.574 | 5.577 | 5.5755 | 594.5112 | 1.49 | 1.491 | 1.4905 | 162.3058 | 0.27301
2020-02-11 | 13:56 | 5.574 | 5.578 | 5.576 | 594.5645 | 1.491 | 1.492 | 1.4915 | 162.4147 | 0.27317
2020-02-11 | 13:57 | 5.571 | 5.575 | 5.573 | 594.2447 | 1.49 | 1.491 | 1.4905 | 162.3058 | 0.27313
2020-02-11 | 13:57 | 5.57 | 5.574 | 5.572 | 594.138 | 1.489 | 1.49 | 1.4895 | 162.1969 | 0.273
2020-02-11 | 13:57 | 5.557 | 5.566 | 5.5615 | 593.0184 | 1.485 | 1.488 | 1.4865 | 161.8702 | 0.27296
2020-02-11 | 13:57 | 5.563 | 5.568 | 5.5655 | 593.4449 | 1.488 | 1.491 | 1.4895 | 162.1969 | 0.27331
2020-02-11 | 13:57 | 5.565 | 5.568 | 5.5665 | 593.5516 | 1.488 | 1.49 | 1.489 | 162.1424 | 0.27317
2020-02-11 | 13:57 | 5.566 | 5.571 | 5.5685 | 593.7648 | 1.489 | 1.491 | 1.49 | 162.2513 | 0.27326
2020-02-11 | 13:57 | 5.571 | 5.572 | 5.5715 | 594.0847 | 1.49 | 1.492 | 1.491 | 162.3602 | 0.27329
2020-02-11 | 13:57 | 5.569 | 5.573 | 5.571 | 594.0314 | 1.491 | 1.492 | 1.4915 | 162.4147 | 0.27341
2020-02-11 | 13:57 | 5.554 | 5.566 | 5.56 | 592.8585 | 1.488 | 1.49 | 1.489 | 162.1424 | 0.27349
2020-02-11 | 13:57 | 5.546 | 5.551 | 5.5485 | 591.6322 | 1.485 | 1.487 | 1.486 | 161.8157 | 0.27351
2020-02-11 | 13:57 | 5.548 | 5.553 | 5.5505 | 591.8455 | 1.485 | 1.486 | 1.4855 | 161.7613 | 0.27332
2020-02-11 | 13:58 | 5.551 | 5.553 | 5.552 | 592.0054 | 1.485 | 1.487 | 1.486 | 161.8157 | 0.27333
2020-02-11 | 13:58 | 5.551 | 5.559 | 5.555 | 592.3253 | 1.485 | 1.486 | 1.4855 | 161.7613 | 0.2731
2020-02-11 | 13:58 | 5.559 | 5.562 | 5.5605 | 592.9118 | 1.486 | 1.488 | 1.487 | 161.9246 | 0.2731
2020-02-11 | 13:58 | 5.554 | 5.557 | 5.5555 | 592.3786 | 1.485 | 1.486 | 1.4855 | 161.7613 | 0.27307
2020-02-11 | 13:58 | 5.552 | 5.559 | 5.5555 | 592.3786 | 1.485 | 1.486 | 1.4855 | 161.7613 | 0.27307
2020-02-11 | 13:58 | 5.551 | 5.556 | 5.5535 | 592.1654 | 1.485 | 1.486 | 1.4855 | 161.7613 | 0.27317
2020-02-11 | 13:58 | 5.551 | 5.556 | 5.5535 | 592.1654 | 1.486 | 1.486 | 1.486 | 161.8157 | 0.27326
2020-02-11 | 13:58 | 5.543 | 5.549 | 5.546 | 591.3657 | 1.483 | 1.486 | 1.4845 | 161.6524 | 0.27335
2020-02-11 | 13:58 | 5.536 | 5.542 | 5.539 | 590.6193 | 1.481 | 1.484 | 1.4825 | 161.4346 | 0.27333
2020-02-11 | 13:58 | 5.532 | 5.536 | 5.534 | 590.0861 | 1.481 | 1.482 | 1.4815 | 161.3257 | 0.27339
2020-02-11 | 13:58 | 5.529 | 5.532 | 5.5305 | 589.7129 | 1.48 | 1.481 | 1.4805 | 161.2168 | 0.27338
2020-02-11 | 13:58 | 5.529 | 5.532 | 5.5305 | 589.7129 | 1.48 | 1.481 | 1.4805 | 161.2168 | 0.27338
2020-02-11 | 13:58 | 5.53 | 5.531 | 5.5305 | 589.7129 | 1.479 | 1.48 | 1.4795 | 161.1079 | 0.2732
2020-02-11 | 13:59 | 5.527 | 5.529 | 5.528 | 589.4463 | 1.477 | 1.48 | 1.4785 | 160.999 | 0.27314
2020-02-11 | 13:59 | 5.523 | 5.528 | 5.5255 | 589.1798 | 1.478 | 1.479 | 1.4785 | 160.999 | 0.27326
2020-02-11 | 13:59 | 5.521 | 5.524 | 5.5225 | 588.8599 | 1.478 | 1.479 | 1.4785 | 160.999 | 0.27341
2020-02-11 | 13:59 | 5.507 | 5.519 | 5.513 | 587.8469 | 1.474 | 1.477 | 1.4755 | 160.6724 | 0.27332
2020-02-11 | 13:59 | 5.505 | 5.507 | 5.506 | 587.1005 | 1.471 | 1.473 | 1.472 | 160.2912 | 0.27302
2020-02-11 | 13:59 | 5.5 | 5.505 | 5.5025 | 586.7273 | 1.471 | 1.473 | 1.472 | 160.2912 | 0.2732
2020-02-11 | 13:59 | 5.495 | 5.498 | 5.4965 | 586.0875 | 1.469 | 1.472 | 1.4705 | 160.1279 | 0.27321
2020-02-11 | 13:59 | 5.495 | 5.503 | 5.499 | 586.3541 | 1.471 | 1.472 | 1.4715 | 160.2368 | 0.27328
2020-02-11 | 13:59 | 5.503 | 5.506 | 5.5045 | 586.9406 | 1.472 | 1.473 | 1.4725 | 160.3457 | 0.27319
2020-02-11 | 13:59 | 5.498 | 5.501 | 5.4995 | 586.4074 | 1.47 | 1.473 | 1.4715 | 160.2368 | 0.27325
2020-02-11 | 13:59 | 5.492 | 5.499 | 5.4955 | 585.9809 | 1.468 | 1.471 | 1.4695 | 160.019 | 0.27308
2020-02-11 | 14:00 | 5.487 | 5.491 | 5.489 | 585.2878 | 1.467 | 1.469 | 1.468 | 159.8557 | 0.27312
2020-02-11 | 14:00 | 5.483 | 5.487 | 5.485 | 584.8613 | 1.467 | 1.468 | 1.4675 | 159.8012 | 0.27323
2020-02-11 | 14:00 | 5.48 | 5.483 | 5.4815 | 584.4881 | 1.465 | 1.467 | 1.466 | 159.6379 | 0.27312
2020-02-11 | 14:00 | 5.483 | 5.485 | 5.484 | 584.7547 | 1.466 | 1.468 | 1.467 | 159.7468 | 0.27319
2020-02-11 | 14:00 | 5.482 | 5.484 | 5.483 | 584.648 | 1.465 | 1.466 | 1.4655 | 159.5834 | 0.27296
2020-02-11 | 14:00 | 5.482 | 5.486 | 5.484 | 584.7547 | 1.466 | 1.468 | 1.467 | 159.7468 | 0.27319
2020-02-11 | 14:00 | 5.484 | 5.486 | 5.485 | 584.8613 | 1.467 | 1.468 | 1.4675 | 159.8012 | 0.27323
2020-02-11 | 14:00 | 5.487 | 5.489 | 5.488 | 585.1812 | 1.467 | 1.469 | 1.468 | 159.8557 | 0.27317
2020-02-11 | 14:00 | 5.486 | 5.489 | 5.4875 | 585.1279 | 1.468 | 1.469 | 1.4685 | 159.9101 | 0.27329
2020-02-11 | 14:00 | 5.488 | 5.489 | 5.4885 | 585.2345 | 1.468 | 1.469 | 1.4685 | 159.9101 | 0.27324
2020-02-11 | 14:00 | 5.484 | 5.488 | 5.486 | 584.9679 | 1.468 | 1.468 | 1.468 | 159.8557 | 0.27327
2020-02-11 | 14:00 | 5.48 | 5.484 | 5.482 | 584.5414 | 1.466 | 1.467 | 1.4665 | 159.6923 | 0.27319
2020-02-11 | 14:00 | 5.475 | 5.477 | 5.476 | 583.9016 | 1.464 | 1.464 | 1.464 | 159.4201 | 0.27303
2020-02-11 | 14:01 | 5.461 | 5.472 | 5.4665 | 582.8886 | 1.462 | 1.465 | 1.4635 | 159.3656 | 0.27341
2020-02-11 | 14:01 | 5.448 | 5.458 | 5.453 | 581.4492 | 1.457 | 1.462 | 1.4595 | 158.9301 | 0.27333
2020-02-11 | 14:01 | 5.438 | 5.445 | 5.4415 | 580.2229 | 1.455 | 1.457 | 1.456 | 158.5489 | 0.27326
2020-02-11 | 14:01 | 5.385 | 5.436 | 5.4105 | 576.9174 | 1.454 | 1.456 | 1.455 | 158.44 | 0.27463
2020-02-11 | 14:01 | 5.339 | 5.417 | 5.378 | 573.452 | 1.452 | 1.454 | 1.453 | 158.2223 | 0.27591
2020-02-11 | 14:01 | 5.413 | 5.434 | 5.4235 | 578.3036 | 1.452 | 1.454 | 1.453 | 158.2223 | 0.2736
2020-02-11 | 14:01 | 5.428 | 5.433 | 5.4305 | 579.05 | 1.452 | 1.455 | 1.4535 | 158.2767 | 0.27334
2020-02-11 | 14:01 | 5.413 | 5.425 | 5.419 | 577.8238 | 1.45 | 1.452 | 1.451 | 158.0045 | 0.27345
2020-02-11 | 14:01 | 5.396 | 5.41 | 5.403 | 576.1177 | 1.444 | 1.448 | 1.446 | 157.46 | 0.27331
2020-02-11 | 14:01 | 5.379 | 5.39 | 5.3845 | 574.1451 | 1.439 | 1.443 | 1.441 | 156.9155 | 0.2733
2020-02-11 | 14:01 | 5.377 | 5.38 | 5.3785 | 573.5053 | 1.44 | 1.441 | 1.4405 | 156.8611 | 0.27351
2020-02-11 | 14:02 | 5.38 | 5.385 | 5.3825 | 573.9318 | 1.44 | 1.441 | 1.4405 | 156.8611 | 0.27331
2020-02-11 | 14:02 | 5.386 | 5.388 | 5.387 | 574.4116 | 1.441 | 1.442 | 1.4415 | 156.97 | 0.27327
2020-02-11 | 14:02 | 5.369 | 5.384 | 5.3765 | 573.292 | 1.437 | 1.443 | 1.44 | 156.8066 | 0.27352
2020-02-11 | 14:02 | 5.352 | 5.366 | 5.359 | 571.426 | 1.433 | 1.436 | 1.4345 | 156.2077 | 0.27336
2020-02-11 | 14:02 | 5.34 | 5.348 | 5.344 | 569.8266 | 1.429 | 1.433 | 1.431 | 155.8266 | 0.27346
2020-02-11 | 14:02 | 5.332 | 5.337 | 5.3345 | 568.8136 | 1.427 | 1.429 | 1.428 | 155.4999 | 0.27338
2020-02-11 | 14:02 | 5.32 | 5.33 | 5.325 | 567.8006 | 1.424 | 1.426 | 1.425 | 155.1732 | 0.27329
2020-02-11 | 14:02 | 5.312 | 5.318 | 5.315 | 566.7343 | 1.422 | 1.424 | 1.423 | 154.9555 | 0.27342
2020-02-11 | 14:02 | 5.304 | 5.312 | 5.308 | 565.9879 | 1.419 | 1.421 | 1.42 | 154.6288 | 0.2732
2020-02-11 | 14:02 | 5.295 | 5.303 | 5.299 | 565.0283 | 1.417 | 1.42 | 1.4185 | 154.4654 | 0.27338
2020-02-11 | 14:02 | 5.282 | 5.291 | 5.2865 | 563.6954 | 1.414 | 1.416 | 1.415 | 154.0843 | 0.27335
2020-02-11 | 14:02 | 5.277 | 5.281 | 5.279 | 562.8957 | 1.411 | 1.415 | 1.413 | 153.8665 | 0.27335
2020-02-11 | 14:02 | 5.252 | 5.272 | 5.262 | 561.083 | 1.405 | 1.41 | 1.4075 | 153.2676 | 0.27316
2020-02-11 | 14:03 | 5.241 | 5.25 | 5.2455 | 559.3236 | 1.401 | 1.404 | 1.4025 | 152.7231 | 0.27305
2020-02-11 | 14:03 | 5.24 | 5.254 | 5.247 | 559.4835 | 1.4 | 1.404 | 1.402 | 152.6687 | 0.27287
2020-02-11 | 14:03 | 5.256 | 5.267 | 5.2615 | 561.0297 | 1.405 | 1.409 | 1.407 | 153.2132 | 0.27309
2020-02-11 | 14:03 | 5.268 | 5.269 | 5.2685 | 561.7761 | 1.409 | 1.411 | 1.41 | 153.5398 | 0.27331
2020-02-11 | 14:03 | 5.269 | 5.269 | 5.269 | 561.8294 | 1.41 | 1.411 | 1.4105 | 153.5943 | 0.27338
2020-02-11 | 14:03 | 5.264 | 5.268 | 5.266 | 561.5095 | 1.409 | 1.411 | 1.41 | 153.5398 | 0.27344
2020-02-11 | 14:03 | 5.253 | 5.262 | 5.2575 | 560.6031 | 1.407 | 1.409 | 1.408 | 153.3221 | 0.27349
2020-02-11 | 14:03 | 5.241 | 5.252 | 5.2465 | 559.4302 | 1.404 | 1.407 | 1.4055 | 153.0498 | 0.27358
2020-02-11 | 14:03 | 5.23 | 5.238 | 5.234 | 558.0974 | 1.401 | 1.402 | 1.4015 | 152.6142 | 0.27345
2020-02-11 | 14:03 | 5.218 | 5.227 | 5.2225 | 556.8711 | 1.397 | 1.399 | 1.398 | 152.2331 | 0.27337
2020-02-11 | 14:03 | 5.216 | 5.221 | 5.2185 | 556.4446 | 1.396 | 1.397 | 1.3965 | 152.0698 | 0.27329
2020-02-11 | 14:04 | 5.222 | 5.227 | 5.2245 | 557.0844 | 1.398 | 1.399 | 1.3985 | 152.2876 | 0.27337
2020-02-11 | 14:04 | 5.213 | 5.221 | 5.217 | 556.2847 | 1.395 | 1.396 | 1.3955 | 151.9609 | 0.27317
2020-02-11 | 14:04 | 5.211 | 5.216 | 5.2135 | 555.9115 | 1.394 | 1.396 | 1.395 | 151.9064 | 0.27326
2020-02-11 | 14:04 | 5.211 | 5.217 | 5.214 | 555.9648 | 1.395 | 1.398 | 1.3965 | 152.0698 | 0.27352
2020-02-11 | 14:04 | 5.202 | 5.209 | 5.2055 | 555.0584 | 1.392 | 1.396 | 1.394 | 151.7975 | 0.27348
2020-02-11 | 14:04 | 5.192 | 5.198 | 5.195 | 553.9388 | 1.39 | 1.392 | 1.391 | 151.4709 | 0.27344
2020-02-11 | 14:04 | 5.184 | 5.191 | 5.1875 | 553.1391 | 1.385 | 1.39 | 1.3875 | 151.0897 | 0.27315
2020-02-11 | 14:04 | 5.169 | 5.182 | 5.1755 | 551.8595 | 1.384 | 1.387 | 1.3855 | 150.8719 | 0.27339
2020-02-11 | 14:04 | 5.16 | 5.168 | 5.164 | 550.6333 | 1.381 | 1.384 | 1.3825 | 150.5453 | 0.2734
2020-02-11 | 14:04 | 5.154 | 5.159 | 5.1565 | 549.8336 | 1.38 | 1.381 | 1.3805 | 150.3275 | 0.27341
2020-02-11 | 14:04 | 5.154 | 5.156 | 5.155 | 549.6736 | 1.379 | 1.381 | 1.38 | 150.273 | 0.27339
2020-02-11 | 14:04 | 5.154 | 5.156 | 5.155 | 549.6736 | 1.379 | 1.38 | 1.3795 | 150.2186 | 0.27329
2020-02-11 | 14:04 | 5.152 | 5.154 | 5.153 | 549.4604 | 1.378 | 1.38 | 1.379 | 150.1641 | 0.27329
2020-02-11 | 14:05 | 5.153 | 5.156 | 5.1545 | 549.6203 | 1.378 | 1.381 | 1.3795 | 150.2186 | 0.27331
2020-02-11 | 14:05 | 5.155 | 5.156 | 5.1555 | 549.727 | 1.378 | 1.381 | 1.3795 | 150.2186 | 0.27326
2020-02-11 | 14:05 | 5.156 | 5.158 | 5.157 | 549.8869 | 1.379 | 1.38 | 1.3795 | 150.2186 | 0.27318
2020-02-11 | 14:05 | 5.161 | 5.168 | 5.1645 | 550.6866 | 1.381 | 1.383 | 1.382 | 150.4908 | 0.27328
2020-02-11 | 14:05 | 5.169 | 5.175 | 5.172 | 551.4863 | 1.382 | 1.385 | 1.3835 | 150.6542 | 0.27318
2020-02-11 | 14:05 | 5.175 | 5.182 | 5.1785 | 552.1794 | 1.385 | 1.388 | 1.3865 | 150.9808 | 0.27343
2020-02-11 | 14:05 | 5.184 | 5.19 | 5.187 | 553.0858 | 1.389 | 1.39 | 1.3895 | 151.3075 | 0.27357
2020-02-11 | 14:05 | 5.182 | 5.185 | 5.1835 | 552.7126 | 1.387 | 1.389 | 1.388 | 151.1442 | 0.27346
2020-02-11 | 14:05 | 5.185 | 5.186 | 5.1855 | 552.9258 | 1.389 | 1.39 | 1.3895 | 151.3075 | 0.27365
2020-02-11 | 14:05 | 5.188 | 5.2 | 5.194 | 553.8322 | 1.39 | 1.392 | 1.391 | 151.4709 | 0.2735
2020-02-11 | 14:05 | 5.204 | 5.215 | 5.2095 | 555.4849 | 1.394 | 1.398 | 1.396 | 152.0153 | 0.27366
2020-02-11 | 14:06 | 5.217 | 5.226 | 5.2215 | 556.7645 | 1.397 | 1.4 | 1.3985 | 152.2876 | 0.27352
2020-02-11 | 14:06 | 5.232 | 5.238 | 5.235 | 558.204 | 1.402 | 1.404 | 1.403 | 152.7776 | 0.27369
2020-02-11 | 14:06 | 5.238 | 5.241 | 5.2395 | 558.6838 | 1.404 | 1.407 | 1.4055 | 153.0498 | 0.27395
2020-02-11 | 14:06 | 5.243 | 5.252 | 5.2475 | 559.5368 | 1.406 | 1.408 | 1.407 | 153.2132 | 0.27382
2020-02-11 | 14:06 | 5.255 | 5.259 | 5.257 | 560.5498 | 1.409 | 1.411 | 1.41 | 153.5398 | 0.27391
2020-02-11 | 14:06 | 5.257 | 5.259 | 5.258 | 560.6565 | 1.408 | 1.411 | 1.4095 | 153.4854 | 0.27376
2020-02-11 | 14:06 | 5.253 | 5.256 | 5.2545 | 560.2833 | 1.408 | 1.409 | 1.4085 | 153.3765 | 0.27375
2020-02-11 | 14:06 | 5.258 | 5.267 | 5.2625 | 561.1363 | 1.409 | 1.412 | 1.4105 | 153.5943 | 0.27372
2020-02-11 | 14:06 | 5.269 | 5.279 | 5.274 | 562.3625 | 1.412 | 1.415 | 1.4135 | 153.921 | 0.2737
2020-02-11 | 14:06 | 5.277 | 5.281 | 5.279 | 562.8957 | 1.415 | 1.416 | 1.4155 | 154.1388 | 0.27383
2020-02-11 | 14:06 | 5.264 | 5.274 | 5.269 | 561.8294 | 1.41 | 1.416 | 1.413 | 153.8665 | 0.27387
2020-02-11 | 14:06 | 5.256 | 5.261 | 5.2585 | 560.7098 | 1.409 | 1.41 | 1.4095 | 153.4854 | 0.27373
2020-02-11 | 14:06 | 5.256 | 5.257 | 5.2565 | 560.4965 | 1.407 | 1.408 | 1.4075 | 153.2676 | 0.27345
2020-02-11 | 14:07 | 5.251 | 5.255 | 5.253 | 560.1233 | 1.407 | 1.408 | 1.4075 | 153.2676 | 0.27363
2020-02-11 | 14:07 | 5.252 | 5.253 | 5.2525 | 560.07 | 1.407 | 1.407 | 1.407 | 153.2132 | 0.27356
2020-02-11 | 14:07 | 5.248 | 5.252 | 5.25 | 559.8034 | 1.406 | 1.407 | 1.4065 | 153.1587 | 0.27359
2020-02-11 | 14:07 | 5.246 | 5.249 | 5.2475 | 559.5368 | 1.405 | 1.406 | 1.4055 | 153.0498 | 0.27353
2020-02-11 | 14:07 | 5.253 | 5.254 | 5.2535 | 560.1766 | 1.405 | 1.407 | 1.406 | 153.1043 | 0.27331
2020-02-11 | 14:07 | 5.252 | 5.254 | 5.253 | 560.1233 | 1.404 | 1.407 | 1.4055 | 153.0498 | 0.27324
2020-02-11 | 14:07 | 5.252 | 5.268 | 5.26 | 560.8697 | 1.407 | 1.411 | 1.409 | 153.4309 | 0.27356
2020-02-11 | 14:07 | 5.271 | 5.28 | 5.2755 | 562.5225 | 1.412 | 1.415 | 1.4135 | 153.921 | 0.27363
2020-02-11 | 14:07 | 5.277 | 5.28 | 5.2785 | 562.8424 | 1.412 | 1.415 | 1.4135 | 153.921 | 0.27347
2020-02-11 | 14:07 | 5.272 | 5.276 | 5.274 | 562.3625 | 1.41 | 1.412 | 1.411 | 153.6487 | 0.27322
2020-02-11 | 14:07 | 5.266 | 5.27 | 5.268 | 561.7227 | 1.41 | 1.411 | 1.4105 | 153.5943 | 0.27343
2020-02-11 | 14:08 | 5.26 | 5.266 | 5.263 | 561.1896 | 1.408 | 1.411 | 1.4095 | 153.4854 | 0.2735
2020-02-11 | 14:08 | 5.252 | 5.258 | 5.255 | 560.3366 | 1.406 | 1.409 | 1.4075 | 153.2676 | 0.27353
2020-02-11 | 14:08 | 5.243 | 5.251 | 5.247 | 559.4835 | 1.403 | 1.406 | 1.4045 | 152.9409 | 0.27336
2020-02-11 | 14:08 | 5.227 | 5.238 | 5.2325 | 557.9374 | 1.398 | 1.402 | 1.4 | 152.4509 | 0.27324
2020-02-11 | 14:08 | 5.224 | 5.227 | 5.2255 | 557.191 | 1.398 | 1.399 | 1.3985 | 152.2876 | 0.27331
2020-02-11 | 14:08 | 5.226 | 5.227 | 5.2265 | 557.2976 | 1.397 | 1.399 | 1.398 | 152.2331 | 0.27316
2020-02-11 | 14:08 | 5.224 | 5.226 | 5.225 | 557.1377 | 1.397 | 1.399 | 1.398 | 152.2331 | 0.27324
2020-02-11 | 14:08 | 5.228 | 5.231 | 5.2295 | 557.6175 | 1.398 | 1.4 | 1.399 | 152.342 | 0.2732
2020-02-11 | 14:08 | 5.233 | 5.235 | 5.234 | 558.0974 | 1.401 | 1.401 | 1.401 | 152.5598 | 0.27336
2020-02-11 | 14:08 | 5.232 | 5.233 | 5.2325 | 557.9374 | 1.4 | 1.402 | 1.401 | 152.5598 | 0.27344
2020-02-11 | 14:08 | 5.235 | 5.246 | 5.2405 | 558.7904 | 1.402 | 1.406 | 1.404 | 152.8865 | 0.2736
2020-02-11 | 14:08 | 5.248 | 5.255 | 5.2515 | 559.9634 | 1.406 | 1.409 | 1.4075 | 153.2676 | 0.27371
2020-02-11 | 14:09 | 5.256 | 5.259 | 5.2575 | 560.6031 | 1.408 | 1.41 | 1.409 | 153.4309 | 0.27369
2020-02-11 | 14:09 | 5.248 | 5.258 | 5.253 | 560.1233 | 1.406 | 1.409 | 1.4075 | 153.2676 | 0.27363
2020-02-11 | 14:09 | 5.235 | 5.243 | 5.239 | 558.6305 | 1.401 | 1.404 | 1.4025 | 152.7231 | 0.27339
2020-02-11 | 14:09 | 5.235 | 5.236 | 5.2355 | 558.2573 | 1.401 | 1.402 | 1.4015 | 152.6142 | 0.27338
2020-02-11 | 14:09 | 5.233 | 5.236 | 5.2345 | 558.1507 | 1.401 | 1.402 | 1.4015 | 152.6142 | 0.27343
2020-02-11 | 14:09 | 5.206 | 5.228 | 5.217 | 556.2847 | 1.393 | 1.4 | 1.3965 | 152.0698 | 0.27337
Average |  |  |  |  | 647.6788 |  |  |  | 173.882 | 0.26861
[5 empty] | W/m^2 |  |  |  | W/m^2

## Measurements - Kumasi
Albedo Measurements
[7 empty] | Cloudless time periods |  | Average Albedo
Site: | Kumasi Technical University | [5 empty] | 12:30pm-14:30pm |  | 0.31022
Location: | Latitude: 6.678546°; Longitude: -1.05285° | [7 empty] | 0.31022 | Inclination
Date: | 14.02.2020; 12:30pm-2:30pm | [7 empty] | 0.31022 | 0%
|  | Top Pyranometer |  |  |  | Bottom Pyranometer
|  | Min. Value [V] | Max. Value [V] | Average Value [V] | Irradiance [W/m^2] | Min. Value [V] | Max. Value [V] | Average Value [V] | Irradiance [W/m^2] | Albedo | Average Albedo: | 0.31022
2020-02-14 | 12:30 | 6.749 | 6.756 | 6.7525 | 720.0138 | 1.874 | 1.875 | 1.8745 | 204.1209 | 0.2835
2020-02-14 | 12:30 | 6.738 | 6.748 | 6.743 | 719.0009 | 1.871 | 1.874 | 1.8725 | 203.9031 | 0.28359
2020-02-14 | 12:30 | 6.726 | 6.736 | 6.731 | 717.7213 | 1.867 | 1.871 | 1.869 | 203.522 | 0.28357
2020-02-14 | 12:30 | 6.717 | 6.722 | 6.7195 | 716.4951 | 1.865 | 1.867 | 1.866 | 203.1953 | 0.2836
2020-02-14 | 12:30 | 6.713 | 6.715 | 6.714 | 715.9086 | 1.864 | 1.865 | 1.8645 | 203.0319 | 0.2836
2020-02-14 | 12:30 | 6.714 | 6.718 | 6.716 | 716.1219 | 1.864 | 1.866 | 1.865 | 203.0864 | 0.28359
2020-02-14 | 12:30 | 6.717 | 6.72 | 6.7185 | 716.3884 | 1.865 | 1.866 | 1.8655 | 203.1408 | 0.28356
2020-02-14 | 12:31 | 6.696 | 6.715 | 6.7055 | 715.0023 | 1.858 | 1.864 | 1.861 | 202.6508 | 0.28343
2020-02-14 | 12:31 | 6.671 | 6.691 | 6.681 | 712.3898 | 1.851 | 1.858 | 1.8545 | 201.943 | 0.28347
2020-02-14 | 12:31 | 6.658 | 6.665 | 6.6615 | 710.3106 | 1.848 | 1.851 | 1.8495 | 201.3985 | 0.28354
2020-02-14 | 12:31 | 6.661 | 6.671 | 6.666 | 710.7904 | 1.847 | 1.851 | 1.849 | 201.3441 | 0.28327
2020-02-14 | 12:31 | 6.673 | 6.674 | 6.6735 | 711.5901 | 1.852 | 1.854 | 1.853 | 201.7797 | 0.28356
2020-02-14 | 12:31 | 6.645 | 6.669 | 6.657 | 709.8307 | 1.844 | 1.851 | 1.8475 | 201.1807 | 0.28342
2020-02-14 | 12:31 | 6.614 | 6.638 | 6.626 | 706.5252 | 1.836 | 1.844 | 1.84 | 200.364 | 0.28359
2020-02-14 | 12:31 | 6.58 | 6.605 | 6.5925 | 702.9532 | 1.826 | 1.834 | 1.83 | 199.2751 | 0.28348
2020-02-14 | 12:31 | 6.548 | 6.575 | 6.5615 | 699.6476 | 1.816 | 1.825 | 1.8205 | 198.2406 | 0.28334
2020-02-14 | 12:31 | 6.536 | 6.543 | 6.5395 | 697.3018 | 1.812 | 1.816 | 1.814 | 197.5328 | 0.28328
2020-02-14 | 12:31 | 6.528 | 6.534 | 6.531 | 696.3955 | 1.81 | 1.813 | 1.8115 | 197.2606 | 0.28326
2020-02-14 | 12:32 | 6.513 | 6.523 | 6.518 | 695.0093 | 1.806 | 1.81 | 1.808 | 196.8795 | 0.28328
2020-02-14 | 12:32 | 6.496 | 6.512 | 6.504 | 693.5165 | 1.802 | 1.806 | 1.804 | 196.4439 | 0.28326
2020-02-14 | 12:32 | 6.459 | 6.488 | 6.4735 | 690.2643 | 1.791 | 1.8 | 1.7955 | 195.5183 | 0.28325
2020-02-14 | 12:32 | 6.445 | 6.454 | 6.4495 | 687.7052 | 1.787 | 1.791 | 1.789 | 194.8105 | 0.28328
2020-02-14 | 12:32 | 6.427 | 6.441 | 6.434 | 686.0524 | 1.782 | 1.787 | 1.7845 | 194.3205 | 0.28324
2020-02-14 | 12:32 | 6.415 | 6.425 | 6.42 | 684.5596 | 1.778 | 1.781 | 1.7795 | 193.776 | 0.28307
2020-02-14 | 12:32 | 6.396 | 6.411 | 6.4035 | 682.8002 | 1.773 | 1.778 | 1.7755 | 193.3404 | 0.28316
2020-02-14 | 12:32 | 6.381 | 6.392 | 6.3865 | 680.9875 | 1.768 | 1.771 | 1.7695 | 192.6871 | 0.28295
2020-02-14 | 12:32 | 6.349 | 6.375 | 6.362 | 678.3751 | 1.759 | 1.767 | 1.763 | 191.9792 | 0.283
2020-02-14 | 12:32 | 6.326 | 6.344 | 6.335 | 675.4961 | 1.753 | 1.758 | 1.7555 | 191.1625 | 0.283
2020-02-14 | 12:32 | 6.308 | 6.32 | 6.314 | 673.2569 | 1.748 | 1.752 | 1.75 | 190.5636 | 0.28305
2020-02-14 | 12:32 | 6.312 | 6.314 | 6.313 | 673.1503 | 1.749 | 1.751 | 1.75 | 190.5636 | 0.28309
2020-02-14 | 12:32 | 6.311 | 6.314 | 6.3125 | 673.097 | 1.748 | 1.751 | 1.7495 | 190.5092 | 0.28303
2020-02-14 | 12:33 | 6.31 | 6.314 | 6.312 | 673.0437 | 1.748 | 1.751 | 1.7495 | 190.5092 | 0.28306
2020-02-14 | 12:33 | 6.309 | 6.316 | 6.3125 | 673.097 | 1.749 | 1.751 | 1.75 | 190.5636 | 0.28311
2020-02-14 | 12:33 | 6.312 | 6.322 | 6.317 | 673.5768 | 1.75 | 1.753 | 1.7515 | 190.727 | 0.28316
2020-02-14 | 12:33 | 6.268 | 6.303 | 6.2855 | 670.218 | 1.736 | 1.747 | 1.7415 | 189.638 | 0.28295
2020-02-14 | 12:33 | 6.238 | 6.26 | 6.249 | 666.326 | 1.728 | 1.736 | 1.732 | 188.6035 | 0.28305
2020-02-14 | 12:33 | 6.206 | 6.231 | 6.2185 | 663.0738 | 1.718 | 1.725 | 1.7215 | 187.4602 | 0.28271
2020-02-14 | 12:33 | 6.178 | 6.201 | 6.1895 | 659.9816 | 1.71 | 1.718 | 1.714 | 186.6435 | 0.2828
2020-02-14 | 12:33 | 6.124 | 6.168 | 6.146 | 655.3432 | 1.694 | 1.706 | 1.7 | 185.119 | 0.28248
2020-02-14 | 12:33 | 6.087 | 6.113 | 6.1 | 650.4383 | 1.682 | 1.692 | 1.687 | 183.7033 | 0.28243
2020-02-14 | 12:33 | 6.081 | 6.085 | 6.083 | 648.6256 | 1.664 | 1.678 | 1.671 | 181.961 | 0.28053
2020-02-14 | 12:33 | 6.087 | 6.09 | 6.0885 | 649.212 | 1.664 | 1.666 | 1.665 | 181.3077 | 0.27927
2020-02-14 | 12:34 | 6.084 | 6.087 | 6.0855 | 648.8921 | 1.663 | 1.678 | 1.6705 | 181.9066 | 0.28033
2020-02-14 | 12:34 | 6.09 | 6.098 | 6.094 | 649.7985 | 1.682 | 1.688 | 1.685 | 183.4856 | 0.28237
2020-02-14 | 12:34 | 6.067 | 6.093 | 6.08 | 648.3057 | 1.678 | 1.687 | 1.6825 | 183.2133 | 0.2826
2020-02-14 | 12:34 | 6.011 | 6.058 | 6.0345 | 643.454 | 1.664 | 1.677 | 1.6705 | 181.9066 | 0.2827
2020-02-14 | 12:34 | 5.92 | 5.993 | 5.9565 | 635.137 | 1.639 | 1.658 | 1.6485 | 179.5109 | 0.28263
2020-02-14 | 12:34 | 5.84 | 5.904 | 5.872 | 626.1268 | 1.614 | 1.635 | 1.6245 | 176.8975 | 0.28253
2020-02-14 | 12:34 | 5.733 | 5.824 | 5.7785 | 616.157 | 1.585 | 1.612 | 1.5985 | 174.0663 | 0.2825
2020-02-14 | 12:34 | 5.579 | 5.702 | 5.6405 | 601.4421 | 1.543 | 1.576 | 1.5595 | 169.8194 | 0.28235
2020-02-14 | 12:34 | 5.467 | 5.553 | 5.51 | 587.527 | 1.51 | 1.534 | 1.522 | 165.7359 | 0.28209
2020-02-14 | 12:34 | 5.407 | 5.447 | 5.427 | 578.6768 | 1.496 | 1.505 | 1.5005 | 163.3947 | 0.28236
2020-02-14 | 12:34 | 5.331 | 5.397 | 5.364 | 571.9592 | 1.474 | 1.493 | 1.4835 | 161.5435 | 0.28244
2020-02-14 | 12:34 | 5.3 | 5.315 | 5.3075 | 565.9346 | 1.465 | 1.469 | 1.467 | 159.7468 | 0.28227
2020-02-14 | 12:34 | 5.326 | 5.386 | 5.356 | 571.1061 | 1.473 | 1.489 | 1.481 | 161.2713 | 0.28238
2020-02-14 | 12:35 | 5.396 | 5.417 | 5.4065 | 576.4909 | 1.493 | 1.499 | 1.496 | 162.9047 | 0.28258
2020-02-14 | 12:35 | 5.415 | 5.464 | 5.4395 | 580.0097 | 1.498 | 1.512 | 1.505 | 163.8847 | 0.28256
2020-02-14 | 12:35 | 5.483 | 5.525 | 5.504 | 586.8872 | 1.517 | 1.529 | 1.523 | 165.8448 | 0.28258
2020-02-14 | 12:35 | 5.483 | 5.521 | 5.502 | 586.674 | 1.516 | 1.529 | 1.5225 | 165.7904 | 0.28259
2020-02-14 | 12:35 | 5.471 | 5.477 | 5.474 | 583.6884 | 1.513 | 1.515 | 1.514 | 164.8648 | 0.28245
2020-02-14 | 12:35 | 5.402 | 5.463 | 5.4325 | 579.2633 | 1.495 | 1.511 | 1.503 | 163.6669 | 0.28254
2020-02-14 | 12:35 | 5.31 | 5.391 | 5.3505 | 570.5197 | 1.469 | 1.492 | 1.4805 | 161.2168 | 0.28258
2020-02-14 | 12:35 | 5.227 | 5.279 | 5.253 | 560.1233 | 1.446 | 1.459 | 1.4525 | 158.1678 | 0.28238
2020-02-14 | 12:35 | 5.256 | 5.327 | 5.2915 | 564.2285 | 1.454 | 1.471 | 1.4625 | 159.2567 | 0.28226
2020-02-14 | 12:35 | 5.338 | 5.379 | 5.3585 | 571.3727 | 1.466 | 1.471 | 1.4685 | 159.9101 | 0.27987
2020-02-14 | 12:35 | 5.384 | 5.394 | 5.389 | 574.6249 | 1.467 | 1.474 | 1.4705 | 160.1279 | 0.27867
2020-02-14 | 12:36 | 5.367 | 5.396 | 5.3815 | 573.8252 | 1.464 | 1.473 | 1.4685 | 159.9101 | 0.27867
2020-02-14 | 12:36 | 5.348 | 5.36 | 5.354 | 570.8929 | 1.458 | 1.463 | 1.4605 | 159.039 | 0.27858
2020-02-14 | 12:36 | 5.331 | 5.346 | 5.3385 | 569.2401 | 1.454 | 1.458 | 1.456 | 158.5489 | 0.27853
2020-02-14 | 12:36 | 5.33 | 5.353 | 5.3415 | 569.56 | 1.459 | 1.464 | 1.4615 | 159.1479 | 0.27942
2020-02-14 | 12:36 | 5.357 | 5.39 | 5.3735 | 572.9721 | 1.466 | 1.472 | 1.469 | 159.9646 | 0.27918
2020-02-14 | 12:36 | 5.404 | 5.447 | 5.4255 | 578.5168 | 1.477 | 1.489 | 1.483 | 161.4891 | 0.27914
2020-02-14 | 12:36 | 5.456 | 5.475 | 5.4655 | 582.782 | 1.488 | 1.502 | 1.495 | 162.7958 | 0.27934
2020-02-14 | 12:36 | 5.468 | 5.474 | 5.471 | 583.3685 | 1.499 | 1.51 | 1.5045 | 163.8303 | 0.28083
2020-02-14 | 12:36 | 5.484 | 5.52 | 5.502 | 586.674 | 1.513 | 1.522 | 1.5175 | 165.2459 | 0.28167
2020-02-14 | 12:36 | 5.518 | 5.535 | 5.5265 | 589.2864 | 1.522 | 1.526 | 1.524 | 165.9537 | 0.28162
2020-02-14 | 12:36 | 5.542 | 5.563 | 5.5525 | 592.0588 | 1.531 | 1.541 | 1.536 | 167.2604 | 0.28251
2020-02-14 | 12:36 | 5.57 | 5.625 | 5.5975 | 596.8571 | 1.543 | 1.558 | 1.5505 | 168.8394 | 0.28288
2020-02-14 | 12:36 | 5.64 | 5.682 | 5.661 | 603.628 | 1.563 | 1.574 | 1.5685 | 170.7995 | 0.28295
2020-02-14 | 12:37 | 5.689 | 5.725 | 5.707 | 608.533 | 1.572 | 1.576 | 1.574 | 171.3984 | 0.28166
2020-02-14 | 12:37 | 5.738 | 5.764 | 5.751 | 613.2247 | 1.576 | 1.583 | 1.5795 | 171.9973 | 0.28048
2020-02-14 | 12:37 | 5.723 | 5.759 | 5.741 | 612.1584 | 1.57 | 1.58 | 1.575 | 171.5073 | 0.28017
2020-02-14 | 12:37 | 5.715 | 5.719 | 5.717 | 609.5993 | 1.567 | 1.569 | 1.568 | 170.745 | 0.28009
2020-02-14 | 12:37 | 5.705 | 5.714 | 5.7095 | 608.7995 | 1.565 | 1.567 | 1.566 | 170.5272 | 0.2801
2020-02-14 | 12:37 | 5.662 | 5.701 | 5.6815 | 605.8139 | 1.553 | 1.564 | 1.5585 | 169.7105 | 0.28014
2020-02-14 | 12:37 | 5.597 | 5.649 | 5.623 | 599.5761 | 1.528 | 1.548 | 1.538 | 167.4782 | 0.27933
2020-02-14 | 12:37 | 5.542 | 5.586 | 5.564 | 593.285 | 1.514 | 1.525 | 1.5195 | 165.4637 | 0.27889
2020-02-14 | 12:37 | 5.523 | 5.535 | 5.529 | 589.553 | 1.507 | 1.512 | 1.5095 | 164.3747 | 0.27881
2020-02-14 | 12:37 | 5.53 | 5.57 | 5.55 | 591.7922 | 1.508 | 1.517 | 1.5125 | 164.7014 | 0.27831
2020-02-14 | 12:37 | 5.578 | 5.589 | 5.5835 | 595.3643 | 1.507 | 1.517 | 1.512 | 164.647 | 0.27655
2020-02-14 | 12:38 | 5.557 | 5.585 | 5.571 | 594.0314 | 1.496 | 1.505 | 1.5005 | 163.3947 | 0.27506
2020-02-14 | 12:38 | 5.552 | 5.571 | 5.5615 | 593.0184 | 1.5 | 1.514 | 1.507 | 164.1025 | 0.27672
2020-02-14 | 12:38 | 5.568 | 5.58 | 5.574 | 594.3513 | 1.516 | 1.518 | 1.517 | 165.1914 | 0.27794
2020-02-14 | 12:38 | 5.506 | 5.557 | 5.5315 | 589.8195 | 1.498 | 1.514 | 1.506 | 163.9936 | 0.27804
2020-02-14 | 12:38 | 5.487 | 5.497 | 5.492 | 585.6077 | 1.494 | 1.496 | 1.495 | 162.7958 | 0.27799
2020-02-14 | 12:38 | 5.499 | 5.542 | 5.5205 | 588.6466 | 1.496 | 1.51 | 1.503 | 163.6669 | 0.27804
2020-02-14 | 12:38 | 5.537 | 5.549 | 5.543 | 591.0458 | 1.508 | 1.511 | 1.5095 | 164.3747 | 0.27811
2020-02-14 | 12:38 | 5.532 | 5.534 | 5.533 | 589.9795 | 1.507 | 1.508 | 1.5075 | 164.157 | 0.27824
2020-02-14 | 12:38 | 5.513 | 5.528 | 5.5205 | 588.6466 | 1.506 | 1.507 | 1.5065 | 164.0481 | 0.27869
2020-02-14 | 12:38 | 5.503 | 5.51 | 5.5065 | 587.1538 | 1.507 | 1.51 | 1.5085 | 164.2658 | 0.27977
2020-02-14 | 12:38 | 5.509 | 5.529 | 5.519 | 588.4867 | 1.512 | 1.517 | 1.5145 | 164.9192 | 0.28024
2020-02-14 | 12:38 | 5.533 | 5.542 | 5.5375 | 590.4593 | 1.517 | 1.52 | 1.5185 | 165.3548 | 0.28004
2020-02-14 | 12:39 | 5.478 | 5.53 | 5.504 | 586.8872 | 1.499 | 1.514 | 1.5065 | 164.0481 | 0.27952
2020-02-14 | 12:39 | 5.442 | 5.469 | 5.4555 | 581.7157 | 1.488 | 1.497 | 1.4925 | 162.5236 | 0.27939
2020-02-14 | 12:39 | 5.425 | 5.436 | 5.4305 | 579.05 | 1.484 | 1.486 | 1.485 | 161.7069 | 0.27926
2020-02-14 | 12:39 | 5.411 | 5.426 | 5.4185 | 577.7704 | 1.479 | 1.484 | 1.4815 | 161.3257 | 0.27922
2020-02-14 | 12:39 | 5.357 | 5.398 | 5.3775 | 573.3986 | 1.466 | 1.476 | 1.471 | 160.1823 | 0.27936
2020-02-14 | 12:39 | 5.317 | 5.352 | 5.3345 | 568.8136 | 1.454 | 1.461 | 1.4575 | 158.7123 | 0.27902
2020-02-14 | 12:39 | 5.266 | 5.306 | 5.286 | 563.6421 | 1.44 | 1.449 | 1.4445 | 157.2967 | 0.27907
2020-02-14 | 12:39 | 5.201 | 5.258 | 5.2295 | 557.6175 | 1.419 | 1.437 | 1.428 | 155.4999 | 0.27886
2020-02-14 | 12:39 | 5.177 | 5.189 | 5.183 | 552.6593 | 1.413 | 1.417 | 1.415 | 154.0843 | 0.27881
2020-02-14 | 12:39 | 5.192 | 5.21 | 5.201 | 554.5786 | 1.418 | 1.423 | 1.4205 | 154.6832 | 0.27892
2020-02-14 | 12:39 | 5.189 | 5.202 | 5.1955 | 553.9921 | 1.421 | 1.424 | 1.4225 | 154.901 | 0.27961
2020-02-14 | 12:39 | 5.14 | 5.187 | 5.1635 | 550.58 | 1.413 | 1.426 | 1.4195 | 154.5743 | 0.28075
2020-02-14 | 12:39 | 5.095 | 5.122 | 5.1085 | 544.7154 | 1.403 | 1.41 | 1.4065 | 153.1587 | 0.28117
2020-02-14 | 12:40 | 5.075 | 5.091 | 5.083 | 541.9963 | 1.398 | 1.402 | 1.4 | 152.4509 | 0.28128
2020-02-14 | 12:40 | 5.071 | 5.075 | 5.073 | 540.93 | 1.396 | 1.397 | 1.3965 | 152.0698 | 0.28113
2020-02-14 | 12:40 | 5.031 | 5.07 | 5.0505 | 538.5309 | 1.385 | 1.395 | 1.39 | 151.362 | 0.28106
2020-02-14 | 12:40 | 5.028 | 5.037 | 5.0325 | 536.6116 | 1.383 | 1.386 | 1.3845 | 150.7631 | 0.28095
2020-02-14 | 12:40 | 4.937 | 5.019 | 4.978 | 530.8003 | 1.359 | 1.381 | 1.37 | 149.1841 | 0.28106
2020-02-14 | 12:40 | 4.93 | 4.946 | 4.938 | 526.5351 | 1.357 | 1.361 | 1.359 | 147.9863 | 0.28106
2020-02-14 | 12:40 | 4.919 | 4.946 | 4.9325 | 525.9486 | 1.352 | 1.361 | 1.3565 | 147.714 | 0.28085
2020-02-14 | 12:40 | 4.789 | 4.9 | 4.8445 | 516.5653 | 1.317 | 1.349 | 1.333 | 145.155 | 0.281
2020-02-14 | 12:40 | 4.725 | 4.761 | 4.743 | 505.7424 | 1.299 | 1.309 | 1.304 | 141.9971 | 0.28077
2020-02-14 | 12:40 | 4.738 | 4.785 | 4.7615 | 507.715 | 1.302 | 1.317 | 1.3095 | 142.596 | 0.28086
2020-02-14 | 12:40 | 4.788 | 4.792 | 4.79 | 510.754 | 1.317 | 1.319 | 1.318 | 143.5216 | 0.281
2020-02-14 | 12:41 | 4.745 | 4.784 | 4.7645 | 508.0349 | 1.306 | 1.316 | 1.311 | 142.7594 | 0.281
2020-02-14 | 12:41 | 4.715 | 4.734 | 4.7245 | 503.7698 | 1.299 | 1.305 | 1.302 | 141.7793 | 0.28144
2020-02-14 | 12:41 | 4.747 | 4.825 | 4.786 | 510.3275 | 1.308 | 1.33 | 1.319 | 143.6305 | 0.28145
2020-02-14 | 12:41 | 4.858 | 5.029 | 4.9435 | 527.1216 | 1.34 | 1.388 | 1.364 | 148.5307 | 0.28178
2020-02-14 | 12:41 | 5.04 | 5.063 | 5.0515 | 538.6375 | 1.392 | 1.397 | 1.3945 | 151.852 | 0.28192
2020-02-14 | 12:41 | 4.985 | 5.025 | 5.005 | 533.6793 | 1.378 | 1.388 | 1.383 | 150.5997 | 0.28219
2020-02-14 | 12:41 | 4.919 | 4.976 | 4.9475 | 527.5481 | 1.359 | 1.375 | 1.367 | 148.8574 | 0.28217
2020-02-14 | 12:41 | 4.904 | 4.944 | 4.924 | 525.0423 | 1.356 | 1.367 | 1.3615 | 148.2585 | 0.28237
2020-02-14 | 12:41 | 4.966 | 5.042 | 5.004 | 533.5726 | 1.372 | 1.393 | 1.3825 | 150.5453 | 0.28215
2020-02-14 | 12:41 | 5.058 | 5.108 | 5.083 | 541.9963 | 1.398 | 1.412 | 1.405 | 152.9954 | 0.28228
2020-02-14 | 12:41 | 5.114 | 5.123 | 5.1185 | 545.7817 | 1.414 | 1.417 | 1.4155 | 154.1388 | 0.28242
2020-02-14 | 12:41 | 5.122 | 5.152 | 5.137 | 547.7543 | 1.416 | 1.425 | 1.4205 | 154.6832 | 0.2824
2020-02-14 | 12:41 | 5.183 | 5.375 | 5.279 | 562.8957 | 1.434 | 1.488 | 1.461 | 159.0934 | 0.28263
2020-02-14 | 12:42 | 5.424 | 5.574 | 5.499 | 586.3541 | 1.502 | 1.542 | 1.522 | 165.7359 | 0.28265
2020-02-14 | 12:42 | 5.597 | 5.654 | 5.6255 | 599.8427 | 1.549 | 1.564 | 1.5565 | 169.4927 | 0.28256
2020-02-14 | 12:42 | 5.656 | 5.662 | 5.659 | 603.4148 | 1.565 | 1.567 | 1.566 | 170.5272 | 0.2826
2020-02-14 | 12:42 | 5.655 | 5.67 | 5.6625 | 603.788 | 1.564 | 1.568 | 1.566 | 170.5272 | 0.28243
2020-02-14 | 12:42 | 5.675 | 5.687 | 5.681 | 605.7606 | 1.569 | 1.573 | 1.571 | 171.0717 | 0.28241
2020-02-14 | 12:42 | 5.689 | 5.702 | 5.6955 | 607.3067 | 1.573 | 1.577 | 1.575 | 171.5073 | 0.28241
2020-02-14 | 12:42 | 5.707 | 5.739 | 5.723 | 610.239 | 1.577 | 1.586 | 1.5815 | 172.2151 | 0.28221
2020-02-14 | 12:42 | 5.744 | 5.764 | 5.754 | 613.5445 | 1.588 | 1.593 | 1.5905 | 173.1951 | 0.28229
2020-02-14 | 12:42 | 5.769 | 5.799 | 5.784 | 616.7434 | 1.595 | 1.605 | 1.6 | 174.2296 | 0.2825
2020-02-14 | 12:42 | 5.811 | 5.856 | 5.8335 | 622.0216 | 1.607 | 1.62 | 1.6135 | 175.6997 | 0.28247
2020-02-14 | 12:42 | 5.869 | 5.913 | 5.891 | 628.1528 | 1.623 | 1.635 | 1.629 | 177.3875 | 0.2824
2020-02-14 | 12:43 | 5.925 | 5.972 | 5.9485 | 634.2839 | 1.64 | 1.652 | 1.646 | 179.2387 | 0.28258
2020-02-14 | 12:43 | 5.977 | 6.004 | 5.9905 | 638.7624 | 1.652 | 1.66 | 1.656 | 180.3276 | 0.28231
2020-02-14 | 12:43 | 6.01 | 6.017 | 6.0135 | 641.2148 | 1.663 | 1.665 | 1.664 | 181.1988 | 0.28259
2020-02-14 | 12:43 | 6.018 | 6.068 | 6.043 | 644.3604 | 1.664 | 1.679 | 1.6715 | 182.0155 | 0.28247
2020-02-14 | 12:43 | 6.087 | 6.133 | 6.11 | 651.5046 | 1.684 | 1.696 | 1.69 | 184.03 | 0.28247
2020-02-14 | 12:43 | 6.141 | 6.16 | 6.1505 | 655.823 | 1.697 | 1.704 | 1.7005 | 185.1734 | 0.28235
2020-02-14 | 12:43 | 6.165 | 6.181 | 6.173 | 658.2222 | 1.705 | 1.71 | 1.7075 | 185.9357 | 0.28248
2020-02-14 | 12:43 | 6.185 | 6.203 | 6.194 | 660.4614 | 1.711 | 1.715 | 1.713 | 186.5346 | 0.28243
2020-02-14 | 12:43 | 6.206 | 6.216 | 6.211 | 662.2741 | 1.717 | 1.719 | 1.718 | 187.079 | 0.28248
2020-02-14 | 12:43 | 6.219 | 6.227 | 6.223 | 663.5537 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.28243
2020-02-14 | 12:43 | 6.225 | 6.227 | 6.226 | 663.8735 | 1.72 | 1.723 | 1.7215 | 187.4602 | 0.28237
2020-02-14 | 12:43 | 6.222 | 6.224 | 6.223 | 663.5537 | 1.719 | 1.721 | 1.72 | 187.2968 | 0.28226
2020-02-14 | 12:43 | 6.223 | 6.224 | 6.2235 | 663.607 | 1.72 | 1.72 | 1.72 | 187.2968 | 0.28224
2020-02-14 | 12:44 | 6.223 | 6.226 | 6.2245 | 663.7136 | 1.718 | 1.722 | 1.72 | 187.2968 | 0.2822
2020-02-14 | 12:44 | 6.228 | 6.233 | 6.2305 | 664.3534 | 1.721 | 1.722 | 1.7215 | 187.4602 | 0.28217
2020-02-14 | 12:44 | 6.237 | 6.244 | 6.2405 | 665.4197 | 1.723 | 1.725 | 1.724 | 187.7324 | 0.28213
2020-02-14 | 12:44 | 6.246 | 6.257 | 6.2515 | 666.5926 | 1.726 | 1.729 | 1.7275 | 188.1135 | 0.2822
2020-02-14 | 12:44 | 6.257 | 6.26 | 6.2585 | 667.339 | 1.729 | 1.73 | 1.7295 | 188.3313 | 0.28221
2020-02-14 | 12:44 | 6.256 | 6.26 | 6.258 | 667.2857 | 1.728 | 1.731 | 1.7295 | 188.3313 | 0.28223
2020-02-14 | 12:44 | 6.25 | 6.255 | 6.2525 | 666.6992 | 1.727 | 1.729 | 1.728 | 188.168 | 0.28224
2020-02-14 | 12:44 | 6.236 | 6.248 | 6.242 | 665.5796 | 1.724 | 1.727 | 1.7255 | 187.8957 | 0.2823
2020-02-14 | 12:44 | 6.226 | 6.233 | 6.2295 | 664.2467 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.2823
2020-02-14 | 12:44 | 6.216 | 6.223 | 6.2195 | 663.1805 | 1.717 | 1.721 | 1.719 | 187.1879 | 0.28226
2020-02-14 | 12:44 | 6.205 | 6.213 | 6.209 | 662.0608 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.28232
2020-02-14 | 12:45 | 6.202 | 6.204 | 6.203 | 661.4211 | 1.714 | 1.716 | 1.715 | 186.7524 | 0.28235
2020-02-14 | 12:45 | 6.198 | 6.201 | 6.1995 | 661.0479 | 1.714 | 1.715 | 1.7145 | 186.6979 | 0.28243
2020-02-14 | 12:45 | 6.192 | 6.198 | 6.195 | 660.568 | 1.712 | 1.714 | 1.713 | 186.5346 | 0.28239
2020-02-14 | 12:45 | 6.192 | 6.193 | 6.1925 | 660.3015 | 1.712 | 1.714 | 1.713 | 186.5346 | 0.2825
2020-02-14 | 12:45 | 6.192 | 6.194 | 6.193 | 660.3548 | 1.711 | 1.713 | 1.712 | 186.4257 | 0.28231
2020-02-14 | 12:45 | 6.195 | 6.196 | 6.1955 | 660.6214 | 1.712 | 1.713 | 1.7125 | 186.4801 | 0.28228
2020-02-14 | 12:45 | 6.192 | 6.194 | 6.193 | 660.3548 | 1.711 | 1.713 | 1.712 | 186.4257 | 0.28231
2020-02-14 | 12:45 | 6.19 | 6.193 | 6.1915 | 660.1948 | 1.711 | 1.713 | 1.712 | 186.4257 | 0.28238
2020-02-14 | 12:45 | 6.192 | 6.195 | 6.1935 | 660.4081 | 1.711 | 1.713 | 1.712 | 186.4257 | 0.28229
2020-02-14 | 12:45 | 6.193 | 6.195 | 6.194 | 660.4614 | 1.712 | 1.713 | 1.7125 | 186.4801 | 0.28235
2020-02-14 | 12:45 | 6.183 | 6.19 | 6.1865 | 659.6617 | 1.71 | 1.711 | 1.7105 | 186.2623 | 0.28236
2020-02-14 | 12:45 | 6.171 | 6.18 | 6.1755 | 658.4888 | 1.706 | 1.709 | 1.7075 | 185.9357 | 0.28237
2020-02-14 | 12:45 | 6.158 | 6.169 | 6.1635 | 657.2092 | 1.702 | 1.706 | 1.704 | 185.5545 | 0.28234
2020-02-14 | 12:46 | 6.147 | 6.155 | 6.151 | 655.8764 | 1.699 | 1.702 | 1.7005 | 185.1734 | 0.28233
2020-02-14 | 12:46 | 6.136 | 6.144 | 6.14 | 654.7034 | 1.698 | 1.699 | 1.6985 | 184.9556 | 0.2825
2020-02-14 | 12:46 | 6.128 | 6.135 | 6.1315 | 653.7971 | 1.694 | 1.697 | 1.6955 | 184.6289 | 0.28239
2020-02-14 | 12:46 | 6.125 | 6.127 | 6.126 | 653.2106 | 1.694 | 1.695 | 1.6945 | 184.52 | 0.28248
2020-02-14 | 12:46 | 6.127 | 6.135 | 6.131 | 653.7438 | 1.694 | 1.697 | 1.6955 | 184.6289 | 0.28242
2020-02-14 | 12:46 | 6.137 | 6.147 | 6.142 | 654.9167 | 1.699 | 1.701 | 1.7 | 185.119 | 0.28266
2020-02-14 | 12:46 | 6.142 | 6.149 | 6.1455 | 655.2899 | 1.698 | 1.702 | 1.7 | 185.119 | 0.2825
2020-02-14 | 12:46 | 6.112 | 6.136 | 6.124 | 652.9974 | 1.691 | 1.698 | 1.6945 | 184.52 | 0.28257
2020-02-14 | 12:46 | 6.106 | 6.11 | 6.108 | 651.2913 | 1.689 | 1.691 | 1.69 | 184.03 | 0.28256
2020-02-14 | 12:46 | 6.113 | 6.136 | 6.1245 | 653.0507 | 1.692 | 1.7 | 1.696 | 184.6834 | 0.2828
2020-02-14 | 12:46 | 6.141 | 6.156 | 6.1485 | 655.6098 | 1.701 | 1.705 | 1.703 | 185.4456 | 0.28286
2020-02-14 | 12:47 | 6.158 | 6.165 | 6.1615 | 656.996 | 1.706 | 1.707 | 1.7065 | 185.8268 | 0.28284
2020-02-14 | 12:47 | 6.163 | 6.169 | 6.166 | 657.4758 | 1.707 | 1.708 | 1.7075 | 185.9357 | 0.2828
2020-02-14 | 12:47 | 6.171 | 6.185 | 6.178 | 658.7553 | 1.709 | 1.714 | 1.7115 | 186.3712 | 0.28291
2020-02-14 | 12:47 | 6.19 | 6.2 | 6.195 | 660.568 | 1.715 | 1.718 | 1.7165 | 186.9157 | 0.28296
2020-02-14 | 12:47 | 6.199 | 6.203 | 6.201 | 661.2078 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.28285
2020-02-14 | 12:47 | 6.197 | 6.199 | 6.198 | 660.8879 | 1.716 | 1.717 | 1.7165 | 186.9157 | 0.28283
2020-02-14 | 12:47 | 6.2 | 6.202 | 6.201 | 661.2078 | 1.717 | 1.718 | 1.7175 | 187.0246 | 0.28285
2020-02-14 | 12:47 | 6.203 | 6.206 | 6.2045 | 661.581 | 1.718 | 1.719 | 1.7185 | 187.1335 | 0.28286
2020-02-14 | 12:47 | 6.206 | 6.212 | 6.209 | 662.0608 | 1.719 | 1.721 | 1.72 | 187.2968 | 0.2829
2020-02-14 | 12:47 | 6.215 | 6.219 | 6.217 | 662.9139 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.28286
2020-02-14 | 12:47 | 6.209 | 6.216 | 6.2125 | 662.434 | 1.72 | 1.722 | 1.721 | 187.4057 | 0.2829
2020-02-14 | 12:47 | 6.202 | 6.208 | 6.205 | 661.6343 | 1.717 | 1.72 | 1.7185 | 187.1335 | 0.28284
2020-02-14 | 12:47 | 6.2 | 6.208 | 6.204 | 661.5277 | 1.717 | 1.719 | 1.718 | 187.079 | 0.2828
2020-02-14 | 12:48 | 6.212 | 6.223 | 6.2175 | 662.9672 | 1.721 | 1.723 | 1.722 | 187.5146 | 0.28284
2020-02-14 | 12:48 | 6.224 | 6.231 | 6.2275 | 664.0335 | 1.725 | 1.727 | 1.726 | 187.9502 | 0.28304
2020-02-14 | 12:48 | 6.235 | 6.245 | 6.24 | 665.3664 | 1.727 | 1.73 | 1.7285 | 188.2224 | 0.28289
2020-02-14 | 12:48 | 6.242 | 6.247 | 6.2445 | 665.8462 | 1.728 | 1.73 | 1.729 | 188.2769 | 0.28276
2020-02-14 | 12:48 | 6.234 | 6.242 | 6.238 | 665.1531 | 1.728 | 1.728 | 1.728 | 188.168 | 0.28289
2020-02-14 | 12:48 | 6.229 | 6.234 | 6.2315 | 664.46 | 1.725 | 1.727 | 1.726 | 187.9502 | 0.28286
2020-02-14 | 12:48 | 6.22 | 6.228 | 6.224 | 663.6603 | 1.722 | 1.725 | 1.7235 | 187.678 | 0.28279
2020-02-14 | 12:48 | 6.192 | 6.214 | 6.203 | 661.4211 | 1.715 | 1.721 | 1.718 | 187.079 | 0.28284
2020-02-14 | 12:48 | 6.169 | 6.187 | 6.178 | 658.7553 | 1.709 | 1.716 | 1.7125 | 186.4801 | 0.28308
2020-02-14 | 12:48 | 6.15 | 6.165 | 6.1575 | 656.5694 | 1.703 | 1.708 | 1.7055 | 185.7179 | 0.28286
2020-02-14 | 12:48 | 6.132 | 6.147 | 6.1395 | 654.6501 | 1.699 | 1.703 | 1.701 | 185.2278 | 0.28294
2020-02-14 | 12:49 | 6.093 | 6.126 | 6.1095 | 651.4512 | 1.688 | 1.698 | 1.693 | 184.3567 | 0.28299
2020-02-14 | 12:49 | 6.083 | 6.086 | 6.0845 | 648.7855 | 1.686 | 1.688 | 1.687 | 183.7033 | 0.28315
2020-02-14 | 12:49 | 6.058 | 6.087 | 6.0725 | 647.506 | 1.678 | 1.687 | 1.6825 | 183.2133 | 0.28295
2020-02-14 | 12:49 | 6.004 | 6.047 | 6.0255 | 642.4944 | 1.664 | 1.676 | 1.67 | 181.8521 | 0.28304
2020-02-14 | 12:49 | 5.957 | 5.994 | 5.9755 | 637.1629 | 1.649 | 1.66 | 1.6545 | 180.1643 | 0.28276
2020-02-14 | 12:49 | 5.943 | 5.952 | 5.9475 | 634.1773 | 1.646 | 1.649 | 1.6475 | 179.402 | 0.28289
2020-02-14 | 12:49 | 5.94 | 5.944 | 5.942 | 633.5908 | 1.646 | 1.647 | 1.6465 | 179.2932 | 0.28298
2020-02-14 | 12:49 | 5.951 | 5.99 | 5.9705 | 636.6298 | 1.649 | 1.66 | 1.6545 | 180.1643 | 0.283
2020-02-14 | 12:49 | 5.978 | 5.992 | 5.985 | 638.1759 | 1.657 | 1.661 | 1.659 | 180.6543 | 0.28308
2020-02-14 | 12:49 | 5.98 | 5.988 | 5.984 | 638.0693 | 1.658 | 1.661 | 1.6595 | 180.7088 | 0.28321
2020-02-14 | 12:49 | 5.976 | 5.985 | 5.9805 | 637.6961 | 1.657 | 1.66 | 1.6585 | 180.5999 | 0.28321
2020-02-14 | 12:49 | 5.978 | 5.996 | 5.987 | 638.3892 | 1.658 | 1.663 | 1.6605 | 180.8177 | 0.28324
2020-02-14 | 12:49 | 5.988 | 5.996 | 5.992 | 638.9223 | 1.659 | 1.663 | 1.661 | 180.8721 | 0.28309
2020-02-14 | 12:50 | 5.951 | 5.98 | 5.9655 | 636.0966 | 1.649 | 1.658 | 1.6535 | 180.0554 | 0.28306
2020-02-14 | 12:50 | 5.944 | 5.946 | 5.945 | 633.9107 | 1.647 | 1.649 | 1.648 | 179.4565 | 0.28309
2020-02-14 | 12:50 | 5.887 | 5.942 | 5.9145 | 630.6585 | 1.631 | 1.647 | 1.639 | 178.4764 | 0.283
2020-02-14 | 12:50 | 5.742 | 5.863 | 5.8025 | 618.7161 | 1.589 | 1.624 | 1.6065 | 174.9374 | 0.28274
2020-02-14 | 12:50 | 5.658 | 5.713 | 5.6855 | 606.2404 | 1.565 | 1.581 | 1.573 | 171.2895 | 0.28254
2020-02-14 | 12:50 | 5.629 | 5.657 | 5.643 | 601.7087 | 1.557 | 1.565 | 1.561 | 169.9828 | 0.2825
2020-02-14 | 12:50 | 5.611 | 5.62 | 5.6155 | 598.7764 | 1.552 | 1.554 | 1.553 | 169.1116 | 0.28243
2020-02-14 | 12:50 | 5.613 | 5.618 | 5.6155 | 598.7764 | 1.553 | 1.554 | 1.5535 | 169.1661 | 0.28252
2020-02-14 | 12:50 | 5.625 | 5.644 | 5.6345 | 600.8024 | 1.556 | 1.56 | 1.558 | 169.6561 | 0.28238
2020-02-14 | 12:50 | 5.55 | 5.632 | 5.591 | 596.164 | 1.535 | 1.558 | 1.5465 | 168.4038 | 0.28248
2020-02-14 | 12:50 | 5.459 | 5.523 | 5.491 | 585.5011 | 1.508 | 1.526 | 1.517 | 165.1914 | 0.28214
2020-02-14 | 12:51 | 5.41 | 5.454 | 5.432 | 579.2099 | 1.494 | 1.507 | 1.5005 | 163.3947 | 0.2821
2020-02-14 | 12:51 | 5.242 | 5.381 | 5.3115 | 566.3611 | 1.446 | 1.486 | 1.466 | 159.6379 | 0.28187
2020-02-14 | 12:51 | 5.207 | 5.246 | 5.2265 | 557.2976 | 1.438 | 1.451 | 1.4445 | 157.2967 | 0.28225
2020-02-14 | 12:51 | 5.278 | 5.398 | 5.338 | 569.1868 | 1.46 | 1.492 | 1.476 | 160.7268 | 0.28238
2020-02-14 | 12:51 | 5.319 | 5.4 | 5.3595 | 571.4793 | 1.47 | 1.493 | 1.4815 | 161.3257 | 0.28229
2020-02-14 | 12:51 | 5.31 | 5.387 | 5.3485 | 570.3064 | 1.468 | 1.491 | 1.4795 | 161.1079 | 0.28249
2020-02-14 | 12:51 | 5.413 | 5.493 | 5.453 | 581.4492 | 1.497 | 1.521 | 1.509 | 164.3203 | 0.2826
2020-02-14 | 12:51 | 5.509 | 5.575 | 5.542 | 590.9392 | 1.525 | 1.543 | 1.534 | 167.0426 | 0.28267
2020-02-14 | 12:51 | 5.588 | 5.617 | 5.6025 | 597.3902 | 1.547 | 1.556 | 1.5515 | 168.9483 | 0.28281
2020-02-14 | 12:51 | 5.631 | 5.708 | 5.6695 | 604.5344 | 1.56 | 1.582 | 1.571 | 171.0717 | 0.28298
2020-02-14 | 12:51 | 5.726 | 5.797 | 5.7615 | 614.3443 | 1.585 | 1.607 | 1.596 | 173.794 | 0.28289
2020-02-14 | 12:51 | 5.812 | 5.86 | 5.836 | 622.2881 | 1.61 | 1.623 | 1.6165 | 176.0263 | 0.28287
2020-02-14 | 12:51 | 5.874 | 5.936 | 5.905 | 629.6456 | 1.628 | 1.646 | 1.637 | 178.2587 | 0.28311
2020-02-14 | 12:52 | 5.945 | 5.97 | 5.9575 | 635.2436 | 1.647 | 1.655 | 1.651 | 179.7832 | 0.28301
2020-02-14 | 12:52 | 5.977 | 5.998 | 5.9875 | 638.4425 | 1.656 | 1.662 | 1.659 | 180.6543 | 0.28296
2020-02-14 | 12:52 | 6.001 | 6.012 | 6.0065 | 640.4684 | 1.662 | 1.666 | 1.664 | 181.1988 | 0.28292
2020-02-14 | 12:52 | 6.013 | 6.019 | 6.016 | 641.4814 | 1.666 | 1.668 | 1.667 | 181.5255 | 0.28298
2020-02-14 | 12:52 | 5.994 | 6.016 | 6.005 | 640.3085 | 1.66 | 1.667 | 1.6635 | 181.1443 | 0.2829
2020-02-14 | 12:52 | 5.991 | 6.009 | 6 | 639.7753 | 1.66 | 1.665 | 1.6625 | 181.0354 | 0.28297
2020-02-14 | 12:52 | 6.015 | 6.021 | 6.018 | 641.6947 | 1.666 | 1.668 | 1.667 | 181.5255 | 0.28288
2020-02-14 | 12:52 | 6.019 | 6.022 | 6.0205 | 641.9612 | 1.667 | 1.668 | 1.6675 | 181.5799 | 0.28285
2020-02-14 | 12:52 | 6.012 | 6.023 | 6.0175 | 641.6413 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.28299
2020-02-14 | 12:52 | 6.008 | 6.009 | 6.0085 | 640.6817 | 1.664 | 1.665 | 1.6645 | 181.2532 | 0.28291
2020-02-14 | 12:52 | 6.01 | 6.026 | 6.018 | 641.6947 | 1.664 | 1.669 | 1.6665 | 181.471 | 0.2828
2020-02-14 | 12:52 | 6.032 | 6.04 | 6.036 | 643.614 | 1.671 | 1.673 | 1.672 | 182.0699 | 0.28289
2020-02-14 | 12:53 | 6.037 | 6.04 | 6.0385 | 643.8806 | 1.671 | 1.672 | 1.6715 | 182.0155 | 0.28269
2020-02-14 | 12:53 | 6.044 | 6.062 | 6.053 | 645.4267 | 1.674 | 1.678 | 1.676 | 182.5055 | 0.28277
2020-02-14 | 12:53 | 6.067 | 6.085 | 6.076 | 647.8792 | 1.68 | 1.686 | 1.683 | 183.2678 | 0.28287
2020-02-14 | 12:53 | 6.088 | 6.094 | 6.091 | 649.4786 | 1.686 | 1.688 | 1.687 | 183.7033 | 0.28285
2020-02-14 | 12:53 | 6.084 | 6.096 | 6.09 | 649.372 | 1.683 | 1.688 | 1.6855 | 183.54 | 0.28264
2020-02-14 | 12:53 | 6.066 | 6.077 | 6.0715 | 647.3993 | 1.679 | 1.683 | 1.681 | 183.05 | 0.28275
2020-02-14 | 12:53 | 6.054 | 6.066 | 6.06 | 646.1731 | 1.676 | 1.679 | 1.6775 | 182.6688 | 0.28269
2020-02-14 | 12:53 | 6.002 | 6.048 | 6.025 | 642.4411 | 1.662 | 1.674 | 1.668 | 181.6344 | 0.28273
2020-02-14 | 12:53 | 5.989 | 6.009 | 5.999 | 639.6687 | 1.658 | 1.664 | 1.661 | 180.8721 | 0.28276
2020-02-14 | 12:53 | 6.027 | 6.114 | 6.0705 | 647.2927 | 1.669 | 1.696 | 1.6825 | 183.2133 | 0.28305
2020-02-14 | 12:53 | 6.13 | 6.181 | 6.1555 | 656.3562 | 1.698 | 1.714 | 1.706 | 185.7723 | 0.28304
2020-02-14 | 12:54 | 6.192 | 6.225 | 6.2085 | 662.0075 | 1.716 | 1.726 | 1.721 | 187.4057 | 0.28309
2020-02-14 | 12:54 | 6.232 | 6.263 | 6.2475 | 666.1661 | 1.728 | 1.737 | 1.7325 | 188.658 | 0.2832
2020-02-14 | 12:54 | 6.268 | 6.281 | 6.2745 | 669.0451 | 1.739 | 1.741 | 1.74 | 189.4747 | 0.2832
2020-02-14 | 12:54 | 6.285 | 6.295 | 6.29 | 670.6978 | 1.742 | 1.746 | 1.744 | 189.9103 | 0.28315
2020-02-14 | 12:54 | 6.301 | 6.318 | 6.3095 | 672.7771 | 1.748 | 1.752 | 1.75 | 190.5636 | 0.28325
2020-02-14 | 12:54 | 6.312 | 6.319 | 6.3155 | 673.4169 | 1.751 | 1.752 | 1.7515 | 190.727 | 0.28322
2020-02-14 | 12:54 | 6.305 | 6.313 | 6.309 | 672.7238 | 1.748 | 1.75 | 1.749 | 190.4547 | 0.28311
2020-02-14 | 12:54 | 6.308 | 6.321 | 6.3145 | 673.3102 | 1.749 | 1.753 | 1.751 | 190.6725 | 0.28319
2020-02-14 | 12:54 | 6.325 | 6.338 | 6.3315 | 675.1229 | 1.753 | 1.758 | 1.7555 | 191.1625 | 0.28315
2020-02-14 | 12:54 | 6.342 | 6.373 | 6.3575 | 677.8953 | 1.759 | 1.769 | 1.764 | 192.0881 | 0.28336
2020-02-14 | 12:54 | 6.385 | 6.431 | 6.408 | 683.2801 | 1.771 | 1.784 | 1.7775 | 193.5582 | 0.28328
2020-02-14 | 12:54 | 6.437 | 6.466 | 6.4515 | 687.9184 | 1.786 | 1.794 | 1.79 | 194.9194 | 0.28335
2020-02-14 | 12:54 | 6.474 | 6.496 | 6.485 | 691.4905 | 1.796 | 1.802 | 1.799 | 195.8994 | 0.2833
2020-02-14 | 12:55 | 6.497 | 6.507 | 6.502 | 693.3032 | 1.801 | 1.805 | 1.803 | 196.335 | 0.28319
2020-02-14 | 12:55 | 6.51 | 6.523 | 6.5165 | 694.8493 | 1.806 | 1.81 | 1.808 | 196.8795 | 0.28334
2020-02-14 | 12:55 | 6.523 | 6.531 | 6.527 | 695.9689 | 1.81 | 1.811 | 1.8105 | 197.1517 | 0.28328
2020-02-14 | 12:55 | 6.533 | 6.541 | 6.537 | 697.0352 | 1.812 | 1.815 | 1.8135 | 197.4784 | 0.28331
2020-02-14 | 12:55 | 6.544 | 6.551 | 6.5475 | 698.1548 | 1.816 | 1.817 | 1.8165 | 197.805 | 0.28333
2020-02-14 | 12:55 | 6.55 | 6.553 | 6.5515 | 698.5814 | 1.817 | 1.819 | 1.818 | 197.9684 | 0.28339
2020-02-14 | 12:55 | 6.539 | 6.55 | 6.5445 | 697.8349 | 1.813 | 1.818 | 1.8155 | 197.6962 | 0.2833
2020-02-14 | 12:55 | 6.53 | 6.537 | 6.5335 | 696.662 | 1.811 | 1.813 | 1.812 | 197.315 | 0.28323
2020-02-14 | 12:55 | 6.525 | 6.528 | 6.5265 | 695.9156 | 1.81 | 1.812 | 1.811 | 197.2061 | 0.28338
2020-02-14 | 12:55 | 6.522 | 6.525 | 6.5235 | 695.5957 | 1.809 | 1.81 | 1.8095 | 197.0428 | 0.28327
2020-02-14 | 12:55 | 6.521 | 6.524 | 6.5225 | 695.4891 | 1.808 | 1.81 | 1.809 | 196.9883 | 0.28324
2020-02-14 | 12:56 | 6.525 | 6.528 | 6.5265 | 695.9156 | 1.81 | 1.81 | 1.81 | 197.0972 | 0.28322
2020-02-14 | 12:56 | 6.53 | 6.532 | 6.531 | 696.3955 | 1.811 | 1.812 | 1.8115 | 197.2606 | 0.28326
2020-02-14 | 12:56 | 6.533 | 6.538 | 6.5355 | 696.8753 | 1.812 | 1.814 | 1.813 | 197.4239 | 0.2833
2020-02-14 | 12:56 | 6.539 | 6.54 | 6.5395 | 697.3018 | 1.813 | 1.815 | 1.814 | 197.5328 | 0.28328
2020-02-14 | 12:56 | 6.541 | 6.543 | 6.542 | 697.5684 | 1.814 | 1.815 | 1.8145 | 197.5873 | 0.28325
2020-02-14 | 12:56 | 6.544 | 6.548 | 6.546 | 697.9949 | 1.814 | 1.815 | 1.8145 | 197.5873 | 0.28308
2020-02-14 | 12:56 | 6.547 | 6.55 | 6.5485 | 698.2615 | 1.815 | 1.818 | 1.8165 | 197.805 | 0.28328
2020-02-14 | 12:56 | 6.549 | 6.55 | 6.5495 | 698.3681 | 1.816 | 1.817 | 1.8165 | 197.805 | 0.28324
2020-02-14 | 12:56 | 6.546 | 6.55 | 6.548 | 698.2082 | 1.815 | 1.816 | 1.8155 | 197.6962 | 0.28315
2020-02-14 | 12:56 | 6.546 | 6.548 | 6.547 | 698.1015 | 1.814 | 1.816 | 1.815 | 197.6417 | 0.28311
2020-02-14 | 12:56 | 6.547 | 6.549 | 6.548 | 698.2082 | 1.814 | 1.816 | 1.815 | 197.6417 | 0.28307
2020-02-14 | 12:56 | 6.548 | 6.55 | 6.549 | 698.3148 | 1.815 | 1.816 | 1.8155 | 197.6962 | 0.2831
2020-02-14 | 12:56 | 6.551 | 6.556 | 6.5535 | 698.7946 | 1.816 | 1.818 | 1.817 | 197.8595 | 0.28314
2020-02-14 | 12:57 | 6.557 | 6.561 | 6.559 | 699.3811 | 1.817 | 1.818 | 1.8175 | 197.9139 | 0.28298
2020-02-14 | 12:57 | 6.564 | 6.566 | 6.565 | 700.0208 | 1.818 | 1.821 | 1.8195 | 198.1317 | 0.28304
2020-02-14 | 12:57 | 6.568 | 6.57 | 6.569 | 700.4474 | 1.821 | 1.821 | 1.821 | 198.2951 | 0.2831
2020-02-14 | 12:57 | 6.562 | 6.569 | 6.5655 | 700.0742 | 1.819 | 1.822 | 1.8205 | 198.2406 | 0.28317
2020-02-14 | 12:57 | 6.559 | 6.561 | 6.56 | 699.4877 | 1.818 | 1.82 | 1.819 | 198.0773 | 0.28317
2020-02-14 | 12:57 | 6.56 | 6.563 | 6.5615 | 699.6476 | 1.818 | 1.821 | 1.8195 | 198.1317 | 0.28319
2020-02-14 | 12:57 | 6.556 | 6.559 | 6.5575 | 699.2211 | 1.818 | 1.82 | 1.819 | 198.0773 | 0.28328
2020-02-14 | 12:57 | 6.556 | 6.562 | 6.559 | 699.3811 | 1.818 | 1.82 | 1.819 | 198.0773 | 0.28322
2020-02-14 | 12:57 | 6.562 | 6.564 | 6.563 | 699.8076 | 1.819 | 1.821 | 1.82 | 198.1862 | 0.2832
2020-02-14 | 12:57 | 6.558 | 6.562 | 6.56 | 699.4877 | 1.818 | 1.82 | 1.819 | 198.0773 | 0.28317
2020-02-14 | 12:57 | 6.556 | 6.56 | 6.558 | 699.2744 | 1.817 | 1.819 | 1.818 | 197.9684 | 0.28311
2020-02-14 | 12:58 | 6.561 | 6.563 | 6.562 | 699.701 | 1.819 | 1.821 | 1.82 | 198.1862 | 0.28324
2020-02-14 | 12:58 | 6.564 | 6.57 | 6.567 | 700.2341 | 1.82 | 1.822 | 1.821 | 198.2951 | 0.28318
2020-02-14 | 12:58 | 6.568 | 6.57 | 6.569 | 700.4474 | 1.821 | 1.822 | 1.8215 | 198.3495 | 0.28318
2020-02-14 | 12:58 | 6.564 | 6.57 | 6.567 | 700.2341 | 1.821 | 1.823 | 1.822 | 198.404 | 0.28334
2020-02-14 | 12:58 | 6.55 | 6.563 | 6.5565 | 699.1145 | 1.816 | 1.82 | 1.818 | 197.9684 | 0.28317
2020-02-14 | 12:58 | 6.518 | 6.545 | 6.5315 | 696.4488 | 1.807 | 1.814 | 1.8105 | 197.1517 | 0.28308
2020-02-14 | 12:58 | 6.485 | 6.51 | 6.4975 | 692.8234 | 1.798 | 1.806 | 1.802 | 196.2261 | 0.28323
2020-02-14 | 12:58 | 6.481 | 6.483 | 6.482 | 691.1706 | 1.798 | 1.799 | 1.7985 | 195.845 | 0.28335
2020-02-14 | 12:58 | 6.484 | 6.499 | 6.4915 | 692.1836 | 1.798 | 1.803 | 1.8005 | 196.0628 | 0.28325
2020-02-14 | 12:58 | 6.491 | 6.5 | 6.4955 | 692.6101 | 1.801 | 1.803 | 1.802 | 196.2261 | 0.28331
2020-02-14 | 12:58 | 6.482 | 6.49 | 6.486 | 691.5971 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.28341
2020-02-14 | 12:58 | 6.483 | 6.494 | 6.4885 | 691.8637 | 1.799 | 1.801 | 1.8 | 196.0083 | 0.2833
2020-02-14 | 12:58 | 6.499 | 6.504 | 6.5015 | 693.2499 | 1.791 | 1.799 | 1.795 | 195.4638 | 0.28195
2020-02-14 | 12:59 | 6.506 | 6.518 | 6.512 | 694.3695 | 1.8 | 1.809 | 1.8045 | 196.4983 | 0.28299
2020-02-14 | 12:59 | 6.516 | 6.523 | 6.5195 | 695.1692 | 1.811 | 1.812 | 1.8115 | 197.2606 | 0.28376
2020-02-14 | 12:59 | 6.505 | 6.526 | 6.5155 | 694.7427 | 1.805 | 1.813 | 1.809 | 196.9883 | 0.28354
2020-02-14 | 12:59 | 6.438 | 6.49 | 6.464 | 689.2513 | 1.786 | 1.801 | 1.7935 | 195.3005 | 0.28335
2020-02-14 | 12:59 | 6.423 | 6.43 | 6.4265 | 685.2527 | 1.783 | 1.785 | 1.784 | 194.266 | 0.2835
2020-02-14 | 12:59 | 6.371 | 6.419 | 6.395 | 681.8939 | 1.769 | 1.782 | 1.7755 | 193.3404 | 0.28353
2020-02-14 | 12:59 | 6.266 | 6.353 | 6.3095 | 672.7771 | 1.737 | 1.763 | 1.75 | 190.5636 | 0.28325
2020-02-14 | 12:59 | 6.174 | 6.246 | 6.21 | 662.1675 | 1.712 | 1.731 | 1.7215 | 187.4602 | 0.2831
2020-02-14 | 12:59 | 6.104 | 6.156 | 6.13 | 653.6371 | 1.691 | 1.706 | 1.6985 | 184.9556 | 0.28296
2020-02-14 | 12:59 | 6.054 | 6.096 | 6.075 | 647.7725 | 1.676 | 1.69 | 1.683 | 183.2678 | 0.28292
2020-02-14 | 12:59 | 5.98 | 6.041 | 6.0105 | 640.8949 | 1.656 | 1.673 | 1.6645 | 181.2532 | 0.28281
2020-02-14 | 13:00 | 5.865 | 5.959 | 5.912 | 630.392 | 1.622 | 1.649 | 1.6355 | 178.0953 | 0.28252
2020-02-14 | 13:00 | 5.761 | 5.843 | 5.802 | 618.6628 | 1.594 | 1.616 | 1.605 | 174.7741 | 0.2825
2020-02-14 | 13:00 | 5.697 | 5.745 | 5.721 | 610.0258 | 1.577 | 1.591 | 1.584 | 172.4873 | 0.28275
2020-02-14 | 13:00 | 5.686 | 5.693 | 5.6895 | 606.667 | 1.574 | 1.576 | 1.575 | 171.5073 | 0.2827
2020-02-14 | 13:00 | 5.626 | 5.676 | 5.651 | 602.5617 | 1.557 | 1.572 | 1.5645 | 170.3639 | 0.28273
2020-02-14 | 13:00 | 5.6 | 5.619 | 5.6095 | 598.1366 | 1.551 | 1.556 | 1.5535 | 169.1661 | 0.28282
2020-02-14 | 13:00 | 5.587 | 5.595 | 5.591 | 596.164 | 1.547 | 1.549 | 1.548 | 168.5671 | 0.28275
2020-02-14 | 13:00 | 5.588 | 5.594 | 5.591 | 596.164 | 1.546 | 1.55 | 1.548 | 168.5671 | 0.28275
2020-02-14 | 13:00 | 5.526 | 5.579 | 5.5525 | 592.0588 | 1.531 | 1.545 | 1.538 | 167.4782 | 0.28287
2020-02-14 | 13:00 | 5.489 | 5.512 | 5.5005 | 586.514 | 1.522 | 1.525 | 1.5235 | 165.8993 | 0.28286
2020-02-14 | 13:00 | 5.497 | 5.549 | 5.523 | 588.9132 | 1.523 | 1.539 | 1.531 | 166.716 | 0.28309
2020-02-14 | 13:00 | 5.571 | 5.719 | 5.645 | 601.922 | 1.546 | 1.59 | 1.568 | 170.745 | 0.28367
2020-02-14 | 13:00 | 5.765 | 5.938 | 5.8515 | 623.9409 | 1.602 | 1.653 | 1.6275 | 177.2242 | 0.28404
2020-02-14 | 13:01 | 5.969 | 6.059 | 6.014 | 641.2681 | 1.661 | 1.686 | 1.6735 | 182.2333 | 0.28418
2020-02-14 | 13:01 | 6.082 | 6.208 | 6.145 | 655.2366 | 1.694 | 1.729 | 1.7115 | 186.3712 | 0.28443
2020-02-14 | 13:01 | 6.236 | 6.288 | 6.262 | 667.7122 | 1.738 | 1.751 | 1.7445 | 189.9647 | 0.2845
2020-02-14 | 13:01 | 6.291 | 6.325 | 6.308 | 672.6171 | 1.753 | 1.763 | 1.758 | 191.4348 | 0.28461
2020-02-14 | 13:01 | 6.338 | 6.36 | 6.349 | 676.9889 | 1.765 | 1.772 | 1.7685 | 192.5782 | 0.28446
2020-02-14 | 13:01 | 6.365 | 6.379 | 6.372 | 679.4414 | 1.773 | 1.777 | 1.775 | 193.286 | 0.28448
2020-02-14 | 13:01 | 6.384 | 6.426 | 6.405 | 682.9602 | 1.779 | 1.791 | 1.785 | 194.3749 | 0.28461
2020-02-14 | 13:01 | 6.428 | 6.439 | 6.4335 | 685.9991 | 1.791 | 1.795 | 1.793 | 195.246 | 0.28462
2020-02-14 | 13:01 | 6.398 | 6.418 | 6.408 | 683.2801 | 1.783 | 1.789 | 1.786 | 194.4838 | 0.28463
2020-02-14 | 13:01 | 6.417 | 6.451 | 6.434 | 686.0524 | 1.789 | 1.798 | 1.7935 | 195.3005 | 0.28467
2020-02-14 | 13:01 | 6.411 | 6.448 | 6.4295 | 685.5726 | 1.786 | 1.796 | 1.791 | 195.0283 | 0.28447
2020-02-14 | 13:02 | 6.375 | 6.402 | 6.3885 | 681.2008 | 1.776 | 1.784 | 1.78 | 193.8304 | 0.28454
2020-02-14 | 13:02 | 6.372 | 6.389 | 6.3805 | 680.3478 | 1.775 | 1.781 | 1.778 | 193.6126 | 0.28458
2020-02-14 | 13:02 | 6.377 | 6.393 | 6.385 | 680.8276 | 1.776 | 1.781 | 1.7785 | 193.6671 | 0.28446
2020-02-14 | 13:02 | 6.359 | 6.37 | 6.3645 | 678.6417 | 1.772 | 1.774 | 1.773 | 193.0682 | 0.28449
2020-02-14 | 13:02 | 6.356 | 6.358 | 6.357 | 677.842 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.28467
2020-02-14 | 13:02 | 6.347 | 6.358 | 6.3525 | 677.3621 | 1.768 | 1.773 | 1.7705 | 192.7959 | 0.28463
2020-02-14 | 13:02 | 6.274 | 6.336 | 6.305 | 672.2973 | 1.748 | 1.768 | 1.758 | 191.4348 | 0.28475
2020-02-14 | 13:02 | 6.198 | 6.258 | 6.228 | 664.0868 | 1.727 | 1.744 | 1.7355 | 188.9847 | 0.28458
2020-02-14 | 13:02 | 6.109 | 6.184 | 6.1465 | 655.3965 | 1.703 | 1.724 | 1.7135 | 186.589 | 0.2847
2020-02-14 | 13:02 | 6.05 | 6.094 | 6.072 | 647.4526 | 1.686 | 1.698 | 1.692 | 184.2478 | 0.28457
2020-02-14 | 13:02 | 6.051 | 6.109 | 6.08 | 648.3057 | 1.687 | 1.705 | 1.696 | 184.6834 | 0.28487
2020-02-14 | 13:02 | 6.125 | 6.151 | 6.138 | 654.4902 | 1.708 | 1.717 | 1.7125 | 186.4801 | 0.28492
2020-02-14 | 13:02 | 6.14 | 6.15 | 6.145 | 655.2366 | 1.713 | 1.717 | 1.715 | 186.7524 | 0.28502
2020-02-14 | 13:03 | 6.148 | 6.203 | 6.1755 | 658.4888 | 1.716 | 1.731 | 1.7235 | 187.678 | 0.28501
2020-02-14 | 13:03 | 6.212 | 6.237 | 6.2245 | 663.7136 | 1.735 | 1.741 | 1.738 | 189.2569 | 0.28515
2020-02-14 | 13:03 | 6.224 | 6.239 | 6.2315 | 664.46 | 1.738 | 1.742 | 1.74 | 189.4747 | 0.28516
2020-02-14 | 13:03 | 6.138 | 6.213 | 6.1755 | 658.4888 | 1.712 | 1.734 | 1.723 | 187.6235 | 0.28493
2020-02-14 | 13:03 | 6.009 | 6.11 | 6.0595 | 646.1198 | 1.676 | 1.705 | 1.6905 | 184.0845 | 0.28491
2020-02-14 | 13:03 | 5.888 | 5.986 | 5.937 | 633.0577 | 1.641 | 1.67 | 1.6555 | 180.2732 | 0.28477
2020-02-14 | 13:03 | 5.753 | 5.862 | 5.8075 | 619.2492 | 1.603 | 1.632 | 1.6175 | 176.1352 | 0.28443
2020-02-14 | 13:03 | 5.613 | 5.726 | 5.6695 | 604.5344 | 1.562 | 1.595 | 1.5785 | 171.8884 | 0.28433
2020-02-14 | 13:03 | 5.538 | 5.586 | 5.562 | 593.0717 | 1.542 | 1.555 | 1.5485 | 168.6216 | 0.28432
2020-02-14 | 13:03 | 5.513 | 5.533 | 5.523 | 588.9132 | 1.534 | 1.539 | 1.5365 | 167.3149 | 0.28411
2020-02-14 | 13:03 | 5.434 | 5.507 | 5.4705 | 583.3152 | 1.513 | 1.534 | 1.5235 | 165.8993 | 0.28441
2020-02-14 | 13:04 | 5.3 | 5.408 | 5.354 | 570.8929 | 1.474 | 1.505 | 1.4895 | 162.1969 | 0.28411
2020-02-14 | 13:04 | 5.142 | 5.269 | 5.2055 | 555.0584 | 1.43 | 1.465 | 1.4475 | 157.6233 | 0.28398
2020-02-14 | 13:04 | 5.009 | 5.111 | 5.06 | 539.5439 | 1.392 | 1.421 | 1.4065 | 153.1587 | 0.28387
2020-02-14 | 13:04 | 4.951 | 4.994 | 4.9725 | 530.2138 | 1.378 | 1.389 | 1.3835 | 150.6542 | 0.28414
2020-02-14 | 13:04 | 4.949 | 5.007 | 4.978 | 530.8003 | 1.378 | 1.395 | 1.3865 | 150.9808 | 0.28444
2020-02-14 | 13:04 | 5.032 | 5.123 | 5.0775 | 541.4099 | 1.401 | 1.427 | 1.414 | 153.9754 | 0.2844
2020-02-14 | 13:04 | 5.137 | 5.156 | 5.1465 | 548.7673 | 1.432 | 1.437 | 1.4345 | 156.2077 | 0.28465
2020-02-14 | 13:04 | 5.117 | 5.149 | 5.133 | 547.3278 | 1.426 | 1.435 | 1.4305 | 155.7722 | 0.2846
2020-02-14 | 13:04 | 5.056 | 5.111 | 5.0835 | 542.0497 | 1.408 | 1.424 | 1.416 | 154.1932 | 0.28446
2020-02-14 | 13:04 | 4.951 | 5.035 | 4.993 | 532.3997 | 1.38 | 1.404 | 1.392 | 151.5798 | 0.28471
2020-02-14 | 13:04 | 4.916 | 4.939 | 4.9275 | 525.4155 | 1.369 | 1.376 | 1.3725 | 149.4563 | 0.28445
2020-02-14 | 13:04 | 4.897 | 4.912 | 4.9045 | 522.963 | 1.366 | 1.369 | 1.3675 | 148.9119 | 0.28475
2020-02-14 | 13:04 | 4.897 | 4.914 | 4.9055 | 523.0697 | 1.366 | 1.372 | 1.369 | 149.0752 | 0.285
2020-02-14 | 13:05 | 4.867 | 4.915 | 4.891 | 521.5235 | 1.356 | 1.37 | 1.363 | 148.4218 | 0.28459
2020-02-14 | 13:05 | 4.818 | 4.851 | 4.8345 | 515.499 | 1.344 | 1.354 | 1.349 | 146.8973 | 0.28496
2020-02-14 | 13:05 | 4.731 | 4.81 | 4.7705 | 508.6747 | 1.32 | 1.342 | 1.331 | 144.9373 | 0.28493
2020-02-14 | 13:05 | 4.634 | 4.708 | 4.671 | 498.0651 | 1.293 | 1.316 | 1.3045 | 142.0516 | 0.28521
2020-02-14 | 13:05 | 4.586 | 4.618 | 4.602 | 490.7077 | 1.279 | 1.29 | 1.2845 | 139.8737 | 0.28504
2020-02-14 | 13:05 | 4.534 | 4.58 | 4.557 | 485.9094 | 1.264 | 1.278 | 1.271 | 138.4036 | 0.28483
2020-02-14 | 13:05 | 4.46 | 4.512 | 4.486 | 478.3387 | 1.246 | 1.259 | 1.2525 | 136.3891 | 0.28513
2020-02-14 | 13:05 | 4.472 | 4.588 | 4.53 | 483.0304 | 1.25 | 1.284 | 1.267 | 137.9681 | 0.28563
2020-02-14 | 13:05 | 4.626 | 4.756 | 4.691 | 500.1977 | 1.296 | 1.333 | 1.3145 | 143.1405 | 0.28617
2020-02-14 | 13:05 | 4.773 | 4.824 | 4.7985 | 511.6603 | 1.339 | 1.353 | 1.346 | 146.5707 | 0.28646
2020-02-14 | 13:05 | 4.835 | 4.86 | 4.8475 | 516.8852 | 1.356 | 1.364 | 1.36 | 148.0952 | 0.28651
2020-02-14 | 13:06 | 4.862 | 4.866 | 4.864 | 518.6445 | 1.364 | 1.366 | 1.365 | 148.6396 | 0.28659
2020-02-14 | 13:06 | 4.812 | 4.856 | 4.834 | 515.4457 | 1.351 | 1.364 | 1.3575 | 147.8229 | 0.28679
2020-02-14 | 13:06 | 4.761 | 4.798 | 4.7795 | 509.6344 | 1.337 | 1.347 | 1.342 | 146.1351 | 0.28674
2020-02-14 | 13:06 | 4.685 | 4.749 | 4.717 | 502.97 | 1.315 | 1.333 | 1.324 | 144.175 | 0.28665
2020-02-14 | 13:06 | 4.64 | 4.672 | 4.656 | 496.4657 | 1.302 | 1.312 | 1.307 | 142.3238 | 0.28667
2020-02-14 | 13:06 | 4.611 | 4.634 | 4.6225 | 492.8936 | 1.295 | 1.301 | 1.298 | 141.3438 | 0.28676
2020-02-14 | 13:06 | 4.55 | 4.603 | 4.5765 | 487.9886 | 1.277 | 1.294 | 1.2855 | 139.9826 | 0.28686
2020-02-14 | 13:06 | 4.414 | 4.531 | 4.4725 | 476.8992 | 1.239 | 1.271 | 1.255 | 136.6613 | 0.28656
2020-02-14 | 13:06 | 4.301 | 4.387 | 4.344 | 463.1973 | 1.207 | 1.23 | 1.2185 | 132.6867 | 0.28646
2020-02-14 | 13:06 | 4.227 | 4.285 | 4.256 | 453.814 | 1.187 | 1.201 | 1.194 | 130.0188 | 0.2865
2020-02-14 | 13:06 | 4.181 | 4.216 | 4.1985 | 447.6828 | 1.172 | 1.184 | 1.178 | 128.2765 | 0.28653
2020-02-14 | 13:06 | 4.152 | 4.175 | 4.1635 | 443.9508 | 1.166 | 1.172 | 1.169 | 127.2965 | 0.28674
2020-02-14 | 13:06 | 4.13 | 4.147 | 4.1385 | 441.285 | 1.16 | 1.164 | 1.162 | 126.5342 | 0.28674
2020-02-14 | 13:07 | 4.126 | 4.135 | 4.1305 | 440.432 | 1.159 | 1.162 | 1.1605 | 126.3709 | 0.28692
2020-02-14 | 13:07 | 4.14 | 4.171 | 4.1555 | 443.0977 | 1.164 | 1.176 | 1.17 | 127.4054 | 0.28753
2020-02-14 | 13:07 | 4.18 | 4.213 | 4.1965 | 447.4695 | 1.178 | 1.187 | 1.1825 | 128.7666 | 0.28777
2020-02-14 | 13:07 | 4.22 | 4.241 | 4.2305 | 451.0949 | 1.19 | 1.197 | 1.1935 | 129.9644 | 0.28811
2020-02-14 | 13:07 | 4.231 | 4.241 | 4.236 | 451.6814 | 1.195 | 1.198 | 1.1965 | 130.2911 | 0.28846
2020-02-14 | 13:07 | 4.212 | 4.228 | 4.22 | 449.9753 | 1.191 | 1.194 | 1.1925 | 129.8555 | 0.28858
2020-02-14 | 13:07 | 4.188 | 4.208 | 4.198 | 447.6295 | 1.183 | 1.19 | 1.1865 | 129.2021 | 0.28864
2020-02-14 | 13:07 | 4.164 | 4.182 | 4.173 | 444.9637 | 1.177 | 1.183 | 1.18 | 128.4943 | 0.28877
2020-02-14 | 13:07 | 4.132 | 4.16 | 4.146 | 442.0848 | 1.168 | 1.175 | 1.1715 | 127.5687 | 0.28856
2020-02-14 | 13:07 | 4.11 | 4.125 | 4.1175 | 439.0458 | 1.164 | 1.166 | 1.165 | 126.8609 | 0.28895
2020-02-14 | 13:07 | 4.089 | 4.106 | 4.0975 | 436.9132 | 1.158 | 1.161 | 1.1595 | 126.262 | 0.28899
2020-02-14 | 13:08 | 4.086 | 4.087 | 4.0865 | 435.7403 | 1.157 | 1.158 | 1.1575 | 126.0442 | 0.28926
2020-02-14 | 13:08 | 4.086 | 4.088 | 4.087 | 435.7936 | 1.157 | 1.158 | 1.1575 | 126.0442 | 0.28923
2020-02-14 | 13:08 | 4.079 | 4.089 | 4.084 | 435.4737 | 1.16 | 1.161 | 1.1605 | 126.3709 | 0.29019
2020-02-14 | 13:08 | 4.046 | 4.071 | 4.0585 | 432.7547 | 1.15 | 1.156 | 1.153 | 125.5542 | 0.29013
2020-02-14 | 13:08 | 4.025 | 4.041 | 4.033 | 430.0357 | 1.145 | 1.149 | 1.147 | 124.9008 | 0.29044
2020-02-14 | 13:08 | 4.006 | 4.02 | 4.013 | 427.9031 | 1.141 | 1.144 | 1.1425 | 124.4108 | 0.29075
2020-02-14 | 13:08 | 3.995 | 4.006 | 4.0005 | 426.5702 | 1.137 | 1.141 | 1.139 | 124.0297 | 0.29076
2020-02-14 | 13:08 | 3.972 | 3.991 | 3.9815 | 424.5443 | 1.131 | 1.136 | 1.1335 | 123.4308 | 0.29074
2020-02-14 | 13:08 | 3.949 | 3.967 | 3.958 | 422.0385 | 1.125 | 1.13 | 1.1275 | 122.7774 | 0.29092
2020-02-14 | 13:08 | 3.919 | 3.943 | 3.931 | 419.1595 | 1.118 | 1.124 | 1.121 | 122.0696 | 0.29122
2020-02-14 | 13:08 | 3.886 | 3.912 | 3.899 | 415.7473 | 1.107 | 1.115 | 1.111 | 120.9807 | 0.291
2020-02-14 | 13:08 | 3.864 | 3.878 | 3.871 | 412.7617 | 1.103 | 1.107 | 1.105 | 120.3273 | 0.29152
2020-02-14 | 13:08 | 3.856 | 3.862 | 3.859 | 411.4822 | 1.102 | 1.103 | 1.1025 | 120.0551 | 0.29176
2020-02-14 | 13:09 | 3.856 | 3.869 | 3.8625 | 411.8554 | 1.104 | 1.107 | 1.1055 | 120.3818 | 0.29229
2020-02-14 | 13:09 | 3.876 | 3.905 | 3.8905 | 414.841 | 1.113 | 1.119 | 1.116 | 121.5251 | 0.29294
2020-02-14 | 13:09 | 3.913 | 3.93 | 3.9215 | 418.1465 | 1.124 | 1.131 | 1.1275 | 122.7774 | 0.29362
2020-02-14 | 13:09 | 3.935 | 3.952 | 3.9435 | 420.4923 | 1.133 | 1.137 | 1.135 | 123.5941 | 0.29393
2020-02-14 | 13:09 | 3.957 | 3.981 | 3.969 | 423.2114 | 1.14 | 1.147 | 1.1435 | 124.5197 | 0.29423
2020-02-14 | 13:09 | 3.985 | 4.01 | 3.9975 | 426.2503 | 1.149 | 1.156 | 1.1525 | 125.4998 | 0.29443
2020-02-14 | 13:09 | 4.018 | 4.061 | 4.0395 | 430.7287 | 1.159 | 1.174 | 1.1665 | 127.0243 | 0.29491
2020-02-14 | 13:09 | 4.068 | 4.091 | 4.0795 | 434.9939 | 1.177 | 1.185 | 1.181 | 128.6032 | 0.29564
2020-02-14 | 13:09 | 4.093 | 4.101 | 4.097 | 436.8599 | 1.184 | 1.188 | 1.186 | 129.1477 | 0.29563
2020-02-14 | 13:09 | 4.095 | 4.103 | 4.099 | 437.0732 | 1.187 | 1.189 | 1.188 | 129.3655 | 0.29598
2020-02-14 | 13:09 | 4.048 | 4.089 | 4.0685 | 433.821 | 1.174 | 1.185 | 1.1795 | 128.4399 | 0.29607
2020-02-14 | 13:10 | 4.001 | 4.036 | 4.0185 | 428.4895 | 1.164 | 1.172 | 1.168 | 127.1876 | 0.29683
2020-02-14 | 13:10 | 3.97 | 3.995 | 3.9825 | 424.6509 | 1.155 | 1.161 | 1.158 | 126.0987 | 0.29695
2020-02-14 | 13:10 | 3.962 | 3.965 | 3.9635 | 422.6249 | 1.151 | 1.155 | 1.153 | 125.5542 | 0.29708
2020-02-14 | 13:10 | 3.944 | 3.965 | 3.9545 | 421.6653 | 1.145 | 1.153 | 1.149 | 125.1186 | 0.29673
2020-02-14 | 13:10 | 3.885 | 3.934 | 3.9095 | 416.8669 | 1.131 | 1.143 | 1.137 | 123.8119 | 0.29701
2020-02-14 | 13:10 | 3.834 | 3.874 | 3.854 | 410.949 | 1.114 | 1.129 | 1.1215 | 122.1241 | 0.29718
2020-02-14 | 13:10 | 3.767 | 3.822 | 3.7945 | 404.6046 | 1.099 | 1.113 | 1.106 | 120.4362 | 0.29766
2020-02-14 | 13:10 | 3.675 | 3.751 | 3.713 | 395.9143 | 1.074 | 1.093 | 1.0835 | 117.9861 | 0.29801
2020-02-14 | 13:10 | 3.583 | 3.655 | 3.619 | 385.8912 | 1.045 | 1.068 | 1.0565 | 115.046 | 0.29813
2020-02-14 | 13:10 | 3.53 | 3.571 | 3.5505 | 378.5871 | 1.033 | 1.042 | 1.0375 | 112.977 | 0.29842
2020-02-14 | 13:10 | 3.47 | 3.521 | 3.4955 | 372.7224 | 1.018 | 1.03 | 1.024 | 111.5069 | 0.29917
2020-02-14 | 13:10 | 3.382 | 3.451 | 3.4165 | 364.2987 | 0.991 | 1.014 | 1.0025 | 109.1657 | 0.29966
2020-02-14 | 13:10 | 3.303 | 3.365 | 3.334 | 355.5018 | 0.97 | 0.986 | 0.978 | 106.4978 | 0.29957
2020-02-14 | 13:11 | 3.272 | 3.295 | 3.2835 | 350.1171 | 0.963 | 0.967 | 0.965 | 105.0822 | 0.30013
2020-02-14 | 13:11 | 3.261 | 3.27 | 3.2655 | 348.1977 | 0.959 | 0.962 | 0.9605 | 104.5922 | 0.30038
2020-02-14 | 13:11 | 3.25 | 3.258 | 3.254 | 346.9715 | 0.959 | 0.962 | 0.9605 | 104.5922 | 0.30144
2020-02-14 | 13:11 | 3.233 | 3.245 | 3.239 | 345.3721 | 0.959 | 0.961 | 0.96 | 104.5378 | 0.30268
2020-02-14 | 13:11 | 3.216 | 3.229 | 3.2225 | 343.6127 | 0.956 | 0.961 | 0.9585 | 104.3744 | 0.30376
2020-02-14 | 13:11 | 3.204 | 3.212 | 3.208 | 342.0665 | 0.954 | 0.958 | 0.956 | 104.1022 | 0.30433
2020-02-14 | 13:11 | 3.191 | 3.201 | 3.196 | 340.787 | 0.951 | 0.955 | 0.953 | 103.7755 | 0.30452
2020-02-14 | 13:11 | 3.179 | 3.189 | 3.184 | 339.5074 | 0.949 | 0.952 | 0.9505 | 103.5033 | 0.30486
2020-02-14 | 13:11 | 3.168 | 3.177 | 3.1725 | 338.2812 | 0.954 | 0.956 | 0.955 | 103.9933 | 0.30742
2020-02-14 | 13:11 | 3.159 | 3.167 | 3.163 | 337.2682 | 0.952 | 0.956 | 0.954 | 103.8844 | 0.30802
2020-02-14 | 13:11 | 3.157 | 3.161 | 3.159 | 336.8417 | 0.955 | 0.96 | 0.9575 | 104.2655 | 0.30954
2020-02-14 | 13:12 | 3.162 | 3.167 | 3.1645 | 337.4282 | 0.965 | 0.984 | 0.9745 | 106.1167 | 0.31449
2020-02-14 | 13:12 | 3.168 | 3.176 | 3.172 | 338.2279 | 0.986 | 0.991 | 0.9885 | 107.6412 | 0.31825
2020-02-14 | 13:12 | 3.179 | 3.188 | 3.1835 | 339.4541 | 0.992 | 0.995 | 0.9935 | 108.1857 | 0.3187
2020-02-14 | 13:12 | 3.191 | 3.204 | 3.1975 | 340.9469 | 0.995 | 0.999 | 0.997 | 108.5668 | 0.31843
2020-02-14 | 13:12 | 3.209 | 3.232 | 3.2205 | 343.3994 | 1.001 | 1.007 | 1.004 | 109.3291 | 0.31837
2020-02-14 | 13:12 | 3.235 | 3.255 | 3.245 | 346.0118 | 1.009 | 1.016 | 1.0125 | 110.2547 | 0.31864
2020-02-14 | 13:12 | 3.26 | 3.272 | 3.266 | 348.251 | 1.017 | 1.02 | 1.0185 | 110.908 | 0.31847
2020-02-14 | 13:12 | 3.275 | 3.281 | 3.278 | 349.5306 | 1.021 | 1.024 | 1.0225 | 111.3436 | 0.31855
2020-02-14 | 13:12 | 3.281 | 3.284 | 3.2825 | 350.0104 | 1.022 | 1.024 | 1.023 | 111.3981 | 0.31827
2020-02-14 | 13:12 | 3.273 | 3.281 | 3.277 | 349.424 | 1.021 | 1.023 | 1.022 | 111.2892 | 0.31849
2020-02-14 | 13:12 | 3.266 | 3.271 | 3.2685 | 348.5176 | 1.018 | 1.02 | 1.019 | 110.9625 | 0.31838
2020-02-14 | 13:12 | 3.266 | 3.275 | 3.2705 | 348.7309 | 1.018 | 1.021 | 1.0195 | 111.0169 | 0.31835
2020-02-14 | 13:12 | 3.28 | 3.304 | 3.292 | 351.0234 | 1.023 | 1.031 | 1.027 | 111.8336 | 0.31859
2020-02-14 | 13:13 | 3.311 | 3.34 | 3.3255 | 354.5955 | 1.033 | 1.041 | 1.037 | 112.9226 | 0.31845
2020-02-14 | 13:13 | 3.349 | 3.376 | 3.3625 | 358.5408 | 1.045 | 1.053 | 1.049 | 114.2293 | 0.31859
2020-02-14 | 13:13 | 3.383 | 3.403 | 3.393 | 361.793 | 1.055 | 1.061 | 1.058 | 115.2093 | 0.31844
2020-02-14 | 13:13 | 3.406 | 3.419 | 3.4125 | 363.8722 | 1.062 | 1.066 | 1.064 | 115.8627 | 0.31842
2020-02-14 | 13:13 | 3.423 | 3.444 | 3.4335 | 366.1114 | 1.068 | 1.074 | 1.071 | 116.6249 | 0.31855
2020-02-14 | 13:13 | 3.448 | 3.467 | 3.4575 | 368.6705 | 1.076 | 1.082 | 1.079 | 117.4961 | 0.3187
2020-02-14 | 13:13 | 3.472 | 3.495 | 3.4835 | 371.4429 | 1.082 | 1.091 | 1.0865 | 118.3128 | 0.31852
2020-02-14 | 13:13 | 3.502 | 3.535 | 3.5185 | 375.1749 | 1.092 | 1.104 | 1.098 | 119.5651 | 0.31869
2020-02-14 | 13:13 | 3.544 | 3.575 | 3.5595 | 379.5467 | 1.105 | 1.116 | 1.1105 | 120.9262 | 0.31861
2020-02-14 | 13:13 | 3.581 | 3.606 | 3.5935 | 383.1721 | 1.117 | 1.125 | 1.121 | 122.0696 | 0.31858
2020-02-14 | 13:13 | 3.611 | 3.633 | 3.622 | 386.211 | 1.127 | 1.132 | 1.1295 | 122.9952 | 0.31847
2020-02-14 | 13:14 | 3.636 | 3.645 | 3.6405 | 388.1837 | 1.134 | 1.138 | 1.136 | 123.703 | 0.31867
2020-02-14 | 13:14 | 3.649 | 3.671 | 3.66 | 390.263 | 1.138 | 1.147 | 1.1425 | 124.4108 | 0.31879
2020-02-14 | 13:14 | 3.678 | 3.702 | 3.69 | 393.4618 | 1.148 | 1.155 | 1.1515 | 125.3909 | 0.31869
2020-02-14 | 13:14 | 3.709 | 3.731 | 3.72 | 396.6607 | 1.158 | 1.164 | 1.161 | 126.4254 | 0.31872
2020-02-14 | 13:14 | 3.734 | 3.756 | 3.745 | 399.3264 | 1.165 | 1.172 | 1.1685 | 127.2421 | 0.31864
2020-02-14 | 13:14 | 3.763 | 3.782 | 3.7725 | 402.2587 | 1.174 | 1.181 | 1.1775 | 128.2221 | 0.31876
2020-02-14 | 13:14 | 3.788 | 3.811 | 3.7995 | 405.1377 | 1.184 | 1.189 | 1.1865 | 129.2021 | 0.31891
2020-02-14 | 13:14 | 3.816 | 3.841 | 3.8285 | 408.23 | 1.192 | 1.199 | 1.1955 | 130.1822 | 0.31889
2020-02-14 | 13:14 | 3.847 | 3.857 | 3.852 | 410.7358 | 1.201 | 1.205 | 1.203 | 130.9989 | 0.31894
2020-02-14 | 13:14 | 3.86 | 3.877 | 3.8685 | 412.4951 | 1.205 | 1.21 | 1.2075 | 131.4889 | 0.31876
2020-02-14 | 13:14 | 3.88 | 3.887 | 3.8835 | 414.0946 | 1.212 | 1.213 | 1.2125 | 132.0334 | 0.31885
2020-02-14 | 13:14 | 3.88 | 3.885 | 3.8825 | 413.988 | 1.21 | 1.213 | 1.2115 | 131.9245 | 0.31867
2020-02-14 | 13:14 | 3.883 | 3.908 | 3.8955 | 415.3741 | 1.213 | 1.22 | 1.2165 | 132.4689 | 0.31891
2020-02-14 | 13:15 | 3.918 | 3.939 | 3.9285 | 418.8929 | 1.223 | 1.23 | 1.2265 | 133.5579 | 0.31884
2020-02-14 | 13:15 | 3.94 | 3.952 | 3.946 | 420.7589 | 1.231 | 1.234 | 1.2325 | 134.2112 | 0.31897
2020-02-14 | 13:15 | 3.957 | 3.977 | 3.967 | 422.9981 | 1.236 | 1.242 | 1.239 | 134.919 | 0.31896
2020-02-14 | 13:15 | 3.985 | 4.014 | 3.9995 | 426.4636 | 1.245 | 1.254 | 1.2495 | 136.0624 | 0.31905
2020-02-14 | 13:15 | 4.024 | 4.056 | 4.04 | 430.7821 | 1.257 | 1.267 | 1.262 | 137.4236 | 0.31901
2020-02-14 | 13:15 | 4.066 | 4.111 | 4.0885 | 435.9536 | 1.271 | 1.284 | 1.2775 | 139.1114 | 0.3191
2020-02-14 | 13:15 | 4.123 | 4.163 | 4.143 | 441.7649 | 1.289 | 1.302 | 1.2955 | 141.0715 | 0.31934
2020-02-14 | 13:15 | 4.172 | 4.205 | 4.1885 | 446.6165 | 1.305 | 1.315 | 1.31 | 142.6505 | 0.3194
2020-02-14 | 13:15 | 4.213 | 4.243 | 4.228 | 450.8284 | 1.318 | 1.328 | 1.323 | 144.0661 | 0.31956
2020-02-14 | 13:15 | 4.248 | 4.265 | 4.2565 | 453.8673 | 1.33 | 1.336 | 1.333 | 145.155 | 0.31982
2020-02-14 | 13:15 | 4.266 | 4.269 | 4.2675 | 455.0402 | 1.335 | 1.336 | 1.3355 | 145.4273 | 0.31959
2020-02-14 | 13:16 | 4.272 | 4.299 | 4.2855 | 456.9595 | 1.337 | 1.347 | 1.342 | 146.1351 | 0.3198
2020-02-14 | 13:16 | 4.306 | 4.342 | 4.324 | 461.0648 | 1.349 | 1.36 | 1.3545 | 147.4962 | 0.3199
2020-02-14 | 13:16 | 4.351 | 4.392 | 4.3715 | 466.1296 | 1.362 | 1.376 | 1.369 | 149.0752 | 0.31981
2020-02-14 | 13:16 | 4.399 | 4.417 | 4.408 | 470.0216 | 1.378 | 1.384 | 1.381 | 150.3819 | 0.31995
2020-02-14 | 13:16 | 4.423 | 4.462 | 4.4425 | 473.7003 | 1.385 | 1.399 | 1.392 | 151.5798 | 0.31999
2020-02-14 | 13:16 | 4.475 | 4.518 | 4.4965 | 479.4583 | 1.402 | 1.416 | 1.409 | 153.4309 | 0.32001
2020-02-14 | 13:16 | 4.535 | 4.623 | 4.579 | 488.2552 | 1.422 | 1.45 | 1.436 | 156.3711 | 0.32027
2020-02-14 | 13:16 | 4.645 | 4.745 | 4.695 | 500.6242 | 1.458 | 1.489 | 1.4735 | 160.4546 | 0.32051
2020-02-14 | 13:16 | 4.765 | 4.798 | 4.7815 | 509.8476 | 1.497 | 1.507 | 1.502 | 163.558 | 0.3208
2020-02-14 | 13:16 | 4.768 | 4.792 | 4.78 | 509.6877 | 1.497 | 1.505 | 1.501 | 163.4491 | 0.32068
2020-02-14 | 13:16 | 4.769 | 4.83 | 4.7995 | 511.767 | 1.497 | 1.519 | 1.508 | 164.2114 | 0.32087
2020-02-14 | 13:16 | 4.862 | 5.034 | 4.948 | 527.6014 | 1.528 | 1.585 | 1.5565 | 169.4927 | 0.32125
2020-02-14 | 13:16 | 5.08 | 5.205 | 5.1425 | 548.3408 | 1.599 | 1.64 | 1.6195 | 176.353 | 0.32161
2020-02-14 | 13:17 | 5.224 | 5.307 | 5.2655 | 561.4562 | 1.645 | 1.673 | 1.659 | 180.6543 | 0.32176
2020-02-14 | 13:17 | 5.334 | 5.467 | 5.4005 | 575.8511 | 1.682 | 1.724 | 1.703 | 185.4456 | 0.32204
2020-02-14 | 13:17 | 5.499 | 5.599 | 5.549 | 591.6856 | 1.735 | 1.766 | 1.7505 | 190.6181 | 0.32216
2020-02-14 | 13:17 | 5.617 | 5.669 | 5.643 | 601.7087 | 1.772 | 1.787 | 1.7795 | 193.776 | 0.32204
2020-02-14 | 13:17 | 5.678 | 5.707 | 5.6925 | 606.9869 | 1.79 | 1.801 | 1.7955 | 195.5183 | 0.32211
2020-02-14 | 13:17 | 5.714 | 5.756 | 5.735 | 611.5186 | 1.804 | 1.817 | 1.8105 | 197.1517 | 0.3224
2020-02-14 | 13:17 | 5.767 | 5.805 | 5.786 | 616.9567 | 1.821 | 1.831 | 1.826 | 198.8395 | 0.32229
2020-02-14 | 13:17 | 5.815 | 5.85 | 5.8325 | 621.9149 | 1.835 | 1.847 | 1.841 | 200.4729 | 0.32235
2020-02-14 | 13:17 | 5.858 | 5.879 | 5.8685 | 625.7536 | 1.849 | 1.857 | 1.853 | 201.7797 | 0.32246
2020-02-14 | 13:17 | 5.885 | 5.903 | 5.894 | 628.4726 | 1.857 | 1.865 | 1.861 | 202.6508 | 0.32245
2020-02-14 | 13:17 | 5.91 | 5.93 | 5.92 | 631.245 | 1.865 | 1.872 | 1.8685 | 203.4675 | 0.32233
2020-02-14 | 13:18 | 5.936 | 5.954 | 5.945 | 633.9107 | 1.873 | 1.88 | 1.8765 | 204.3387 | 0.32235
2020-02-14 | 13:18 | 5.96 | 5.983 | 5.9715 | 636.7364 | 1.882 | 1.888 | 1.885 | 205.2643 | 0.32237
2020-02-14 | 13:18 | 5.987 | 6.012 | 5.9995 | 639.722 | 1.891 | 1.9 | 1.8955 | 206.4076 | 0.32265
2020-02-14 | 13:18 | 6.018 | 6.038 | 6.028 | 642.761 | 1.9 | 1.906 | 1.903 | 207.2243 | 0.3224
2020-02-14 | 13:18 | 6.042 | 6.059 | 6.0505 | 645.1601 | 1.909 | 1.914 | 1.9115 | 208.1499 | 0.32263
2020-02-14 | 13:18 | 6.064 | 6.082 | 6.073 | 647.5593 | 1.914 | 1.92 | 1.917 | 208.7488 | 0.32236
2020-02-14 | 13:18 | 6.088 | 6.099 | 6.0935 | 649.7452 | 1.922 | 1.928 | 1.925 | 209.62 | 0.32262
2020-02-14 | 13:18 | 6.097 | 6.1 | 6.0985 | 650.2783 | 1.925 | 1.927 | 1.926 | 209.7289 | 0.32252
2020-02-14 | 13:18 | 6.096 | 6.099 | 6.0975 | 650.1717 | 1.925 | 1.926 | 1.9255 | 209.6744 | 0.32249
2020-02-14 | 13:18 | 6.097 | 6.101 | 6.099 | 650.3316 | 1.924 | 1.927 | 1.9255 | 209.6744 | 0.32241
2020-02-14 | 13:18 | 6.099 | 6.102 | 6.1005 | 650.4916 | 1.926 | 1.927 | 1.9265 | 209.7833 | 0.3225
2020-02-14 | 13:18 | 6.108 | 6.136 | 6.122 | 652.7841 | 1.929 | 1.937 | 1.933 | 210.4911 | 0.32245
2020-02-14 | 13:18 | 6.145 | 6.176 | 6.1605 | 656.8893 | 1.941 | 1.951 | 1.946 | 211.9068 | 0.32259
2020-02-14 | 13:19 | 6.185 | 6.213 | 6.199 | 660.9946 | 1.953 | 1.963 | 1.958 | 213.2135 | 0.32256
2020-02-14 | 13:19 | 6.218 | 6.234 | 6.226 | 663.8735 | 1.964 | 1.97 | 1.967 | 214.1935 | 0.32264
2020-02-14 | 13:19 | 6.237 | 6.251 | 6.244 | 665.7929 | 1.97 | 1.975 | 1.9725 | 214.7924 | 0.32261
2020-02-14 | 13:19 | 6.255 | 6.263 | 6.259 | 667.3923 | 1.976 | 1.979 | 1.9775 | 215.3369 | 0.32265
2020-02-14 | 13:19 | 6.265 | 6.27 | 6.2675 | 668.2987 | 1.979 | 1.982 | 1.9805 | 215.6636 | 0.32271
2020-02-14 | 13:19 | 6.272 | 6.283 | 6.2775 | 669.3649 | 1.982 | 1.985 | 1.9835 | 215.9903 | 0.32268
2020-02-14 | 13:19 | 6.286 | 6.293 | 6.2895 | 670.6445 | 1.986 | 1.988 | 1.987 | 216.3714 | 0.32263
2020-02-14 | 13:19 | 6.297 | 6.309 | 6.303 | 672.084 | 1.989 | 1.993 | 1.991 | 216.807 | 0.32259
2020-02-14 | 13:19 | 6.312 | 6.327 | 6.3195 | 673.8434 | 1.995 | 1.999 | 1.997 | 217.4603 | 0.32272
2020-02-14 | 13:19 | 6.331 | 6.341 | 6.336 | 675.6028 | 2.001 | 2.004 | 2.0025 | 218.0592 | 0.32276
2020-02-14 | 13:19 | 6.341 | 6.345 | 6.343 | 676.3492 | 2.004 | 2.005 | 2.0045 | 218.277 | 0.32273
2020-02-14 | 13:20 | 6.337 | 6.341 | 6.339 | 675.9226 | 2.001 | 2.005 | 2.003 | 218.1137 | 0.32269
2020-02-14 | 13:20 | 6.333 | 6.336 | 6.3345 | 675.4428 | 1.999 | 2.001 | 2 | 217.787 | 0.32244
2020-02-14 | 13:20 | 6.333 | 6.338 | 6.3355 | 675.5494 | 2 | 2.002 | 2.001 | 217.8959 | 0.32255
2020-02-14 | 13:20 | 6.334 | 6.338 | 6.336 | 675.6028 | 1.999 | 2.002 | 2.0005 | 217.8415 | 0.32244
2020-02-14 | 13:20 | 6.327 | 6.333 | 6.33 | 674.963 | 1.998 | 1.999 | 1.9985 | 217.6237 | 0.32242
2020-02-14 | 13:20 | 6.325 | 6.327 | 6.326 | 674.5365 | 1.996 | 1.996 | 1.996 | 217.3514 | 0.32222
2020-02-14 | 13:20 | 6.321 | 6.323 | 6.322 | 674.1099 | 1.996 | 1.997 | 1.9965 | 217.4059 | 0.32251
2020-02-14 | 13:20 | 6.317 | 6.321 | 6.319 | 673.7901 | 1.993 | 1.995 | 1.994 | 217.1336 | 0.32226
2020-02-14 | 13:20 | 6.317 | 6.319 | 6.318 | 673.6834 | 1.993 | 1.995 | 1.994 | 217.1336 | 0.32231
2020-02-14 | 13:20 | 6.293 | 6.317 | 6.305 | 672.2973 | 1.984 | 1.994 | 1.989 | 216.5892 | 0.32216
2020-02-14 | 13:20 | 6.235 | 6.28 | 6.2575 | 667.2324 | 1.965 | 1.981 | 1.973 | 214.8469 | 0.322
2020-02-14 | 13:20 | 6.218 | 6.226 | 6.222 | 663.447 | 1.961 | 1.964 | 1.9625 | 213.7035 | 0.32211
2020-02-14 | 13:20 | 6.234 | 6.247 | 6.2405 | 665.4197 | 1.966 | 1.971 | 1.9685 | 214.3569 | 0.32214
2020-02-14 | 13:21 | 6.242 | 6.249 | 6.2455 | 665.9528 | 1.969 | 1.972 | 1.9705 | 214.5746 | 0.32221
2020-02-14 | 13:21 | 6.221 | 6.237 | 6.229 | 664.1934 | 1.962 | 1.966 | 1.964 | 213.8668 | 0.32199
2020-02-14 | 13:21 | 6.187 | 6.217 | 6.202 | 661.3144 | 1.951 | 1.961 | 1.956 | 212.9957 | 0.32208
2020-02-14 | 13:21 | 6.116 | 6.177 | 6.1465 | 655.3965 | 1.925 | 1.947 | 1.936 | 210.8178 | 0.32166
2020-02-14 | 13:21 | 6.018 | 6.095 | 6.0565 | 645.7999 | 1.894 | 1.919 | 1.9065 | 207.6055 | 0.32147
2020-02-14 | 13:21 | 5.921 | 6.003 | 5.962 | 635.7234 | 1.861 | 1.887 | 1.874 | 204.0664 | 0.321
2020-02-14 | 13:21 | 5.824 | 5.898 | 5.861 | 624.9539 | 1.83 | 1.854 | 1.842 | 200.5818 | 0.32095
2020-02-14 | 13:21 | 5.631 | 5.802 | 5.7165 | 609.546 | 1.765 | 1.821 | 1.793 | 195.246 | 0.32031
2020-02-14 | 13:21 | 5.413 | 5.574 | 5.4935 | 585.7676 | 1.694 | 1.746 | 1.72 | 187.2968 | 0.31975
2020-02-14 | 13:21 | 5.346 | 5.391 | 5.3685 | 572.439 | 1.672 | 1.687 | 1.6795 | 182.8866 | 0.31949
2020-02-14 | 13:21 | 5.25 | 5.341 | 5.2955 | 564.6551 | 1.64 | 1.671 | 1.6555 | 180.2732 | 0.31926
2020-02-14 | 13:22 | 5.039 | 5.206 | 5.1225 | 546.2082 | 1.571 | 1.625 | 1.598 | 174.0118 | 0.31858
2020-02-14 | 13:22 | 4.898 | 5.002 | 4.95 | 527.8147 | 1.528 | 1.561 | 1.5445 | 168.186 | 0.31865
2020-02-14 | 13:22 | 4.847 | 4.887 | 4.867 | 518.9644 | 1.511 | 1.524 | 1.5175 | 165.2459 | 0.31841
2020-02-14 | 13:22 | 4.727 | 4.834 | 4.7805 | 509.741 | 1.471 | 1.506 | 1.4885 | 162.088 | 0.31798
2020-02-14 | 13:22 | 4.607 | 4.697 | 4.652 | 496.0391 | 1.435 | 1.462 | 1.4485 | 157.7322 | 0.31798
2020-02-14 | 13:22 | 4.526 | 4.592 | 4.559 | 486.1226 | 1.408 | 1.429 | 1.4185 | 154.4654 | 0.31775
2020-02-14 | 13:22 | 4.412 | 4.507 | 4.4595 | 475.513 | 1.373 | 1.401 | 1.387 | 151.0353 | 0.31763
2020-02-14 | 13:22 | 4.33 | 4.391 | 4.3605 | 464.9567 | 1.346 | 1.365 | 1.3555 | 147.6051 | 0.31746
2020-02-14 | 13:22 | 4.266 | 4.319 | 4.2925 | 457.7059 | 1.325 | 1.343 | 1.334 | 145.2639 | 0.31737
2020-02-14 | 13:22 | 4.221 | 4.252 | 4.2365 | 451.7347 | 1.311 | 1.322 | 1.3165 | 143.3583 | 0.31735
2020-02-14 | 13:22 | 4.203 | 4.216 | 4.2095 | 448.8557 | 1.306 | 1.31 | 1.308 | 142.4327 | 0.31732
2020-02-14 | 13:22 | 4.167 | 4.195 | 4.181 | 445.8168 | 1.295 | 1.303 | 1.299 | 141.4527 | 0.31729
2020-02-14 | 13:22 | 4.109 | 4.163 | 4.136 | 441.0185 | 1.275 | 1.294 | 1.2845 | 139.8737 | 0.31716
2020-02-14 | 13:23 | 4.036 | 4.09 | 4.063 | 433.2345 | 1.253 | 1.27 | 1.2615 | 137.3692 | 0.31708
2020-02-14 | 13:23 | 3.97 | 4.023 | 3.9965 | 426.1437 | 1.232 | 1.249 | 1.2405 | 135.0824 | 0.31699
2020-02-14 | 13:23 | 3.887 | 3.955 | 3.921 | 418.0932 | 1.207 | 1.228 | 1.2175 | 132.5778 | 0.3171
2020-02-14 | 13:23 | 3.803 | 3.868 | 3.8355 | 408.9764 | 1.181 | 1.201 | 1.191 | 129.6922 | 0.31711
2020-02-14 | 13:23 | 3.773 | 3.793 | 3.783 | 403.3784 | 1.171 | 1.178 | 1.1745 | 127.8954 | 0.31706
2020-02-14 | 13:23 | 3.773 | 3.774 | 3.7735 | 402.3654 | 1.172 | 1.174 | 1.173 | 127.7321 | 0.31745
2020-02-14 | 13:23 | 3.763 | 3.774 | 3.7685 | 401.8322 | 1.17 | 1.173 | 1.1715 | 127.5687 | 0.31747
2020-02-14 | 13:23 | 3.741 | 3.76 | 3.7505 | 399.9129 | 1.162 | 1.168 | 1.165 | 126.8609 | 0.31722
2020-02-14 | 13:23 | 3.712 | 3.735 | 3.7235 | 397.0339 | 1.155 | 1.161 | 1.158 | 126.0987 | 0.3176
2020-02-14 | 13:23 | 3.71 | 3.733 | 3.7215 | 396.8207 | 1.155 | 1.162 | 1.1585 | 126.1531 | 0.31791
2020-02-14 | 13:23 | 3.748 | 3.822 | 3.785 | 403.5916 | 1.166 | 1.19 | 1.178 | 128.2765 | 0.31784
2020-02-14 | 13:24 | 3.842 | 3.904 | 3.873 | 412.975 | 1.195 | 1.215 | 1.205 | 131.2167 | 0.31774
2020-02-14 | 13:24 | 3.917 | 3.975 | 3.946 | 420.7589 | 1.218 | 1.237 | 1.2275 | 133.6668 | 0.31768
2020-02-14 | 13:24 | 3.988 | 4.002 | 3.995 | 425.9837 | 1.24 | 1.244 | 1.242 | 135.2457 | 0.31749
2020-02-14 | 13:24 | 3.864 | 3.971 | 3.9175 | 417.72 | 1.201 | 1.234 | 1.2175 | 132.5778 | 0.31738
2020-02-14 | 13:24 | 3.702 | 3.83 | 3.766 | 401.5657 | 1.149 | 1.19 | 1.1695 | 127.351 | 0.31714
2020-02-14 | 13:24 | 3.576 | 3.674 | 3.625 | 386.5309 | 1.111 | 1.139 | 1.125 | 122.5052 | 0.31694
2020-02-14 | 13:24 | 3.476 | 3.554 | 3.515 | 374.8017 | 1.079 | 1.102 | 1.0905 | 118.7484 | 0.31683
2020-02-14 | 13:24 | 3.407 | 3.461 | 3.434 | 366.1648 | 1.055 | 1.074 | 1.0645 | 115.9171 | 0.31657
2020-02-14 | 13:24 | 3.35 | 3.394 | 3.372 | 359.5537 | 1.04 | 1.051 | 1.0455 | 113.8482 | 0.31664
2020-02-14 | 13:24 | 3.305 | 3.34 | 3.3225 | 354.2756 | 1.026 | 1.035 | 1.0305 | 112.2148 | 0.31674
2020-02-14 | 13:24 | 3.269 | 3.297 | 3.283 | 350.0637 | 1.014 | 1.023 | 1.0185 | 110.908 | 0.31682
2020-02-14 | 13:24 | 3.235 | 3.262 | 3.2485 | 346.385 | 1.003 | 1.012 | 1.0075 | 109.7102 | 0.31673
2020-02-14 | 13:24 | 3.21 | 3.228 | 3.219 | 343.2395 | 0.996 | 1.002 | 0.999 | 108.7846 | 0.31694
2020-02-14 | 13:25 | 3.195 | 3.205 | 3.2 | 341.2135 | 0.993 | 0.996 | 0.9945 | 108.2946 | 0.31738
2020-02-14 | 13:25 | 3.19 | 3.193 | 3.1915 | 340.3072 | 0.991 | 0.992 | 0.9915 | 107.9679 | 0.31727
2020-02-14 | 13:25 | 3.191 | 3.194 | 3.1925 | 340.4138 | 0.992 | 0.993 | 0.9925 | 108.0768 | 0.31749
2020-02-14 | 13:25 | 3.192 | 3.197 | 3.1945 | 340.6271 | 0.993 | 0.996 | 0.9945 | 108.2946 | 0.31793
2020-02-14 | 13:25 | 3.199 | 3.204 | 3.2015 | 341.3735 | 0.996 | 0.999 | 0.9975 | 108.6213 | 0.31819
2020-02-14 | 13:25 | 3.205 | 3.213 | 3.209 | 342.1732 | 0.999 | 1.003 | 1.001 | 109.0024 | 0.31856
2020-02-14 | 13:25 | 3.217 | 3.246 | 3.2315 | 344.5723 | 1.005 | 1.013 | 1.009 | 109.8735 | 0.31887
2020-02-14 | 13:25 | 3.255 | 3.318 | 3.2865 | 350.4369 | 1.017 | 1.039 | 1.028 | 111.9425 | 0.31944
2020-02-14 | 13:25 | 3.34 | 3.438 | 3.389 | 361.3664 | 1.046 | 1.076 | 1.061 | 115.536 | 0.31972
2020-02-14 | 13:25 | 3.463 | 3.574 | 3.5185 | 375.1749 | 1.085 | 1.122 | 1.1035 | 120.164 | 0.32029
2020-02-14 | 13:25 | 3.607 | 3.767 | 3.687 | 393.1419 | 1.132 | 1.183 | 1.1575 | 126.0442 | 0.32061
2020-02-14 | 13:26 | 3.812 | 3.979 | 3.8955 | 415.3741 | 1.199 | 1.251 | 1.225 | 133.3945 | 0.32114
2020-02-14 | 13:26 | 4.016 | 4.112 | 4.064 | 433.3412 | 1.262 | 1.294 | 1.278 | 139.1659 | 0.32115
2020-02-14 | 13:26 | 4.127 | 4.19 | 4.1585 | 443.4176 | 1.299 | 1.319 | 1.309 | 142.5416 | 0.32146
2020-02-14 | 13:26 | 4.211 | 4.313 | 4.262 | 454.4537 | 1.327 | 1.359 | 1.343 | 146.244 | 0.3218
2020-02-14 | 13:26 | 4.321 | 4.333 | 4.327 | 461.3846 | 1.361 | 1.365 | 1.363 | 148.4218 | 0.32169
2020-02-14 | 13:26 | 4.319 | 4.36 | 4.3395 | 462.7175 | 1.362 | 1.375 | 1.3685 | 149.0208 | 0.32206
2020-02-14 | 13:26 | 4.37 | 4.386 | 4.378 | 466.8227 | 1.376 | 1.381 | 1.3785 | 150.1097 | 0.32156
2020-02-14 | 13:26 | 4.333 | 4.384 | 4.3585 | 464.7435 | 1.362 | 1.379 | 1.3705 | 149.2385 | 0.32112
2020-02-14 | 13:26 | 4.224 | 4.307 | 4.2655 | 454.827 | 1.325 | 1.353 | 1.339 | 145.8084 | 0.32058
2020-02-14 | 13:26 | 4.11 | 4.206 | 4.158 | 443.3643 | 1.288 | 1.32 | 1.304 | 141.9971 | 0.32027
2020-02-14 | 13:26 | 4.002 | 4.081 | 4.0415 | 430.942 | 1.254 | 1.28 | 1.267 | 137.9681 | 0.32015
2020-02-14 | 13:26 | 3.951 | 3.988 | 3.9695 | 423.2647 | 1.237 | 1.249 | 1.243 | 135.3546 | 0.31979
2020-02-14 | 13:26 | 3.928 | 3.942 | 3.935 | 419.586 | 1.23 | 1.235 | 1.2325 | 134.2112 | 0.31987
2020-02-14 | 13:27 | 3.928 | 3.948 | 3.938 | 419.9059 | 1.231 | 1.236 | 1.2335 | 134.3201 | 0.31988
2020-02-14 | 13:27 | 3.958 | 4.002 | 3.98 | 424.3843 | 1.24 | 1.254 | 1.247 | 135.7902 | 0.31997
2020-02-14 | 13:27 | 4.014 | 4.054 | 4.034 | 430.1423 | 1.258 | 1.27 | 1.264 | 137.6414 | 0.31999
2020-02-14 | 13:27 | 4.042 | 4.059 | 4.0505 | 431.9017 | 1.267 | 1.273 | 1.27 | 138.2947 | 0.3202
2020-02-14 | 13:27 | 3.999 | 4.033 | 4.016 | 428.223 | 1.254 | 1.264 | 1.259 | 137.0969 | 0.32015
2020-02-14 | 13:27 | 3.999 | 4.027 | 4.013 | 427.9031 | 1.254 | 1.263 | 1.2585 | 137.0425 | 0.32027
2020-02-14 | 13:27 | 4.034 | 4.058 | 4.046 | 431.4218 | 1.265 | 1.274 | 1.2695 | 138.2403 | 0.32043
2020-02-14 | 13:27 | 4.067 | 4.099 | 4.083 | 435.3671 | 1.274 | 1.285 | 1.2795 | 139.3292 | 0.32003
2020-02-14 | 13:27 | 4.104 | 4.126 | 4.115 | 438.7793 | 1.286 | 1.295 | 1.2905 | 140.5271 | 0.32027
2020-02-14 | 13:27 | 4.138 | 4.193 | 4.1655 | 444.164 | 1.298 | 1.316 | 1.307 | 142.3238 | 0.32043
2020-02-14 | 13:27 | 4.179 | 4.202 | 4.1905 | 446.8298 | 1.309 | 1.319 | 1.314 | 143.0861 | 0.32023
2020-02-14 | 13:28 | 4.107 | 4.166 | 4.1365 | 441.0718 | 1.287 | 1.306 | 1.2965 | 141.1804 | 0.32008
2020-02-14 | 13:28 | 4.087 | 4.101 | 4.094 | 436.54 | 1.28 | 1.286 | 1.283 | 139.7104 | 0.32004
2020-02-14 | 13:28 | 4.023 | 4.08 | 4.0515 | 432.0083 | 1.26 | 1.278 | 1.269 | 138.1859 | 0.31987
2020-02-14 | 13:28 | 3.926 | 4.003 | 3.9645 | 422.7316 | 1.229 | 1.253 | 1.241 | 135.1368 | 0.31968
2020-02-14 | 13:28 | 3.867 | 3.911 | 3.889 | 414.681 | 1.209 | 1.223 | 1.216 | 132.4145 | 0.31932
2020-02-14 | 13:28 | 3.844 | 3.86 | 3.852 | 410.7358 | 1.201 | 1.207 | 1.204 | 131.1078 | 0.3192
2020-02-14 | 13:28 | 3.822 | 3.841 | 3.8315 | 408.5499 | 1.194 | 1.201 | 1.1975 | 130.4 | 0.31918
2020-02-14 | 13:28 | 3.753 | 3.812 | 3.7825 | 403.325 | 1.172 | 1.191 | 1.1815 | 128.6577 | 0.31899
2020-02-14 | 13:28 | 3.672 | 3.737 | 3.7045 | 395.008 | 1.147 | 1.166 | 1.1565 | 125.9353 | 0.31882
2020-02-14 | 13:28 | 3.622 | 3.662 | 3.642 | 388.3436 | 1.13 | 1.143 | 1.1365 | 123.7575 | 0.31868
2020-02-14 | 13:28 | 3.599 | 3.616 | 3.6075 | 384.6649 | 1.122 | 1.128 | 1.125 | 122.5052 | 0.31847
2020-02-14 | 13:28 | 3.581 | 3.596 | 3.5885 | 382.639 | 1.118 | 1.122 | 1.12 | 121.9607 | 0.31874
2020-02-14 | 13:28 | 3.568 | 3.579 | 3.5735 | 381.0395 | 1.113 | 1.116 | 1.1145 | 121.3618 | 0.3185
2020-02-14 | 13:29 | 3.563 | 3.565 | 3.564 | 380.0266 | 1.111 | 1.112 | 1.1115 | 121.0351 | 0.31849
2020-02-14 | 13:29 | 3.564 | 3.572 | 3.568 | 380.4531 | 1.112 | 1.114 | 1.113 | 121.1985 | 0.31856
2020-02-14 | 13:29 | 3.575 | 3.596 | 3.5855 | 382.3191 | 1.115 | 1.122 | 1.1185 | 121.7974 | 0.31858
2020-02-14 | 13:29 | 3.602 | 3.622 | 3.612 | 385.1448 | 1.125 | 1.131 | 1.128 | 122.8319 | 0.31892
2020-02-14 | 13:29 | 3.625 | 3.63 | 3.6275 | 386.7975 | 1.132 | 1.135 | 1.1335 | 123.4308 | 0.31911
2020-02-14 | 13:29 | 3.612 | 3.629 | 3.6205 | 386.0511 | 1.128 | 1.132 | 1.13 | 123.0497 | 0.31874
2020-02-14 | 13:29 | 3.572 | 3.606 | 3.589 | 382.6923 | 1.114 | 1.125 | 1.1195 | 121.9063 | 0.31855
2020-02-14 | 13:29 | 3.529 | 3.563 | 3.546 | 378.1072 | 1.102 | 1.11 | 1.106 | 120.4362 | 0.31852
2020-02-14 | 13:29 | 3.514 | 3.524 | 3.519 | 375.2282 | 1.096 | 1.099 | 1.0975 | 119.5106 | 0.3185
2020-02-14 | 13:29 | 3.514 | 3.532 | 3.523 | 375.6548 | 1.097 | 1.103 | 1.1 | 119.7829 | 0.31886
2020-02-14 | 13:29 | 3.542 | 3.602 | 3.572 | 380.8796 | 1.107 | 1.126 | 1.1165 | 121.5796 | 0.31921
2020-02-14 | 13:30 | 3.62 | 3.681 | 3.6505 | 389.25 | 1.132 | 1.151 | 1.1415 | 124.3019 | 0.31934
2020-02-14 | 13:30 | 3.694 | 3.743 | 3.7185 | 396.5008 | 1.155 | 1.172 | 1.1635 | 126.6976 | 0.31954
2020-02-14 | 13:30 | 3.756 | 3.807 | 3.7815 | 403.2184 | 1.177 | 1.194 | 1.1855 | 129.0932 | 0.32016
2020-02-14 | 13:30 | 3.819 | 3.87 | 3.8445 | 409.936 | 1.198 | 1.215 | 1.2065 | 131.38 | 0.32049
2020-02-14 | 13:30 | 3.882 | 3.935 | 3.9085 | 416.7603 | 1.217 | 1.235 | 1.226 | 133.5034 | 0.32034
2020-02-14 | 13:30 | 3.948 | 4.005 | 3.9765 | 424.0111 | 1.238 | 1.258 | 1.248 | 135.8991 | 0.32051
2020-02-14 | 13:30 | 4.02 | 4.076 | 4.048 | 431.6351 | 1.263 | 1.282 | 1.2725 | 138.567 | 0.32103
2020-02-14 | 13:30 | 4.09 | 4.138 | 4.114 | 438.6726 | 1.286 | 1.301 | 1.2935 | 140.8537 | 0.32109
2020-02-14 | 13:30 | 4.146 | 4.166 | 4.156 | 443.1511 | 1.304 | 1.309 | 1.3065 | 142.2694 | 0.32104
2020-02-14 | 13:30 | 4.168 | 4.177 | 4.1725 | 444.9104 | 1.31 | 1.312 | 1.311 | 142.7594 | 0.32087
2020-02-14 | 13:30 | 4.182 | 4.211 | 4.1965 | 447.4695 | 1.313 | 1.324 | 1.3185 | 143.5761 | 0.32086
2020-02-14 | 13:30 | 4.222 | 4.271 | 4.2465 | 452.801 | 1.328 | 1.344 | 1.336 | 145.4817 | 0.32129
2020-02-14 | 13:30 | 4.283 | 4.329 | 4.306 | 459.1454 | 1.347 | 1.361 | 1.354 | 147.4418 | 0.32112
2020-02-14 | 13:31 | 4.338 | 4.374 | 4.356 | 464.4769 | 1.364 | 1.377 | 1.3705 | 149.2385 | 0.3213
2020-02-14 | 13:31 | 4.381 | 4.416 | 4.3985 | 469.0086 | 1.379 | 1.389 | 1.384 | 150.7086 | 0.32133
2020-02-14 | 13:31 | 4.421 | 4.442 | 4.4315 | 472.5274 | 1.392 | 1.398 | 1.395 | 151.9064 | 0.32148
2020-02-14 | 13:31 | 4.433 | 4.445 | 4.439 | 473.3271 | 1.395 | 1.401 | 1.398 | 152.2331 | 0.32162
2020-02-14 | 13:31 | 4.393 | 4.425 | 4.409 | 470.1282 | 1.382 | 1.393 | 1.3875 | 151.0897 | 0.32138
2020-02-14 | 13:31 | 4.38 | 4.386 | 4.383 | 467.3559 | 1.377 | 1.379 | 1.378 | 150.0552 | 0.32107
2020-02-14 | 13:31 | 4.389 | 4.435 | 4.412 | 470.4481 | 1.38 | 1.396 | 1.388 | 151.1442 | 0.32128
2020-02-14 | 13:31 | 4.451 | 4.504 | 4.4775 | 477.4323 | 1.401 | 1.418 | 1.4095 | 153.4854 | 0.32148
2020-02-14 | 13:31 | 4.515 | 4.555 | 4.535 | 483.5635 | 1.423 | 1.435 | 1.429 | 155.6088 | 0.3218
2020-02-14 | 13:31 | 4.564 | 4.598 | 4.581 | 488.4685 | 1.438 | 1.448 | 1.443 | 157.1333 | 0.32169
2020-02-14 | 13:31 | 4.606 | 4.645 | 4.6255 | 493.2135 | 1.451 | 1.466 | 1.4585 | 158.8212 | 0.32201
2020-02-14 | 13:32 | 4.655 | 4.695 | 4.675 | 498.4916 | 1.467 | 1.48 | 1.4735 | 160.4546 | 0.32188
2020-02-14 | 13:32 | 4.705 | 4.745 | 4.725 | 503.8231 | 1.484 | 1.495 | 1.4895 | 162.1969 | 0.32193
2020-02-14 | 13:32 | 4.753 | 4.784 | 4.7685 | 508.4614 | 1.498 | 1.508 | 1.503 | 163.6669 | 0.32189
2020-02-14 | 13:32 | 4.789 | 4.807 | 4.798 | 511.607 | 1.509 | 1.516 | 1.5125 | 164.7014 | 0.32193
2020-02-14 | 13:32 | 4.809 | 4.824 | 4.8165 | 513.5797 | 1.516 | 1.52 | 1.518 | 165.3003 | 0.32186
2020-02-14 | 13:32 | 4.829 | 4.838 | 4.8335 | 515.3923 | 1.521 | 1.525 | 1.523 | 165.8448 | 0.32178
2020-02-14 | 13:32 | 4.838 | 4.84 | 4.839 | 515.9788 | 1.524 | 1.526 | 1.525 | 166.0626 | 0.32184
2020-02-14 | 13:32 | 4.834 | 4.838 | 4.836 | 515.6589 | 1.523 | 1.524 | 1.5235 | 165.8993 | 0.32172
2020-02-14 | 13:32 | 4.833 | 4.836 | 4.8345 | 515.499 | 1.523 | 1.524 | 1.5235 | 165.8993 | 0.32182
2020-02-14 | 13:32 | 4.837 | 4.841 | 4.839 | 515.9788 | 1.524 | 1.526 | 1.525 | 166.0626 | 0.32184
2020-02-14 | 13:32 | 4.843 | 4.85 | 4.8465 | 516.7785 | 1.527 | 1.53 | 1.5285 | 166.4437 | 0.32208
2020-02-14 | 13:32 | 4.853 | 4.868 | 4.8605 | 518.2713 | 1.529 | 1.534 | 1.5315 | 166.7704 | 0.32178
2020-02-14 | 13:32 | 4.873 | 4.893 | 4.883 | 520.6705 | 1.536 | 1.542 | 1.539 | 167.5871 | 0.32187
2020-02-14 | 13:33 | 4.899 | 4.916 | 4.9075 | 523.2829 | 1.544 | 1.549 | 1.5465 | 168.4038 | 0.32182
2020-02-14 | 13:33 | 4.921 | 4.936 | 4.9285 | 525.5221 | 1.551 | 1.557 | 1.554 | 169.2205 | 0.322
2020-02-14 | 13:33 | 4.941 | 4.957 | 4.949 | 527.708 | 1.559 | 1.564 | 1.5615 | 170.0372 | 0.32222
2020-02-14 | 13:33 | 4.962 | 4.988 | 4.975 | 530.4804 | 1.565 | 1.575 | 1.57 | 170.9628 | 0.32228
2020-02-14 | 13:33 | 4.996 | 5.029 | 5.0125 | 534.479 | 1.576 | 1.587 | 1.5815 | 172.2151 | 0.32221
2020-02-14 | 13:33 | 5.037 | 5.064 | 5.0505 | 538.5309 | 1.59 | 1.599 | 1.5945 | 173.6307 | 0.32242
2020-02-14 | 13:33 | 5.069 | 5.087 | 5.078 | 541.4632 | 1.6 | 1.608 | 1.604 | 174.6652 | 0.32258
2020-02-14 | 13:33 | 5.095 | 5.116 | 5.1055 | 544.3955 | 1.609 | 1.616 | 1.6125 | 175.5908 | 0.32254
2020-02-14 | 13:33 | 5.12 | 5.136 | 5.128 | 546.7947 | 1.617 | 1.622 | 1.6195 | 176.353 | 0.32252
2020-02-14 | 13:33 | 5.137 | 5.142 | 5.1395 | 548.0209 | 1.622 | 1.625 | 1.6235 | 176.7886 | 0.32259
2020-02-14 | 13:33 | 5.145 | 5.147 | 5.146 | 548.714 | 1.624 | 1.626 | 1.625 | 176.9519 | 0.32248
2020-02-14 | 13:34 | 5.142 | 5.147 | 5.1445 | 548.554 | 1.624 | 1.626 | 1.625 | 176.9519 | 0.32258
2020-02-14 | 13:34 | 5.126 | 5.138 | 5.132 | 547.2212 | 1.619 | 1.623 | 1.621 | 176.5164 | 0.32257
2020-02-14 | 13:34 | 5.119 | 5.124 | 5.1215 | 546.1016 | 1.616 | 1.618 | 1.617 | 176.0808 | 0.32243
2020-02-14 | 13:34 | 5.109 | 5.12 | 5.1145 | 545.3552 | 1.613 | 1.616 | 1.6145 | 175.8086 | 0.32237
2020-02-14 | 13:34 | 5.108 | 5.117 | 5.1125 | 545.1419 | 1.613 | 1.618 | 1.6155 | 175.9175 | 0.3227
2020-02-14 | 13:34 | 5.121 | 5.133 | 5.127 | 546.688 | 1.618 | 1.622 | 1.62 | 176.4075 | 0.32268
2020-02-14 | 13:34 | 5.137 | 5.155 | 5.146 | 548.714 | 1.623 | 1.63 | 1.6265 | 177.1153 | 0.32278
2020-02-14 | 13:34 | 5.162 | 5.201 | 5.1815 | 552.4993 | 1.632 | 1.645 | 1.6385 | 178.422 | 0.32294
2020-02-14 | 13:34 | 5.211 | 5.234 | 5.2225 | 556.8711 | 1.648 | 1.655 | 1.6515 | 179.8376 | 0.32294
2020-02-14 | 13:34 | 5.238 | 5.249 | 5.2435 | 559.1103 | 1.657 | 1.662 | 1.6595 | 180.7088 | 0.32321
2020-02-14 | 13:34 | 5.252 | 5.259 | 5.2555 | 560.3899 | 1.663 | 1.665 | 1.664 | 181.1988 | 0.32334
2020-02-14 | 13:34 | 5.261 | 5.273 | 5.267 | 561.6161 | 1.666 | 1.669 | 1.6675 | 181.5799 | 0.32332
2020-02-14 | 13:34 | 5.277 | 5.297 | 5.287 | 563.7487 | 1.671 | 1.676 | 1.6735 | 182.2333 | 0.32325
2020-02-14 | 13:35 | 5.301 | 5.319 | 5.31 | 566.2012 | 1.678 | 1.684 | 1.681 | 183.05 | 0.32329
2020-02-14 | 13:35 | 5.32 | 5.324 | 5.322 | 567.4807 | 1.685 | 1.686 | 1.6855 | 183.54 | 0.32343
2020-02-14 | 13:35 | 5.323 | 5.325 | 5.324 | 567.694 | 1.685 | 1.687 | 1.686 | 183.5944 | 0.3234
2020-02-14 | 13:35 | 5.323 | 5.327 | 5.325 | 567.8006 | 1.686 | 1.687 | 1.6865 | 183.6489 | 0.32344
2020-02-14 | 13:35 | 5.323 | 5.328 | 5.3255 | 567.8539 | 1.685 | 1.688 | 1.6865 | 183.6489 | 0.32341
2020-02-14 | 13:35 | 5.333 | 5.349 | 5.341 | 569.5067 | 1.688 | 1.694 | 1.691 | 184.1389 | 0.32333
2020-02-14 | 13:35 | 5.352 | 5.362 | 5.357 | 571.2127 | 1.695 | 1.7 | 1.6975 | 184.8467 | 0.3236
2020-02-14 | 13:35 | 5.365 | 5.375 | 5.37 | 572.5989 | 1.699 | 1.703 | 1.701 | 185.2278 | 0.32349
2020-02-14 | 13:35 | 5.377 | 5.401 | 5.389 | 574.6249 | 1.704 | 1.711 | 1.7075 | 185.9357 | 0.32358
2020-02-14 | 13:35 | 5.408 | 5.433 | 5.4205 | 577.9837 | 1.715 | 1.722 | 1.7185 | 187.1335 | 0.32377
2020-02-14 | 13:35 | 5.442 | 5.48 | 5.461 | 582.3022 | 1.725 | 1.738 | 1.7315 | 188.5491 | 0.3238
2020-02-14 | 13:36 | 5.487 | 5.522 | 5.5045 | 586.9406 | 1.741 | 1.751 | 1.746 | 190.1281 | 0.32393
2020-02-14 | 13:36 | 5.526 | 5.533 | 5.5295 | 589.6063 | 1.754 | 1.756 | 1.755 | 191.1081 | 0.32413
2020-02-14 | 13:36 | 5.53 | 5.534 | 5.532 | 589.8729 | 1.755 | 1.756 | 1.7555 | 191.1625 | 0.32407
2020-02-14 | 13:36 | 5.524 | 5.532 | 5.528 | 589.4463 | 1.751 | 1.755 | 1.753 | 190.8903 | 0.32385
2020-02-14 | 13:36 | 5.516 | 5.521 | 5.5185 | 588.4334 | 1.75 | 1.753 | 1.7515 | 190.727 | 0.32413
2020-02-14 | 13:36 | 5.5 | 5.515 | 5.5075 | 587.2604 | 1.745 | 1.75 | 1.7475 | 190.2914 | 0.32403
2020-02-14 | 13:36 | 5.48 | 5.495 | 5.4875 | 585.1279 | 1.738 | 1.743 | 1.7405 | 189.5291 | 0.32391
2020-02-14 | 13:36 | 5.481 | 5.495 | 5.488 | 585.1812 | 1.738 | 1.743 | 1.7405 | 189.5291 | 0.32388
2020-02-14 | 13:36 | 5.492 | 5.498 | 5.495 | 585.9276 | 1.742 | 1.746 | 1.744 | 189.9103 | 0.32412
2020-02-14 | 13:36 | 5.463 | 5.488 | 5.4755 | 583.8483 | 1.732 | 1.741 | 1.7365 | 189.0936 | 0.32387
2020-02-14 | 13:36 | 5.428 | 5.456 | 5.442 | 580.2762 | 1.721 | 1.729 | 1.725 | 187.8413 | 0.32371
2020-02-14 | 13:36 | 5.402 | 5.424 | 5.413 | 577.184 | 1.712 | 1.719 | 1.7155 | 186.8068 | 0.32365
2020-02-14 | 13:36 | 5.365 | 5.395 | 5.38 | 573.6652 | 1.699 | 1.709 | 1.704 | 185.5545 | 0.32345
2020-02-14 | 13:37 | 5.334 | 5.357 | 5.3455 | 569.9865 | 1.69 | 1.697 | 1.6935 | 184.4111 | 0.32354
2020-02-14 | 13:37 | 5.307 | 5.33 | 5.3185 | 567.1075 | 1.68 | 1.689 | 1.6845 | 183.4311 | 0.32345
2020-02-14 | 13:37 | 5.269 | 5.3 | 5.2845 | 563.4821 | 1.668 | 1.676 | 1.672 | 182.0699 | 0.32312
2020-02-14 | 13:37 | 5.241 | 5.262 | 5.2515 | 559.9634 | 1.658 | 1.666 | 1.662 | 180.981 | 0.3232
2020-02-14 | 13:37 | 5.223 | 5.236 | 5.2295 | 557.6175 | 1.652 | 1.658 | 1.655 | 180.2187 | 0.32319
2020-02-14 | 13:37 | 5.201 | 5.218 | 5.2095 | 555.4849 | 1.646 | 1.652 | 1.649 | 179.5654 | 0.32326
2020-02-14 | 13:37 | 5.199 | 5.207 | 5.203 | 554.7918 | 1.645 | 1.648 | 1.6465 | 179.2932 | 0.32317
2020-02-14 | 13:37 | 5.212 | 5.233 | 5.2225 | 556.8711 | 1.649 | 1.657 | 1.653 | 180.001 | 0.32324
2020-02-14 | 13:37 | 5.236 | 5.243 | 5.2395 | 558.6838 | 1.659 | 1.661 | 1.66 | 180.7632 | 0.32355
2020-02-14 | 13:37 | 5.232 | 5.242 | 5.237 | 558.4172 | 1.657 | 1.662 | 1.6595 | 180.7088 | 0.32361
2020-02-14 | 13:37 | 5.195 | 5.227 | 5.211 | 555.6449 | 1.643 | 1.657 | 1.65 | 179.6743 | 0.32336
2020-02-14 | 13:38 | 5.152 | 5.187 | 5.1695 | 551.2198 | 1.631 | 1.642 | 1.6365 | 178.2042 | 0.32329
2020-02-14 | 13:38 | 5.126 | 5.145 | 5.1355 | 547.5944 | 1.622 | 1.628 | 1.625 | 176.9519 | 0.32314
2020-02-14 | 13:38 | 5.105 | 5.121 | 5.113 | 545.1952 | 1.615 | 1.619 | 1.617 | 176.0808 | 0.32297
2020-02-14 | 13:38 | 5.073 | 5.098 | 5.0855 | 542.2629 | 1.605 | 1.613 | 1.609 | 175.2096 | 0.32311
2020-02-14 | 13:38 | 5.056 | 5.069 | 5.0625 | 539.8104 | 1.599 | 1.602 | 1.6005 | 174.284 | 0.32286
2020-02-14 | 13:38 | 5.026 | 5.052 | 5.039 | 537.3047 | 1.589 | 1.596 | 1.5925 | 173.4129 | 0.32275
2020-02-14 | 13:38 | 4.993 | 5.019 | 5.006 | 533.7859 | 1.579 | 1.586 | 1.5825 | 172.324 | 0.32283
2020-02-14 | 13:38 | 4.971 | 4.987 | 4.979 | 530.9069 | 1.572 | 1.576 | 1.574 | 171.3984 | 0.32284
2020-02-14 | 13:38 | 4.94 | 4.966 | 4.953 | 528.1345 | 1.56 | 1.57 | 1.565 | 170.4183 | 0.32268
2020-02-14 | 13:38 | 4.913 | 4.935 | 4.924 | 525.0423 | 1.552 | 1.56 | 1.556 | 169.4383 | 0.32271
2020-02-14 | 13:38 | 4.886 | 4.907 | 4.8965 | 522.11 | 1.543 | 1.55 | 1.5465 | 168.4038 | 0.32254
2020-02-14 | 13:38 | 4.868 | 4.881 | 4.8745 | 519.7641 | 1.538 | 1.542 | 1.54 | 167.696 | 0.32264
2020-02-14 | 13:38 | 4.85 | 4.864 | 4.857 | 517.8981 | 1.532 | 1.537 | 1.5345 | 167.0971 | 0.32264
2020-02-14 | 13:39 | 4.828 | 4.847 | 4.8375 | 515.8189 | 1.524 | 1.531 | 1.5275 | 166.3348 | 0.32247
2020-02-14 | 13:39 | 4.795 | 4.822 | 4.8085 | 512.7266 | 1.513 | 1.522 | 1.5175 | 165.2459 | 0.32229
2020-02-14 | 13:39 | 4.782 | 4.792 | 4.787 | 510.4341 | 1.509 | 1.513 | 1.511 | 164.5381 | 0.32235
2020-02-14 | 13:39 | 4.781 | 4.786 | 4.7835 | 510.0609 | 1.509 | 1.511 | 1.51 | 164.4292 | 0.32237
2020-02-14 | 13:39 | 4.766 | 4.785 | 4.7755 | 509.2079 | 1.503 | 1.51 | 1.5065 | 164.0481 | 0.32216
2020-02-14 | 13:39 | 4.728 | 4.759 | 4.7435 | 505.7957 | 1.491 | 1.501 | 1.496 | 162.9047 | 0.32208
2020-02-14 | 13:39 | 4.677 | 4.72 | 4.6985 | 500.9974 | 1.473 | 1.489 | 1.481 | 161.2713 | 0.3219
2020-02-14 | 13:39 | 4.615 | 4.665 | 4.64 | 494.7596 | 1.452 | 1.469 | 1.4605 | 159.039 | 0.32145
2020-02-14 | 13:39 | 4.535 | 4.6 | 4.5675 | 487.029 | 1.427 | 1.448 | 1.4375 | 156.5344 | 0.32141
2020-02-14 | 13:39 | 4.456 | 4.518 | 4.487 | 478.4453 | 1.4 | 1.421 | 1.4105 | 153.5943 | 0.32103
2020-02-14 | 13:39 | 4.355 | 4.439 | 4.397 | 468.8487 | 1.366 | 1.394 | 1.38 | 150.273 | 0.32051
2020-02-14 | 13:40 | 4.252 | 4.33 | 4.291 | 457.546 | 1.333 | 1.36 | 1.3465 | 146.6251 | 0.32046
2020-02-14 | 13:40 | 4.186 | 4.236 | 4.211 | 449.0157 | 1.312 | 1.329 | 1.3205 | 143.7939 | 0.32024
2020-02-14 | 13:40 | 4.115 | 4.171 | 4.143 | 441.7649 | 1.29 | 1.307 | 1.2985 | 141.3982 | 0.32008
2020-02-14 | 13:40 | 4.07 | 4.103 | 4.0865 | 435.7403 | 1.274 | 1.285 | 1.2795 | 139.3292 | 0.31975
2020-02-14 | 13:40 | 4.045 | 4.064 | 4.0545 | 432.3282 | 1.266 | 1.274 | 1.27 | 138.2947 | 0.31988
2020-02-14 | 13:40 | 4.01 | 4.041 | 4.0255 | 429.2359 | 1.255 | 1.265 | 1.26 | 137.2058 | 0.31965
2020-02-14 | 13:40 | 3.963 | 3.999 | 3.981 | 424.4909 | 1.241 | 1.253 | 1.247 | 135.7902 | 0.31989
2020-02-14 | 13:40 | 3.927 | 3.954 | 3.9405 | 420.1725 | 1.229 | 1.238 | 1.2335 | 134.3201 | 0.31968
2020-02-14 | 13:40 | 3.868 | 3.918 | 3.893 | 415.1076 | 1.209 | 1.226 | 1.2175 | 132.5778 | 0.31938
2020-02-14 | 13:40 | 3.782 | 3.853 | 3.8175 | 407.0571 | 1.182 | 1.204 | 1.193 | 129.9099 | 0.31914
2020-02-14 | 13:40 | 3.688 | 3.765 | 3.7265 | 397.3538 | 1.151 | 1.176 | 1.1635 | 126.6976 | 0.31885
2020-02-14 | 13:40 | 3.568 | 3.666 | 3.617 | 385.6779 | 1.113 | 1.144 | 1.1285 | 122.8863 | 0.31862
2020-02-14 | 13:40 | 3.466 | 3.546 | 3.506 | 373.8421 | 1.079 | 1.106 | 1.0925 | 118.9662 | 0.31823
2020-02-14 | 13:41 | 3.359 | 3.446 | 3.4025 | 362.8059 | 1.044 | 1.072 | 1.058 | 115.2093 | 0.31755
2020-02-14 | 13:41 | 3.255 | 3.336 | 3.2955 | 351.3966 | 1.012 | 1.037 | 1.0245 | 111.5614 | 0.31748
2020-02-14 | 13:41 | 3.185 | 3.239 | 3.212 | 342.4931 | 0.99 | 1.006 | 0.998 | 108.6757 | 0.31731
2020-02-14 | 13:41 | 3.134 | 3.174 | 3.154 | 336.3086 | 0.974 | 0.985 | 0.9795 | 106.6612 | 0.31715
2020-02-14 | 13:41 | 3.085 | 3.125 | 3.105 | 331.0837 | 0.957 | 0.971 | 0.964 | 104.9733 | 0.31706
2020-02-14 | 13:41 | 3.03 | 3.074 | 3.052 | 325.4324 | 0.94 | 0.953 | 0.9465 | 103.0677 | 0.31671
2020-02-14 | 13:41 | 2.984 | 3.019 | 3.0015 | 320.0476 | 0.927 | 0.938 | 0.9325 | 101.5432 | 0.31728
2020-02-14 | 13:41 | 2.949 | 2.976 | 2.9625 | 315.8891 | 0.916 | 0.924 | 0.92 | 100.182 | 0.31714
2020-02-14 | 13:41 | 2.91 | 2.942 | 2.926 | 311.9971 | 0.904 | 0.913 | 0.9085 | 98.9297 | 0.31709
2020-02-14 | 13:41 | 2.865 | 2.902 | 2.8835 | 307.4654 | 0.888 | 0.901 | 0.8945 | 97.4052 | 0.3168
2020-02-14 | 13:41 | 2.821 | 2.856 | 2.8385 | 302.667 | 0.876 | 0.886 | 0.881 | 95.9352 | 0.31697
2020-02-14 | 13:42 | 2.785 | 2.814 | 2.7995 | 298.5085 | 0.865 | 0.875 | 0.87 | 94.7373 | 0.31737
2020-02-14 | 13:42 | 2.756 | 2.778 | 2.767 | 295.0431 | 0.856 | 0.863 | 0.8595 | 93.594 | 0.31722
2020-02-14 | 13:42 | 2.733 | 2.752 | 2.7425 | 292.4306 | 0.848 | 0.855 | 0.8515 | 92.7228 | 0.31708
2020-02-14 | 13:42 | 2.708 | 2.728 | 2.718 | 289.8182 | 0.84 | 0.847 | 0.8435 | 91.8517 | 0.31693
2020-02-14 | 13:42 | 2.684 | 2.704 | 2.694 | 287.2591 | 0.832 | 0.839 | 0.8355 | 90.9805 | 0.31672
2020-02-14 | 13:42 | 2.664 | 2.68 | 2.672 | 284.9133 | 0.827 | 0.832 | 0.8295 | 90.3272 | 0.31703
2020-02-14 | 13:42 | 2.642 | 2.658 | 2.65 | 282.5674 | 0.819 | 0.825 | 0.822 | 89.5105 | 0.31678
2020-02-14 | 13:42 | 2.622 | 2.637 | 2.6295 | 280.3815 | 0.814 | 0.819 | 0.8165 | 88.9115 | 0.31711
2020-02-14 | 13:42 | 2.606 | 2.62 | 2.613 | 278.6222 | 0.807 | 0.812 | 0.8095 | 88.1493 | 0.31638
2020-02-14 | 13:42 | 2.588 | 2.601 | 2.5945 | 276.6495 | 0.803 | 0.807 | 0.805 | 87.6593 | 0.31686
2020-02-14 | 13:42 | 2.572 | 2.586 | 2.579 | 274.9968 | 0.797 | 0.802 | 0.7995 | 87.0604 | 0.31659
2020-02-14 | 13:42 | 2.559 | 2.57 | 2.5645 | 273.4506 | 0.793 | 0.796 | 0.7945 | 86.5159 | 0.31639
2020-02-14 | 13:42 | 2.547 | 2.556 | 2.5515 | 272.0645 | 0.789 | 0.791 | 0.79 | 86.0259 | 0.3162
2020-02-14 | 13:43 | 2.536 | 2.545 | 2.5405 | 270.8915 | 0.785 | 0.789 | 0.787 | 85.6992 | 0.31636
2020-02-14 | 13:43 | 2.529 | 2.535 | 2.532 | 269.9852 | 0.783 | 0.786 | 0.7845 | 85.427 | 0.31641
2020-02-14 | 13:43 | 2.52 | 2.527 | 2.5235 | 269.0788 | 0.78 | 0.783 | 0.7815 | 85.1003 | 0.31627
2020-02-14 | 13:43 | 2.513 | 2.518 | 2.5155 | 268.2258 | 0.778 | 0.78 | 0.779 | 84.828 | 0.31626
2020-02-14 | 13:43 | 2.506 | 2.512 | 2.509 | 267.5327 | 0.776 | 0.778 | 0.777 | 84.6103 | 0.31626
2020-02-14 | 13:43 | 2.501 | 2.505 | 2.503 | 266.8929 | 0.775 | 0.776 | 0.7755 | 84.4469 | 0.31641
2020-02-14 | 13:43 | 2.495 | 2.5 | 2.4975 | 266.3065 | 0.772 | 0.775 | 0.7735 | 84.2291 | 0.31629
2020-02-14 | 13:43 | 2.492 | 2.497 | 2.4945 | 265.9866 | 0.771 | 0.774 | 0.7725 | 84.1202 | 0.31626
2020-02-14 | 13:43 | 2.488 | 2.491 | 2.4895 | 265.4535 | 0.77 | 0.771 | 0.7705 | 83.9024 | 0.31607
2020-02-14 | 13:43 | 2.486 | 2.489 | 2.4875 | 265.2402 | 0.769 | 0.771 | 0.77 | 83.848 | 0.31612
2020-02-14 | 13:43 | 2.484 | 2.485 | 2.4845 | 264.9203 | 0.768 | 0.77 | 0.769 | 83.7391 | 0.31609
2020-02-14 | 13:44 | 2.482 | 2.483 | 2.4825 | 264.707 | 0.769 | 0.77 | 0.7695 | 83.7935 | 0.31655
2020-02-14 | 13:44 | 2.483 | 2.484 | 2.4835 | 264.8137 | 0.769 | 0.77 | 0.7695 | 83.7935 | 0.31642
2020-02-14 | 13:44 | 2.484 | 2.486 | 2.485 | 264.9736 | 0.769 | 0.771 | 0.77 | 83.848 | 0.31644
2020-02-14 | 13:44 | 2.486 | 2.488 | 2.487 | 265.1869 | 0.77 | 0.771 | 0.7705 | 83.9024 | 0.31639
2020-02-14 | 13:44 | 2.489 | 2.492 | 2.4905 | 265.5601 | 0.771 | 0.772 | 0.7715 | 84.0113 | 0.31636
2020-02-14 | 13:44 | 2.492 | 2.497 | 2.4945 | 265.9866 | 0.772 | 0.773 | 0.7725 | 84.1202 | 0.31626
2020-02-14 | 13:44 | 2.498 | 2.505 | 2.5015 | 266.733 | 0.774 | 0.777 | 0.7755 | 84.4469 | 0.3166
2020-02-14 | 13:44 | 2.506 | 2.516 | 2.511 | 267.746 | 0.776 | 0.78 | 0.778 | 84.7191 | 0.31642
2020-02-14 | 13:44 | 2.517 | 2.528 | 2.5225 | 268.9722 | 0.78 | 0.784 | 0.782 | 85.1547 | 0.31659
2020-02-14 | 13:44 | 2.529 | 2.541 | 2.535 | 270.3051 | 0.784 | 0.788 | 0.786 | 85.5903 | 0.31664
2020-02-14 | 13:44 | 2.543 | 2.551 | 2.547 | 271.5846 | 0.789 | 0.791 | 0.79 | 86.0259 | 0.31676
2020-02-14 | 13:44 | 2.552 | 2.555 | 2.5535 | 272.2777 | 0.791 | 0.793 | 0.792 | 86.2437 | 0.31675
2020-02-14 | 13:44 | 2.556 | 2.558 | 2.557 | 272.6509 | 0.792 | 0.794 | 0.793 | 86.3525 | 0.31671
2020-02-14 | 13:45 | 2.549 | 2.554 | 2.5515 | 272.0645 | 0.791 | 0.792 | 0.7915 | 86.1892 | 0.3168
2020-02-14 | 13:45 | 2.538 | 2.547 | 2.5425 | 271.1048 | 0.788 | 0.791 | 0.7895 | 85.9714 | 0.31712
2020-02-14 | 13:45 | 2.53 | 2.535 | 2.5325 | 270.0385 | 0.784 | 0.786 | 0.785 | 85.4814 | 0.31655
2020-02-14 | 13:45 | 2.524 | 2.527 | 2.5255 | 269.2921 | 0.782 | 0.784 | 0.783 | 85.2636 | 0.31662
2020-02-14 | 13:45 | 2.519 | 2.522 | 2.5205 | 268.759 | 0.781 | 0.782 | 0.7815 | 85.1003 | 0.31664
2020-02-14 | 13:45 | 2.516 | 2.518 | 2.517 | 268.3858 | 0.779 | 0.781 | 0.78 | 84.9369 | 0.31647
2020-02-14 | 13:45 | 2.514 | 2.516 | 2.515 | 268.1725 | 0.779 | 0.78 | 0.7795 | 84.8825 | 0.31652
2020-02-14 | 13:45 | 2.513 | 2.514 | 2.5135 | 268.0126 | 0.777 | 0.78 | 0.7785 | 84.7736 | 0.3163
2020-02-14 | 13:45 | 2.513 | 2.516 | 2.5145 | 268.1192 | 0.779 | 0.78 | 0.7795 | 84.8825 | 0.31658
2020-02-14 | 13:45 | 2.517 | 2.521 | 2.519 | 268.599 | 0.78 | 0.782 | 0.781 | 85.0458 | 0.31663
2020-02-14 | 13:45 | 2.525 | 2.533 | 2.529 | 269.6653 | 0.783 | 0.785 | 0.784 | 85.3725 | 0.31659
2020-02-14 | 13:46 | 2.535 | 2.546 | 2.5405 | 270.8915 | 0.787 | 0.791 | 0.789 | 85.917 | 0.31716
2020-02-14 | 13:46 | 2.549 | 2.566 | 2.5575 | 272.7042 | 0.791 | 0.796 | 0.7935 | 86.407 | 0.31685
2020-02-14 | 13:46 | 2.571 | 2.59 | 2.5805 | 275.1567 | 0.798 | 0.804 | 0.801 | 87.2237 | 0.317
2020-02-14 | 13:46 | 2.594 | 2.602 | 2.598 | 277.0227 | 0.805 | 0.808 | 0.8065 | 87.8226 | 0.31702
2020-02-14 | 13:46 | 2.602 | 2.604 | 2.603 | 277.5559 | 0.808 | 0.808 | 0.808 | 87.9859 | 0.317
2020-02-14 | 13:46 | 2.594 | 2.602 | 2.598 | 277.0227 | 0.804 | 0.808 | 0.806 | 87.7682 | 0.31683
2020-02-14 | 13:46 | 2.582 | 2.59 | 2.586 | 275.7432 | 0.801 | 0.803 | 0.802 | 87.3326 | 0.31672
2020-02-14 | 13:46 | 2.58 | 2.582 | 2.581 | 275.21 | 0.799 | 0.801 | 0.8 | 87.1148 | 0.31654
2020-02-14 | 13:46 | 2.58 | 2.58 | 2.58 | 275.1034 | 0.8 | 0.8 | 0.8 | 87.1148 | 0.31666
2020-02-14 | 13:46 | 2.569 | 2.577 | 2.573 | 274.357 | 0.796 | 0.798 | 0.797 | 86.7881 | 0.31633
2020-02-14 | 13:46 | 2.558 | 2.567 | 2.5625 | 273.2374 | 0.791 | 0.796 | 0.7935 | 86.407 | 0.31623
2020-02-14 | 13:46 | 2.549 | 2.556 | 2.5525 | 272.1711 | 0.79 | 0.792 | 0.791 | 86.1348 | 0.31647
2020-02-14 | 13:46 | 2.546 | 2.548 | 2.547 | 271.5846 | 0.788 | 0.789 | 0.7885 | 85.8625 | 0.31615
2020-02-14 | 13:47 | 2.546 | 2.547 | 2.5465 | 271.5313 | 0.789 | 0.789 | 0.789 | 85.917 | 0.31642
2020-02-14 | 13:47 | 2.546 | 2.547 | 2.5465 | 271.5313 | 0.789 | 0.789 | 0.789 | 85.917 | 0.31642
2020-02-14 | 13:47 | 2.546 | 2.552 | 2.549 | 271.7979 | 0.789 | 0.791 | 0.79 | 86.0259 | 0.31651
2020-02-14 | 13:47 | 2.549 | 2.552 | 2.5505 | 271.9578 | 0.789 | 0.791 | 0.79 | 86.0259 | 0.31632
2020-02-14 | 13:47 | 2.544 | 2.548 | 2.546 | 271.478 | 0.788 | 0.789 | 0.7885 | 85.8625 | 0.31628
2020-02-14 | 13:47 | 2.542 | 2.543 | 2.5425 | 271.1048 | 0.786 | 0.788 | 0.787 | 85.6992 | 0.31611
2020-02-14 | 13:47 | 2.544 | 2.547 | 2.5455 | 271.4247 | 0.788 | 0.79 | 0.789 | 85.917 | 0.31654
2020-02-14 | 13:47 | 2.549 | 2.553 | 2.551 | 272.0111 | 0.789 | 0.79 | 0.7895 | 85.9714 | 0.31606
2020-02-14 | 13:47 | 2.554 | 2.557 | 2.5555 | 272.491 | 0.79 | 0.793 | 0.7915 | 86.1892 | 0.3163
2020-02-14 | 13:47 | 2.559 | 2.563 | 2.561 | 273.0774 | 0.792 | 0.794 | 0.793 | 86.3525 | 0.31622
2020-02-14 | 13:47 | 2.564 | 2.567 | 2.5655 | 273.5573 | 0.793 | 0.796 | 0.7945 | 86.5159 | 0.31626
2020-02-14 | 13:48 | 2.568 | 2.571 | 2.5695 | 273.9838 | 0.795 | 0.796 | 0.7955 | 86.6248 | 0.31617
2020-02-14 | 13:48 | 2.571 | 2.58 | 2.5755 | 274.6236 | 0.797 | 0.8 | 0.7985 | 86.9515 | 0.31662
2020-02-14 | 13:48 | 2.583 | 2.592 | 2.5875 | 275.9031 | 0.8 | 0.804 | 0.802 | 87.3326 | 0.31653
2020-02-14 | 13:48 | 2.596 | 2.606 | 2.601 | 277.3426 | 0.803 | 0.808 | 0.8055 | 87.7137 | 0.31626
2020-02-14 | 13:48 | 2.609 | 2.624 | 2.6165 | 278.9954 | 0.81 | 0.815 | 0.8125 | 88.476 | 0.31712
2020-02-14 | 13:48 | 2.629 | 2.648 | 2.6385 | 281.3412 | 0.816 | 0.822 | 0.819 | 89.1838 | 0.317
2020-02-14 | 13:48 | 2.653 | 2.678 | 2.6655 | 284.2202 | 0.824 | 0.831 | 0.8275 | 90.1094 | 0.31704
2020-02-14 | 13:48 | 2.681 | 2.698 | 2.6895 | 286.7793 | 0.832 | 0.839 | 0.8355 | 90.9805 | 0.31725
2020-02-14 | 13:48 | 2.702 | 2.723 | 2.7125 | 289.2318 | 0.84 | 0.846 | 0.843 | 91.7972 | 0.31738
2020-02-14 | 13:48 | 2.729 | 2.752 | 2.7405 | 292.2174 | 0.85 | 0.856 | 0.853 | 92.8862 | 0.31787
2020-02-14 | 13:48 | 2.757 | 2.78 | 2.7685 | 295.203 | 0.857 | 0.867 | 0.862 | 93.8662 | 0.31797
2020-02-14 | 13:48 | 2.786 | 2.811 | 2.7985 | 298.4019 | 0.867 | 0.876 | 0.8715 | 94.9007 | 0.31803
2020-02-14 | 13:48 | 2.818 | 2.841 | 2.8295 | 301.7074 | 0.876 | 0.886 | 0.881 | 95.9352 | 0.31797
2020-02-14 | 13:49 | 2.846 | 2.868 | 2.857 | 304.6397 | 0.887 | 0.894 | 0.8905 | 96.9697 | 0.31831
2020-02-14 | 13:49 | 2.873 | 2.897 | 2.885 | 307.6253 | 0.896 | 0.904 | 0.9 | 98.0042 | 0.31858
2020-02-14 | 13:49 | 2.902 | 2.919 | 2.9105 | 310.3444 | 0.905 | 0.912 | 0.9085 | 98.9297 | 0.31877
2020-02-14 | 13:49 | 2.923 | 2.937 | 2.93 | 312.4236 | 0.913 | 0.918 | 0.9155 | 99.692 | 0.31909
2020-02-14 | 13:49 | 2.941 | 2.961 | 2.951 | 314.6628 | 0.918 | 0.926 | 0.922 | 100.3998 | 0.31907
2020-02-14 | 13:49 | 2.968 | 2.998 | 2.983 | 318.075 | 0.927 | 0.938 | 0.9325 | 101.5432 | 0.31924
2020-02-14 | 13:49 | 3.006 | 3.045 | 3.0255 | 322.6067 | 0.941 | 0.954 | 0.9475 | 103.1766 | 0.31982
2020-02-14 | 13:49 | 3.055 | 3.098 | 3.0765 | 328.0448 | 0.956 | 0.971 | 0.9635 | 104.9189 | 0.31983
2020-02-14 | 13:49 | 3.111 | 3.161 | 3.136 | 334.3892 | 0.975 | 0.992 | 0.9835 | 107.0968 | 0.32028
2020-02-14 | 13:49 | 3.175 | 3.225 | 3.2 | 341.2135 | 0.996 | 1.013 | 1.0045 | 109.3835 | 0.32057
2020-02-14 | 13:49 | 3.237 | 3.281 | 3.259 | 347.5046 | 1.016 | 1.032 | 1.024 | 111.5069 | 0.32088
2020-02-14 | 13:50 | 3.293 | 3.335 | 3.314 | 353.3692 | 1.034 | 1.048 | 1.041 | 113.3581 | 0.32079
2020-02-14 | 13:50 | 3.347 | 3.389 | 3.368 | 359.1272 | 1.052 | 1.068 | 1.06 | 115.4271 | 0.32141
2020-02-14 | 13:50 | 3.402 | 3.462 | 3.432 | 365.9515 | 1.07 | 1.091 | 1.0805 | 117.6594 | 0.32152
2020-02-14 | 13:50 | 3.48 | 3.548 | 3.514 | 374.6951 | 1.096 | 1.118 | 1.107 | 120.5451 | 0.32172
2020-02-14 | 13:50 | 3.566 | 3.634 | 3.6 | 383.8652 | 1.126 | 1.147 | 1.1365 | 123.7575 | 0.3224
2020-02-14 | 13:50 | 3.652 | 3.733 | 3.6925 | 393.7284 | 1.153 | 1.179 | 1.166 | 126.9698 | 0.32248
2020-02-14 | 13:50 | 3.754 | 3.848 | 3.801 | 405.2977 | 1.187 | 1.219 | 1.203 | 130.9989 | 0.32322
2020-02-14 | 13:50 | 3.874 | 3.973 | 3.9235 | 418.3598 | 1.226 | 1.257 | 1.2415 | 135.1913 | 0.32315
2020-02-14 | 13:50 | 3.996 | 4.081 | 4.0385 | 430.6221 | 1.266 | 1.294 | 1.28 | 139.3837 | 0.32368
2020-02-14 | 13:50 | 4.1 | 4.163 | 4.1315 | 440.5386 | 1.3 | 1.321 | 1.3105 | 142.7049 | 0.32393
2020-02-14 | 13:50 | 4.177 | 4.236 | 4.2065 | 448.5358 | 1.325 | 1.344 | 1.3345 | 145.3184 | 0.32398
2020-02-14 | 13:50 | 4.25 | 4.302 | 4.276 | 455.9466 | 1.349 | 1.365 | 1.357 | 147.7685 | 0.32409
2020-02-14 | 13:50 | 4.316 | 4.377 | 4.3465 | 463.4639 | 1.371 | 1.392 | 1.3815 | 150.4364 | 0.32459
2020-02-14 | 13:51 | 4.397 | 4.473 | 4.435 | 472.9006 | 1.395 | 1.422 | 1.4085 | 153.3765 | 0.32433
2020-02-14 | 13:51 | 4.492 | 4.565 | 4.5285 | 482.8704 | 1.43 | 1.452 | 1.441 | 156.9155 | 0.32496
2020-02-14 | 13:51 | 4.583 | 4.653 | 4.618 | 492.4138 | 1.457 | 1.481 | 1.469 | 159.9646 | 0.32486
2020-02-14 | 13:51 | 4.67 | 4.738 | 4.704 | 501.5839 | 1.486 | 1.509 | 1.4975 | 163.068 | 0.32511
2020-02-14 | 13:51 | 4.754 | 4.821 | 4.7875 | 510.4874 | 1.514 | 1.536 | 1.525 | 166.0626 | 0.3253
2020-02-14 | 13:51 | 4.836 | 4.901 | 4.8685 | 519.1244 | 1.54 | 1.56 | 1.55 | 168.7849 | 0.32513
2020-02-14 | 13:51 | 4.916 | 4.973 | 4.9445 | 527.2282 | 1.565 | 1.586 | 1.5755 | 171.5617 | 0.3254
2020-02-14 | 13:51 | 4.986 | 5.028 | 5.007 | 533.8925 | 1.589 | 1.603 | 1.596 | 173.794 | 0.32552
2020-02-14 | 13:51 | 5.04 | 5.107 | 5.0735 | 540.9834 | 1.606 | 1.627 | 1.6165 | 176.0263 | 0.32538
2020-02-14 | 13:51 | 5.127 | 5.198 | 5.1625 | 550.4734 | 1.635 | 1.658 | 1.6465 | 179.2932 | 0.32571
2020-02-14 | 13:51 | 5.212 | 5.277 | 5.2445 | 559.217 | 1.663 | 1.685 | 1.674 | 182.2877 | 0.32597
2020-02-14 | 13:52 | 5.294 | 5.367 | 5.3305 | 568.3871 | 1.69 | 1.714 | 1.702 | 185.3367 | 0.32607
2020-02-14 | 13:52 | 5.383 | 5.439 | 5.411 | 576.9707 | 1.719 | 1.738 | 1.7285 | 188.2224 | 0.32623
2020-02-14 | 13:52 | 5.453 | 5.508 | 5.4805 | 584.3815 | 1.742 | 1.761 | 1.7515 | 190.727 | 0.32637
2020-02-14 | 13:52 | 5.522 | 5.569 | 5.5455 | 591.3124 | 1.764 | 1.78 | 1.772 | 192.9593 | 0.32632
2020-02-14 | 13:52 | 5.577 | 5.598 | 5.5875 | 595.7908 | 1.782 | 1.787 | 1.7845 | 194.3205 | 0.32616
2020-02-14 | 13:52 | 5.602 | 5.613 | 5.6075 | 597.9234 | 1.789 | 1.793 | 1.791 | 195.0283 | 0.32618
2020-02-14 | 13:52 | 5.616 | 5.636 | 5.626 | 599.896 | 1.793 | 1.8 | 1.7965 | 195.6272 | 0.3261
2020-02-14 | 13:52 | 5.643 | 5.668 | 5.6555 | 603.0416 | 1.801 | 1.81 | 1.8055 | 196.6072 | 0.32603
2020-02-14 | 13:52 | 5.674 | 5.681 | 5.6775 | 605.3874 | 1.812 | 1.814 | 1.813 | 197.4239 | 0.32611
2020-02-14 | 13:52 | 5.667 | 5.68 | 5.6735 | 604.9609 | 1.807 | 1.813 | 1.81 | 197.0972 | 0.3258
2020-02-14 | 13:52 | 5.655 | 5.665 | 5.66 | 603.5214 | 1.802 | 1.806 | 1.804 | 196.4439 | 0.3255
2020-02-14 | 13:52 | 5.639 | 5.651 | 5.645 | 601.922 | 1.796 | 1.801 | 1.7985 | 195.845 | 0.32537
2020-02-14 | 13:52 | 5.63 | 5.634 | 5.632 | 600.5358 | 1.793 | 1.795 | 1.794 | 195.3549 | 0.3253
2020-02-14 | 13:53 | 5.636 | 5.641 | 5.6385 | 601.2289 | 1.795 | 1.797 | 1.796 | 195.5727 | 0.32529
2020-02-14 | 13:53 | 5.63 | 5.641 | 5.6355 | 600.909 | 1.791 | 1.796 | 1.7935 | 195.3005 | 0.32501
2020-02-14 | 13:53 | 5.58 | 5.622 | 5.601 | 597.2303 | 1.774 | 1.79 | 1.782 | 194.0482 | 0.32491
2020-02-14 | 13:53 | 5.499 | 5.565 | 5.532 | 589.8729 | 1.746 | 1.769 | 1.7575 | 191.3803 | 0.32444
2020-02-14 | 13:53 | 5.377 | 5.48 | 5.4285 | 578.8367 | 1.705 | 1.74 | 1.7225 | 187.5691 | 0.32404
2020-02-14 | 13:53 | 5.201 | 5.342 | 5.2715 | 562.0959 | 1.647 | 1.693 | 1.67 | 181.8521 | 0.32353
2020-02-14 | 13:53 | 5.026 | 5.167 | 5.0965 | 543.4358 | 1.585 | 1.635 | 1.61 | 175.3185 | 0.32261
2020-02-14 | 13:53 | 4.866 | 4.99 | 4.928 | 525.4688 | 1.535 | 1.575 | 1.555 | 169.3294 | 0.32224
2020-02-14 | 13:53 | 4.753 | 4.84 | 4.7965 | 511.4471 | 1.497 | 1.525 | 1.511 | 164.5381 | 0.32171
2020-02-14 | 13:53 | 4.669 | 4.734 | 4.7015 | 501.3173 | 1.471 | 1.491 | 1.481 | 161.2713 | 0.3217
2020-02-14 | 13:53 | 4.629 | 4.658 | 4.6435 | 495.1328 | 1.458 | 1.467 | 1.4625 | 159.2567 | 0.32164
2020-02-14 | 13:54 | 4.601 | 4.625 | 4.613 | 491.8806 | 1.45 | 1.456 | 1.453 | 158.2223 | 0.32167
2020-02-14 | 13:54 | 4.557 | 4.595 | 4.576 | 487.9353 | 1.435 | 1.447 | 1.441 | 156.9155 | 0.32159
2020-02-14 | 13:54 | 4.505 | 4.546 | 4.5255 | 482.5505 | 1.419 | 1.432 | 1.4255 | 155.2277 | 0.32168
2020-02-14 | 13:54 | 4.456 | 4.495 | 4.4755 | 477.2191 | 1.402 | 1.415 | 1.4085 | 153.3765 | 0.3214
2020-02-14 | 13:54 | 4.414 | 4.445 | 4.4295 | 472.3141 | 1.391 | 1.401 | 1.396 | 152.0153 | 0.32185
2020-02-14 | 13:54 | 4.389 | 4.408 | 4.3985 | 469.0086 | 1.384 | 1.39 | 1.387 | 151.0353 | 0.32203
2020-02-14 | 13:54 | 4.373 | 4.385 | 4.379 | 466.9294 | 1.379 | 1.383 | 1.381 | 150.3819 | 0.32207
2020-02-14 | 13:54 | 4.372 | 4.379 | 4.3755 | 466.5562 | 1.381 | 1.381 | 1.381 | 150.3819 | 0.32232
2020-02-14 | 13:54 | 4.383 | 4.398 | 4.3905 | 468.1556 | 1.385 | 1.39 | 1.3875 | 151.0897 | 0.32273
2020-02-14 | 13:54 | 4.401 | 4.408 | 4.4045 | 469.6484 | 1.39 | 1.395 | 1.3925 | 151.6342 | 0.32287
2020-02-14 | 13:54 | 4.411 | 4.435 | 4.423 | 471.6211 | 1.396 | 1.404 | 1.4 | 152.4509 | 0.32325
2020-02-14 | 13:54 | 4.445 | 4.491 | 4.468 | 476.4194 | 1.408 | 1.425 | 1.4165 | 154.2476 | 0.32376
2020-02-14 | 13:54 | 4.501 | 4.544 | 4.5225 | 482.2307 | 1.428 | 1.443 | 1.4355 | 156.3166 | 0.32415
2020-02-14 | 13:55 | 4.558 | 4.613 | 4.5855 | 488.9483 | 1.447 | 1.467 | 1.457 | 158.6578 | 0.32449
2020-02-14 | 13:55 | 4.627 | 4.683 | 4.655 | 496.359 | 1.472 | 1.491 | 1.4815 | 161.3257 | 0.32502
2020-02-14 | 13:55 | 4.695 | 4.75 | 4.7225 | 503.5565 | 1.495 | 1.514 | 1.5045 | 163.8303 | 0.32535
2020-02-14 | 13:55 | 4.764 | 4.834 | 4.799 | 511.7136 | 1.519 | 1.542 | 1.5305 | 166.6615 | 0.32569
2020-02-14 | 13:55 | 4.852 | 4.924 | 4.888 | 521.2036 | 1.549 | 1.575 | 1.562 | 170.0916 | 0.32634
2020-02-14 | 13:55 | 4.941 | 5.009 | 4.975 | 530.4804 | 1.579 | 1.603 | 1.591 | 173.2496 | 0.32659
2020-02-14 | 13:55 | 5.025 | 5.094 | 5.0595 | 539.4906 | 1.607 | 1.631 | 1.619 | 176.2986 | 0.32679
2020-02-14 | 13:55 | 5.113 | 5.187 | 5.15 | 549.1405 | 1.637 | 1.662 | 1.6495 | 179.6198 | 0.32709
2020-02-14 | 13:55 | 5.207 | 5.285 | 5.246 | 559.3769 | 1.67 | 1.694 | 1.682 | 183.1589 | 0.32743
2020-02-14 | 13:55 | 5.3 | 5.343 | 5.3215 | 567.4274 | 1.699 | 1.713 | 1.706 | 185.7723 | 0.32739
2020-02-14 | 13:55 | 5.348 | 5.368 | 5.358 | 571.3194 | 1.715 | 1.722 | 1.7185 | 187.1335 | 0.32755
2020-02-14 | 13:56 | 5.374 | 5.407 | 5.3905 | 574.7848 | 1.724 | 1.735 | 1.7295 | 188.3313 | 0.32766
2020-02-14 | 13:56 | 5.415 | 5.434 | 5.4245 | 578.4102 | 1.736 | 1.742 | 1.739 | 189.3658 | 0.32739
2020-02-14 | 13:56 | 5.441 | 5.473 | 5.457 | 581.8757 | 1.745 | 1.755 | 1.75 | 190.5636 | 0.3275
2020-02-14 | 13:56 | 5.481 | 5.516 | 5.4985 | 586.3008 | 1.759 | 1.769 | 1.764 | 192.0881 | 0.32763
2020-02-14 | 13:56 | 5.521 | 5.527 | 5.524 | 589.0198 | 1.771 | 1.773 | 1.772 | 192.9593 | 0.32759
2020-02-14 | 13:56 | 5.527 | 5.538 | 5.5325 | 589.9262 | 1.773 | 1.776 | 1.7745 | 193.2315 | 0.32755
2020-02-14 | 13:56 | 5.545 | 5.586 | 5.5655 | 593.4449 | 1.778 | 1.793 | 1.7855 | 194.4293 | 0.32763
2020-02-14 | 13:56 | 5.597 | 5.648 | 5.6225 | 599.5228 | 1.796 | 1.813 | 1.8045 | 196.4983 | 0.32776
2020-02-14 | 13:56 | 5.66 | 5.702 | 5.681 | 605.7606 | 1.817 | 1.833 | 1.825 | 198.7306 | 0.32807
2020-02-14 | 13:56 | 5.711 | 5.751 | 5.731 | 611.0921 | 1.834 | 1.848 | 1.841 | 200.4729 | 0.32806
2020-02-14 | 13:56 | 5.766 | 5.831 | 5.7985 | 618.2895 | 1.853 | 1.875 | 1.864 | 202.9775 | 0.32829
2020-02-14 | 13:56 | 5.846 | 5.904 | 5.875 | 626.4467 | 1.879 | 1.899 | 1.889 | 205.6998 | 0.32836
2020-02-14 | 13:56 | 5.915 | 5.954 | 5.9345 | 632.7911 | 1.903 | 1.914 | 1.9085 | 207.8232 | 0.32842
2020-02-14 | 13:57 | 5.961 | 5.983 | 5.972 | 636.7897 | 1.917 | 1.923 | 1.92 | 209.0755 | 0.32833
2020-02-14 | 13:57 | 5.988 | 6.014 | 6.001 | 639.882 | 1.926 | 1.935 | 1.9305 | 210.2189 | 0.32853
2020-02-14 | 13:57 | 6.022 | 6.046 | 6.034 | 643.4007 | 1.938 | 1.945 | 1.9415 | 211.4167 | 0.32859
2020-02-14 | 13:57 | 6.05 | 6.065 | 6.0575 | 645.9065 | 1.946 | 1.951 | 1.9485 | 212.179 | 0.3285
2020-02-14 | 13:57 | 6.065 | 6.069 | 6.067 | 646.9195 | 1.951 | 1.952 | 1.9515 | 212.5057 | 0.32849
2020-02-14 | 13:57 | 6.051 | 6.07 | 6.0605 | 646.2264 | 1.945 | 1.952 | 1.9485 | 212.179 | 0.32834
2020-02-14 | 13:57 | 6.034 | 6.047 | 6.0405 | 644.0938 | 1.937 | 1.943 | 1.94 | 211.2534 | 0.32799
2020-02-14 | 13:57 | 6.014 | 6.028 | 6.021 | 642.0146 | 1.932 | 1.937 | 1.9345 | 210.6545 | 0.32811
2020-02-14 | 13:57 | 6.011 | 6.014 | 6.0125 | 641.1082 | 1.931 | 1.933 | 1.932 | 210.3822 | 0.32815
2020-02-14 | 13:57 | 6.016 | 6.027 | 6.0215 | 642.0679 | 1.933 | 1.937 | 1.935 | 210.7089 | 0.32817
2020-02-14 | 13:57 | 6.028 | 6.044 | 6.036 | 643.614 | 1.937 | 1.942 | 1.9395 | 211.1989 | 0.32815
2020-02-14 | 13:58 | 6.046 | 6.059 | 6.0525 | 645.3734 | 1.943 | 1.947 | 1.945 | 211.7979 | 0.32818
2020-02-14 | 13:58 | 6.059 | 6.063 | 6.061 | 646.2797 | 1.947 | 1.948 | 1.9475 | 212.0701 | 0.32814
2020-02-14 | 13:58 | 6.062 | 6.065 | 6.0635 | 646.5463 | 1.947 | 1.948 | 1.9475 | 212.0701 | 0.328
2020-02-14 | 13:58 | 6.065 | 6.069 | 6.067 | 646.9195 | 1.948 | 1.951 | 1.9495 | 212.2879 | 0.32815
2020-02-14 | 13:58 | 6.059 | 6.069 | 6.064 | 646.5996 | 1.945 | 1.949 | 1.947 | 212.0156 | 0.32789
2020-02-14 | 13:58 | 6.04 | 6.054 | 6.047 | 644.7869 | 1.938 | 1.943 | 1.9405 | 211.3078 | 0.32772
2020-02-14 | 13:58 | 6.028 | 6.037 | 6.0325 | 643.2408 | 1.936 | 1.937 | 1.9365 | 210.8723 | 0.32783
2020-02-14 | 13:58 | 6.024 | 6.027 | 6.0255 | 642.4944 | 1.935 | 1.936 | 1.9355 | 210.7634 | 0.32804
2020-02-14 | 13:58 | 6.006 | 6.022 | 6.014 | 641.2681 | 1.928 | 1.933 | 1.9305 | 210.2189 | 0.32782
2020-02-14 | 13:58 | 6.003 | 6.005 | 6.004 | 640.2019 | 1.927 | 1.928 | 1.9275 | 209.8922 | 0.32785
2020-02-14 | 13:58 | 5.989 | 6.004 | 5.9965 | 639.4021 | 1.924 | 1.928 | 1.926 | 209.7289 | 0.32801
2020-02-14 | 13:58 | 5.948 | 5.983 | 5.9655 | 636.0966 | 1.909 | 1.921 | 1.915 | 208.5311 | 0.32783
2020-02-14 | 13:58 | 5.896 | 5.939 | 5.9175 | 630.9784 | 1.893 | 1.906 | 1.8995 | 206.8432 | 0.32781
2020-02-14 | 13:59 | 5.853 | 5.887 | 5.87 | 625.9135 | 1.877 | 1.889 | 1.883 | 205.0465 | 0.3276
2020-02-14 | 13:59 | 5.782 | 5.843 | 5.8125 | 619.7824 | 1.854 | 1.874 | 1.864 | 202.9775 | 0.3275
2020-02-14 | 13:59 | 5.695 | 5.766 | 5.7305 | 611.0388 | 1.824 | 1.847 | 1.8355 | 199.874 | 0.32711
2020-02-14 | 13:59 | 5.616 | 5.679 | 5.6475 | 602.1885 | 1.798 | 1.818 | 1.808 | 196.8795 | 0.32694
2020-02-14 | 13:59 | 5.563 | 5.604 | 5.5835 | 595.3643 | 1.779 | 1.793 | 1.786 | 194.4838 | 0.32666
2020-02-14 | 13:59 | 5.504 | 5.554 | 5.529 | 589.553 | 1.761 | 1.778 | 1.7695 | 192.6871 | 0.32684
2020-02-14 | 13:59 | 5.421 | 5.489 | 5.455 | 581.6624 | 1.731 | 1.756 | 1.7435 | 189.8558 | 0.3264
2020-02-14 | 13:59 | 5.318 | 5.401 | 5.3595 | 571.4793 | 1.698 | 1.726 | 1.712 | 186.4257 | 0.32622
2020-02-14 | 13:59 | 5.214 | 5.297 | 5.2555 | 560.3899 | 1.662 | 1.69 | 1.676 | 182.5055 | 0.32568
2020-02-14 | 13:59 | 5.081 | 5.19 | 5.1355 | 547.5944 | 1.618 | 1.654 | 1.636 | 178.1498 | 0.32533
2020-02-14 | 13:59 | 4.94 | 5.054 | 4.997 | 532.8262 | 1.571 | 1.609 | 1.59 | 173.1407 | 0.32495
2020-02-14 | 14:00 | 4.792 | 4.909 | 4.8505 | 517.205 | 1.523 | 1.562 | 1.5425 | 167.9682 | 0.32476
2020-02-14 | 14:00 | 4.685 | 4.765 | 4.725 | 503.8231 | 1.488 | 1.514 | 1.501 | 163.4491 | 0.32442
2020-02-14 | 14:00 | 4.625 | 4.671 | 4.648 | 495.6126 | 1.469 | 1.485 | 1.477 | 160.8357 | 0.32452
2020-02-14 | 14:00 | 4.6 | 4.617 | 4.6085 | 491.4008 | 1.462 | 1.466 | 1.464 | 159.4201 | 0.32442
2020-02-14 | 14:00 | 4.609 | 4.678 | 4.6435 | 495.1328 | 1.468 | 1.49 | 1.479 | 161.0535 | 0.32527
2020-02-14 | 14:00 | 4.7 | 4.766 | 4.733 | 504.6761 | 1.498 | 1.522 | 1.51 | 164.4292 | 0.32581
2020-02-14 | 14:00 | 4.777 | 4.818 | 4.7975 | 511.5537 | 1.525 | 1.542 | 1.5335 | 166.9882 | 0.32643
2020-02-14 | 14:00 | 4.833 | 4.921 | 4.877 | 520.0307 | 1.545 | 1.577 | 1.561 | 169.9828 | 0.32687
2020-02-14 | 14:00 | 4.95 | 5.061 | 5.0055 | 533.7326 | 1.588 | 1.626 | 1.607 | 174.9919 | 0.32786
2020-02-14 | 14:00 | 5.083 | 5.133 | 5.108 | 544.6621 | 1.633 | 1.65 | 1.6415 | 178.7487 | 0.32818
2020-02-14 | 14:00 | 5.141 | 5.171 | 5.156 | 549.7803 | 1.651 | 1.662 | 1.6565 | 180.3821 | 0.3281
2020-02-14 | 14:00 | 5.179 | 5.219 | 5.199 | 554.3653 | 1.666 | 1.678 | 1.672 | 182.0699 | 0.32843
2020-02-14 | 14:00 | 5.231 | 5.288 | 5.2595 | 560.8164 | 1.683 | 1.703 | 1.693 | 184.3567 | 0.32873
2020-02-14 | 14:01 | 5.302 | 5.369 | 5.3355 | 568.9202 | 1.707 | 1.731 | 1.719 | 187.1879 | 0.32902
2020-02-14 | 14:01 | 5.385 | 5.447 | 5.416 | 577.5039 | 1.734 | 1.756 | 1.745 | 190.0192 | 0.32904
2020-02-14 | 14:01 | 5.462 | 5.526 | 5.494 | 585.821 | 1.762 | 1.783 | 1.7725 | 193.0137 | 0.32948
2020-02-14 | 14:01 | 5.54 | 5.588 | 5.564 | 593.285 | 1.788 | 1.803 | 1.7955 | 195.5183 | 0.32955
2020-02-14 | 14:01 | 5.598 | 5.641 | 5.6195 | 599.2029 | 1.807 | 1.821 | 1.814 | 197.5328 | 0.32966
2020-02-14 | 14:01 | 5.653 | 5.702 | 5.6775 | 605.3874 | 1.826 | 1.841 | 1.8335 | 199.6562 | 0.3298
2020-02-14 | 14:01 | 5.713 | 5.763 | 5.738 | 611.8385 | 1.844 | 1.862 | 1.853 | 201.7797 | 0.32979
2020-02-14 | 14:01 | 5.775 | 5.845 | 5.81 | 619.5158 | 1.866 | 1.89 | 1.878 | 204.502 | 0.3301
2020-02-14 | 14:01 | 5.862 | 5.936 | 5.899 | 629.0058 | 1.895 | 1.919 | 1.907 | 207.6599 | 0.33014
2020-02-14 | 14:01 | 5.955 | 6.035 | 5.995 | 639.2422 | 1.926 | 1.952 | 1.939 | 211.1445 | 0.3303
2020-02-14 | 14:01 | 6.056 | 6.126 | 6.091 | 649.4786 | 1.959 | 1.982 | 1.9705 | 214.5746 | 0.33038
2020-02-14 | 14:02 | 6.141 | 6.194 | 6.1675 | 657.6357 | 1.987 | 2.005 | 1.996 | 217.3514 | 0.3305
2020-02-14 | 14:02 | 6.207 | 6.247 | 6.227 | 663.9802 | 2.009 | 2.022 | 2.0155 | 219.4749 | 0.33054
2020-02-14 | 14:02 | 6.259 | 6.296 | 6.2775 | 669.3649 | 2.025 | 2.037 | 2.031 | 221.1627 | 0.33041
2020-02-14 | 14:02 | 6.303 | 6.335 | 6.319 | 673.7901 | 2.041 | 2.05 | 2.0455 | 222.7417 | 0.33058
2020-02-14 | 14:02 | 6.341 | 6.364 | 6.3525 | 677.3621 | 2.053 | 2.059 | 2.056 | 223.885 | 0.33052
2020-02-14 | 14:02 | 6.367 | 6.38 | 6.3735 | 679.6014 | 2.06 | 2.063 | 2.0615 | 224.484 | 0.33032
2020-02-14 | 14:02 | 6.384 | 6.392 | 6.388 | 681.1475 | 2.063 | 2.067 | 2.065 | 224.8651 | 0.33013
2020-02-14 | 14:02 | 6.393 | 6.402 | 6.3975 | 682.1605 | 2.066 | 2.07 | 2.068 | 225.1918 | 0.33012
2020-02-14 | 14:02 | 6.406 | 6.41 | 6.408 | 683.2801 | 2.069 | 2.071 | 2.07 | 225.4095 | 0.32989
2020-02-14 | 14:02 | 6.412 | 6.417 | 6.4145 | 683.9732 | 2.071 | 2.074 | 2.0725 | 225.6818 | 0.32996
2020-02-14 | 14:02 | 6.416 | 6.419 | 6.4175 | 684.293 | 2.072 | 2.074 | 2.073 | 225.7362 | 0.32988
2020-02-14 | 14:02 | 6.419 | 6.424 | 6.4215 | 684.7196 | 2.073 | 2.074 | 2.0735 | 225.7907 | 0.32976
2020-02-14 | 14:02 | 6.425 | 6.427 | 6.426 | 685.1994 | 2.074 | 2.075 | 2.0745 | 225.8996 | 0.32968
2020-02-14 | 14:03 | 6.429 | 6.431 | 6.43 | 685.6259 | 2.074 | 2.075 | 2.0745 | 225.8996 | 0.32948
2020-02-14 | 14:03 | 6.429 | 6.432 | 6.4305 | 685.6792 | 2.074 | 2.075 | 2.0745 | 225.8996 | 0.32945
2020-02-14 | 14:03 | 6.432 | 6.433 | 6.4325 | 685.8925 | 2.075 | 2.075 | 2.075 | 225.954 | 0.32943
2020-02-14 | 14:03 | 6.43 | 6.432 | 6.431 | 685.7325 | 2.073 | 2.075 | 2.074 | 225.8451 | 0.32935
2020-02-14 | 14:03 | 6.423 | 6.429 | 6.426 | 685.1994 | 2.072 | 2.072 | 2.072 | 225.6273 | 0.32929
2020-02-14 | 14:03 | 6.418 | 6.422 | 6.42 | 684.5596 | 2.07 | 2.071 | 2.0705 | 225.464 | 0.32936
2020-02-14 | 14:03 | 6.417 | 6.42 | 6.4185 | 684.3997 | 2.068 | 2.07 | 2.069 | 225.3007 | 0.32919
2020-02-14 | 14:03 | 6.411 | 6.417 | 6.414 | 683.9198 | 2.067 | 2.07 | 2.0685 | 225.2462 | 0.32935
2020-02-14 | 14:03 | 6.404 | 6.41 | 6.407 | 683.1734 | 2.064 | 2.067 | 2.0655 | 224.9195 | 0.32923
2020-02-14 | 14:03 | 6.403 | 6.405 | 6.404 | 682.8535 | 2.063 | 2.065 | 2.064 | 224.7562 | 0.32914
2020-02-14 | 14:03 | 6.388 | 6.4 | 6.394 | 681.7873 | 2.058 | 2.063 | 2.0605 | 224.3751 | 0.3291
2020-02-14 | 14:04 | 6.373 | 6.383 | 6.378 | 680.0812 | 2.053 | 2.057 | 2.055 | 223.7761 | 0.32904
2020-02-14 | 14:04 | 6.365 | 6.372 | 6.3685 | 679.0682 | 2.051 | 2.052 | 2.0515 | 223.395 | 0.32897
2020-02-14 | 14:04 | 6.363 | 6.369 | 6.366 | 678.8016 | 2.05 | 2.051 | 2.0505 | 223.2861 | 0.32894
2020-02-14 | 14:04 | 6.373 | 6.39 | 6.3815 | 680.4544 | 2.053 | 2.061 | 2.057 | 223.9939 | 0.32918
2020-02-14 | 14:04 | 6.395 | 6.407 | 6.401 | 682.5337 | 2.062 | 2.066 | 2.064 | 224.7562 | 0.3293
2020-02-14 | 14:04 | 6.409 | 6.414 | 6.4115 | 683.6533 | 2.067 | 2.069 | 2.068 | 225.1918 | 0.32939
2020-02-14 | 14:04 | 6.406 | 6.413 | 6.4095 | 683.44 | 2.067 | 2.069 | 2.068 | 225.1918 | 0.3295
2020-02-14 | 14:04 | 6.395 | 6.404 | 6.3995 | 682.3737 | 2.063 | 2.065 | 2.064 | 224.7562 | 0.32937
2020-02-14 | 14:04 | 6.373 | 6.391 | 6.382 | 680.5077 | 2.056 | 2.061 | 2.0585 | 224.1573 | 0.3294
2020-02-14 | 14:04 | 6.342 | 6.367 | 6.3545 | 677.5754 | 2.044 | 2.053 | 2.0485 | 223.0683 | 0.32922
2020-02-14 | 14:04 | 6.312 | 6.336 | 6.324 | 674.3232 | 2.034 | 2.043 | 2.0385 | 221.9794 | 0.32919
2020-02-14 | 14:04 | 6.283 | 6.306 | 6.2945 | 671.1776 | 2.024 | 2.033 | 2.0285 | 220.8905 | 0.32911
2020-02-14 | 14:04 | 6.25 | 6.277 | 6.2635 | 667.8721 | 2.014 | 2.022 | 2.018 | 219.7471 | 0.32903
2020-02-14 | 14:05 | 6.223 | 6.246 | 6.2345 | 664.7799 | 2.005 | 2.012 | 2.0085 | 218.7126 | 0.329
2020-02-14 | 14:05 | 6.176 | 6.217 | 6.1965 | 660.728 | 1.988 | 2.002 | 1.995 | 217.2425 | 0.32879
2020-02-14 | 14:05 | 6.135 | 6.167 | 6.151 | 655.8764 | 1.974 | 1.985 | 1.9795 | 215.5547 | 0.32865
2020-02-14 | 14:05 | 6.11 | 6.13 | 6.12 | 652.5708 | 1.966 | 1.973 | 1.9695 | 214.4658 | 0.32865
2020-02-14 | 14:05 | 6.075 | 6.104 | 6.0895 | 649.3187 | 1.955 | 1.965 | 1.96 | 213.4313 | 0.3287
2020-02-14 | 14:05 | 6.044 | 6.066 | 6.055 | 645.6399 | 1.945 | 1.952 | 1.9485 | 212.179 | 0.32863
2020-02-14 | 14:05 | 6.034 | 6.041 | 6.0375 | 643.7739 | 1.942 | 1.944 | 1.943 | 211.5801 | 0.32866
2020-02-14 | 14:05 | 6.015 | 6.032 | 6.0235 | 642.2811 | 1.937 | 1.942 | 1.9395 | 211.1989 | 0.32883
2020-02-14 | 14:05 | 5.985 | 6.011 | 5.998 | 639.5621 | 1.926 | 1.934 | 1.93 | 210.1645 | 0.32861
2020-02-14 | 14:05 | 5.93 | 5.975 | 5.9525 | 634.7104 | 1.907 | 1.923 | 1.915 | 208.5311 | 0.32855
2020-02-14 | 14:05 | 5.85 | 5.918 | 5.884 | 627.4063 | 1.879 | 1.904 | 1.8915 | 205.9721 | 0.32829
2020-02-14 | 14:06 | 5.733 | 5.829 | 5.781 | 616.4235 | 1.841 | 1.872 | 1.8565 | 202.1608 | 0.32796
2020-02-14 | 14:06 | 5.611 | 5.709 | 5.66 | 603.5214 | 1.798 | 1.833 | 1.8155 | 197.6962 | 0.32757
2020-02-14 | 14:06 | 5.463 | 5.584 | 5.5235 | 588.9665 | 1.748 | 1.79 | 1.769 | 192.6326 | 0.32707
2020-02-14 | 14:06 | 5.349 | 5.433 | 5.391 | 574.8381 | 1.712 | 1.737 | 1.7245 | 187.7868 | 0.32668
2020-02-14 | 14:06 | 5.296 | 5.337 | 5.3165 | 566.8943 | 1.694 | 1.707 | 1.7005 | 185.1734 | 0.32665
2020-02-14 | 14:06 | 5.241 | 5.283 | 5.262 | 561.083 | 1.677 | 1.691 | 1.684 | 183.3767 | 0.32683
2020-02-14 | 14:06 | 5.202 | 5.233 | 5.2175 | 556.338 | 1.665 | 1.673 | 1.669 | 181.7433 | 0.32668
2020-02-14 | 14:06 | 5.16 | 5.195 | 5.1775 | 552.0728 | 1.652 | 1.662 | 1.657 | 180.4365 | 0.32683
2020-02-14 | 14:06 | 5.094 | 5.148 | 5.121 | 546.0483 | 1.628 | 1.648 | 1.638 | 178.3676 | 0.32665
2020-02-14 | 14:06 | 5.054 | 5.082 | 5.068 | 540.3969 | 1.617 | 1.625 | 1.621 | 176.5164 | 0.32664
2020-02-14 | 14:06 | 5.054 | 5.083 | 5.0685 | 540.4502 | 1.616 | 1.628 | 1.622 | 176.6253 | 0.32681
2020-02-14 | 14:06 | 5.096 | 5.145 | 5.1205 | 545.9949 | 1.634 | 1.65 | 1.642 | 178.8031 | 0.32748
2020-02-14 | 14:06 | 5.154 | 5.183 | 5.1685 | 551.1131 | 1.653 | 1.664 | 1.6585 | 180.5999 | 0.3277
2020-02-14 | 14:07 | 5.19 | 5.228 | 5.209 | 555.4316 | 1.666 | 1.681 | 1.6735 | 182.2333 | 0.32809
2020-02-14 | 14:07 | 5.236 | 5.25 | 5.243 | 559.057 | 1.682 | 1.688 | 1.685 | 183.4856 | 0.32821
2020-02-14 | 14:07 | 5.245 | 5.247 | 5.246 | 559.3769 | 1.685 | 1.686 | 1.6855 | 183.54 | 0.32812
2020-02-14 | 14:07 | 5.244 | 5.249 | 5.2465 | 559.4302 | 1.686 | 1.687 | 1.6865 | 183.6489 | 0.32828
2020-02-14 | 14:07 | 5.252 | 5.273 | 5.2625 | 561.1363 | 1.688 | 1.695 | 1.6915 | 184.1934 | 0.32825
2020-02-14 | 14:07 | 5.275 | 5.281 | 5.278 | 562.789 | 1.697 | 1.7 | 1.6985 | 184.9556 | 0.32864
2020-02-14 | 14:07 | 5.283 | 5.302 | 5.2925 | 564.3352 | 1.7 | 1.707 | 1.7035 | 185.5001 | 0.32871
2020-02-14 | 14:07 | 5.306 | 5.312 | 5.309 | 566.0945 | 1.708 | 1.711 | 1.7095 | 186.1534 | 0.32884
2020-02-14 | 14:07 | 5.307 | 5.311 | 5.309 | 566.0945 | 1.709 | 1.71 | 1.7095 | 186.1534 | 0.32884
2020-02-14 | 14:07 | 5.313 | 5.344 | 5.3285 | 568.1738 | 1.71 | 1.723 | 1.7165 | 186.9157 | 0.32898
2020-02-14 | 14:07 | 5.356 | 5.387 | 5.3715 | 572.7589 | 1.726 | 1.738 | 1.732 | 188.6035 | 0.32929
2020-02-14 | 14:08 | 5.39 | 5.418 | 5.404 | 576.2243 | 1.737 | 1.748 | 1.7425 | 189.7469 | 0.32929
2020-02-14 | 14:08 | 5.429 | 5.468 | 5.4485 | 580.9693 | 1.751 | 1.765 | 1.758 | 191.4348 | 0.32951
2020-02-14 | 14:08 | 5.479 | 5.522 | 5.5005 | 586.514 | 1.769 | 1.785 | 1.777 | 193.5038 | 0.32992
2020-02-14 | 14:08 | 5.532 | 5.569 | 5.5505 | 591.8455 | 1.788 | 1.801 | 1.7945 | 195.4094 | 0.33017
2020-02-14 | 14:08 | 5.578 | 5.617 | 5.5975 | 596.8571 | 1.803 | 1.818 | 1.8105 | 197.1517 | 0.33032
2020-02-14 | 14:08 | 5.629 | 5.673 | 5.651 | 602.5617 | 1.82 | 1.836 | 1.828 | 199.0573 | 0.33035
2020-02-14 | 14:08 | 5.684 | 5.716 | 5.7 | 607.7866 | 1.84 | 1.852 | 1.846 | 201.0174 | 0.33074
2020-02-14 | 14:08 | 5.719 | 5.742 | 5.7305 | 611.0388 | 1.854 | 1.86 | 1.857 | 202.2152 | 0.33094
2020-02-14 | 14:08 | 5.748 | 5.76 | 5.754 | 613.5445 | 1.861 | 1.866 | 1.8635 | 202.923 | 0.33074
2020-02-14 | 14:08 | 5.751 | 5.759 | 5.755 | 613.6512 | 1.862 | 1.865 | 1.8635 | 202.923 | 0.33068
2020-02-14 | 14:08 | 5.752 | 5.77 | 5.761 | 614.291 | 1.862 | 1.868 | 1.865 | 203.0864 | 0.3306
2020-02-14 | 14:08 | 5.777 | 5.8 | 5.7885 | 617.2233 | 1.87 | 1.878 | 1.874 | 204.0664 | 0.33062
2020-02-14 | 14:09 | 5.804 | 5.813 | 5.8085 | 619.3558 | 1.88 | 1.882 | 1.881 | 204.8287 | 0.33071
2020-02-14 | 14:09 | 5.817 | 5.837 | 5.827 | 621.3285 | 1.883 | 1.89 | 1.8865 | 205.4276 | 0.33063
2020-02-14 | 14:09 | 5.842 | 5.87 | 5.856 | 624.4207 | 1.893 | 1.9 | 1.8965 | 206.5165 | 0.33073
2020-02-14 | 14:09 | 5.874 | 5.887 | 5.8805 | 627.0331 | 1.903 | 1.907 | 1.905 | 207.4421 | 0.33083
2020-02-14 | 14:09 | 5.888 | 5.896 | 5.892 | 628.2594 | 1.907 | 1.909 | 1.908 | 207.7688 | 0.33071
2020-02-14 | 14:09 | 5.895 | 5.898 | 5.8965 | 628.7392 | 1.909 | 1.911 | 1.91 | 207.9866 | 0.3308
2020-02-14 | 14:09 | 5.89 | 5.893 | 5.8915 | 628.2061 | 1.907 | 1.908 | 1.9075 | 207.7144 | 0.33065
2020-02-14 | 14:09 | 5.89 | 5.903 | 5.8965 | 628.7392 | 1.908 | 1.912 | 1.91 | 207.9866 | 0.3308
2020-02-14 | 14:09 | 5.907 | 5.926 | 5.9165 | 630.8718 | 1.914 | 1.92 | 1.917 | 208.7488 | 0.33089
2020-02-14 | 14:09 | 5.93 | 5.943 | 5.9365 | 633.0044 | 1.92 | 1.925 | 1.9225 | 209.3478 | 0.33072
2020-02-14 | 14:09 | 5.946 | 5.951 | 5.9485 | 634.2839 | 1.926 | 1.928 | 1.927 | 209.8378 | 0.33083
2020-02-14 | 14:09 | 5.944 | 5.95 | 5.947 | 634.124 | 1.924 | 1.927 | 1.9255 | 209.6744 | 0.33065
2020-02-14 | 14:09 | 5.941 | 5.944 | 5.9425 | 633.6442 | 1.923 | 1.924 | 1.9235 | 209.4567 | 0.33056
2020-02-14 | 14:10 | 5.947 | 5.958 | 5.9525 | 634.7104 | 1.926 | 1.931 | 1.9285 | 210.0011 | 0.33086
2020-02-14 | 14:10 | 5.961 | 5.965 | 5.963 | 635.8301 | 1.93 | 1.931 | 1.9305 | 210.2189 | 0.33062
2020-02-14 | 14:10 | 5.952 | 5.963 | 5.9575 | 635.2436 | 1.925 | 1.93 | 1.9275 | 209.8922 | 0.33041
2020-02-14 | 14:10 | 5.949 | 5.954 | 5.9515 | 634.6038 | 1.925 | 1.927 | 1.926 | 209.7289 | 0.33049
2020-02-14 | 14:10 | 5.956 | 5.96 | 5.958 | 635.2969 | 1.926 | 1.929 | 1.9275 | 209.8922 | 0.33038
2020-02-14 | 14:10 | 5.953 | 5.958 | 5.9555 | 635.0303 | 1.925 | 1.928 | 1.9265 | 209.7833 | 0.33035
2020-02-14 | 14:10 | 5.955 | 5.966 | 5.9605 | 635.5635 | 1.927 | 1.93 | 1.9285 | 210.0011 | 0.33042
2020-02-14 | 14:10 | 5.968 | 5.97 | 5.969 | 636.4698 | 1.931 | 1.932 | 1.9315 | 210.3278 | 0.33046
2020-02-14 | 14:10 | 5.966 | 5.969 | 5.9675 | 636.3099 | 1.929 | 1.931 | 1.93 | 210.1645 | 0.33029
2020-02-14 | 14:10 | 5.955 | 5.965 | 5.96 | 635.5102 | 1.926 | 1.928 | 1.927 | 209.8378 | 0.33019
2020-02-14 | 14:10 | 5.929 | 5.95 | 5.9395 | 633.3243 | 1.916 | 1.924 | 1.92 | 209.0755 | 0.33012
2020-02-14 | 14:11 | 5.899 | 5.922 | 5.9105 | 630.232 | 1.905 | 1.914 | 1.9095 | 207.9321 | 0.32993
2020-02-14 | 14:11 | 5.865 | 5.892 | 5.8785 | 626.8199 | 1.894 | 1.906 | 1.9 | 206.8977 | 0.33008
2020-02-14 | 14:11 | 5.845 | 5.86 | 5.8525 | 624.0475 | 1.888 | 1.892 | 1.89 | 205.8087 | 0.3298
2020-02-14 | 14:11 | 5.822 | 5.842 | 5.832 | 621.8616 | 1.88 | 1.887 | 1.8835 | 205.1009 | 0.32982
2020-02-14 | 14:11 | 5.774 | 5.814 | 5.794 | 617.8097 | 1.864 | 1.878 | 1.871 | 203.7397 | 0.32978
2020-02-14 | 14:11 | 5.724 | 5.765 | 5.7445 | 612.5316 | 1.846 | 1.86 | 1.853 | 201.7797 | 0.32942
2020-02-14 | 14:11 | 5.649 | 5.71 | 5.6795 | 605.6007 | 1.82 | 1.841 | 1.8305 | 199.3296 | 0.32914
2020-02-14 | 14:11 | 5.582 | 5.63 | 5.606 | 597.7634 | 1.798 | 1.813 | 1.8055 | 196.6072 | 0.3289
2020-02-14 | 14:11 | 5.582 | 5.618 | 5.6 | 597.1236 | 1.798 | 1.814 | 1.806 | 196.6617 | 0.32935
2020-02-14 | 14:11 | 5.633 | 5.692 | 5.6625 | 603.788 | 1.816 | 1.838 | 1.827 | 198.9484 | 0.3295
2020-02-14 | 14:11 | 5.701 | 5.723 | 5.712 | 609.0661 | 1.843 | 1.849 | 1.846 | 201.0174 | 0.33004
2020-02-14 | 14:11 | 5.727 | 5.753 | 5.74 | 612.0517 | 1.852 | 1.861 | 1.8565 | 202.1608 | 0.3303
2020-02-14 | 14:11 | 5.749 | 5.761 | 5.755 | 613.6512 | 1.857 | 1.863 | 1.86 | 202.5419 | 0.33006
2020-02-14 | 14:12 | 5.706 | 5.741 | 5.7235 | 610.2924 | 1.842 | 1.855 | 1.8485 | 201.2896 | 0.32982
2020-02-14 | 14:12 | 5.686 | 5.699 | 5.6925 | 606.9869 | 1.836 | 1.84 | 1.838 | 200.1463 | 0.32974
2020-02-14 | 14:12 | 5.681 | 5.685 | 5.683 | 605.9739 | 1.835 | 1.837 | 1.836 | 199.9285 | 0.32993
2020-02-14 | 14:12 | 5.683 | 5.692 | 5.6875 | 606.4537 | 1.836 | 1.838 | 1.837 | 200.0374 | 0.32985
2020-02-14 | 14:12 | 5.695 | 5.705 | 5.7 | 607.7866 | 1.841 | 1.844 | 1.8425 | 200.6363 | 0.33011
2020-02-14 | 14:12 | 5.708 | 5.735 | 5.7215 | 610.0791 | 1.846 | 1.855 | 1.8505 | 201.5074 | 0.3303
2020-02-14 | 14:12 | 5.742 | 5.773 | 5.7575 | 613.9178 | 1.858 | 1.868 | 1.863 | 202.8686 | 0.33045
2020-02-14 | 14:12 | 5.777 | 5.785 | 5.781 | 616.4235 | 1.87 | 1.872 | 1.871 | 203.7397 | 0.33052
2020-02-14 | 14:12 | 5.786 | 5.794 | 5.79 | 617.3832 | 1.872 | 1.875 | 1.8735 | 204.012 | 0.33045
2020-02-14 | 14:12 | 5.796 | 5.809 | 5.8025 | 618.7161 | 1.877 | 1.881 | 1.879 | 204.6109 | 0.3307
2020-02-14 | 14:12 | 5.811 | 5.824 | 5.8175 | 620.3155 | 1.883 | 1.886 | 1.8845 | 205.2098 | 0.33082
2020-02-14 | 14:13 | 5.828 | 5.847 | 5.8375 | 622.4481 | 1.888 | 1.895 | 1.8915 | 205.9721 | 0.33091
2020-02-14 | 14:13 | 5.852 | 5.862 | 5.857 | 624.5274 | 1.897 | 1.9 | 1.8985 | 206.7343 | 0.33103
2020-02-14 | 14:13 | 5.864 | 5.865 | 5.8645 | 625.3271 | 1.899 | 1.902 | 1.9005 | 206.9521 | 0.33095
2020-02-14 | 14:13 | 5.861 | 5.865 | 5.863 | 625.1671 | 1.9 | 1.901 | 1.9005 | 206.9521 | 0.33103
2020-02-14 | 14:13 | 5.859 | 5.861 | 5.86 | 624.8472 | 1.898 | 1.9 | 1.899 | 206.7888 | 0.33094
2020-02-14 | 14:13 | 5.86 | 5.862 | 5.861 | 624.9539 | 1.899 | 1.902 | 1.9005 | 206.9521 | 0.33115
2020-02-14 | 14:13 | 5.858 | 5.862 | 5.86 | 624.8472 | 1.899 | 1.9 | 1.8995 | 206.8432 | 0.33103
2020-02-14 | 14:13 | 5.849 | 5.857 | 5.853 | 624.1008 | 1.895 | 1.899 | 1.897 | 206.571 | 0.33099
2020-02-14 | 14:13 | 5.833 | 5.846 | 5.8395 | 622.6613 | 1.89 | 1.895 | 1.8925 | 206.081 | 0.33097
2020-02-14 | 14:13 | 5.816 | 5.83 | 5.823 | 620.902 | 1.885 | 1.889 | 1.887 | 205.482 | 0.33094
2020-02-14 | 14:13 | 5.791 | 5.811 | 5.801 | 618.5561 | 1.877 | 1.883 | 1.88 | 204.7198 | 0.33096
2020-02-14 | 14:13 | 5.776 | 5.789 | 5.7825 | 616.5835 | 1.871 | 1.875 | 1.873 | 203.9575 | 0.33079
2020-02-14 | 14:13 | 5.744 | 5.77 | 5.757 | 613.8644 | 1.859 | 1.869 | 1.864 | 202.9775 | 0.33066
2020-02-14 | 14:14 | 5.704 | 5.738 | 5.721 | 610.0258 | 1.845 | 1.858 | 1.8515 | 201.6163 | 0.3305
2020-02-14 | 14:14 | 5.657 | 5.695 | 5.676 | 605.2275 | 1.83 | 1.843 | 1.8365 | 199.9829 | 0.33043
2020-02-14 | 14:14 | 5.605 | 5.645 | 5.625 | 599.7894 | 1.812 | 1.824 | 1.818 | 197.9684 | 0.33006
2020-02-14 | 14:14 | 5.569 | 5.596 | 5.5825 | 595.2576 | 1.801 | 1.809 | 1.805 | 196.5528 | 0.3302
2020-02-14 | 14:14 | 5.55 | 5.566 | 5.558 | 592.6452 | 1.793 | 1.798 | 1.7955 | 195.5183 | 0.32991
2020-02-14 | 14:14 | 5.52 | 5.546 | 5.533 | 589.9795 | 1.784 | 1.792 | 1.788 | 194.7016 | 0.33001
2020-02-14 | 14:14 | 5.461 | 5.512 | 5.4865 | 585.0212 | 1.763 | 1.78 | 1.7715 | 192.9048 | 0.32974
2020-02-14 | 14:14 | 5.378 | 5.447 | 5.4125 | 577.1307 | 1.735 | 1.757 | 1.746 | 190.1281 | 0.32944
2020-02-14 | 14:14 | 5.28 | 5.359 | 5.3195 | 567.2142 | 1.701 | 1.728 | 1.7145 | 186.6979 | 0.32915
2020-02-14 | 14:14 | 5.194 | 5.26 | 5.227 | 557.3509 | 1.672 | 1.695 | 1.6835 | 183.3222 | 0.32892
2020-02-14 | 14:14 | 5.128 | 5.181 | 5.1545 | 549.6203 | 1.652 | 1.669 | 1.6605 | 180.8177 | 0.32899
2020-02-14 | 14:15 | 5.068 | 5.116 | 5.092 | 542.956 | 1.632 | 1.648 | 1.64 | 178.5853 | 0.32891
2020-02-14 | 14:15 | 4.971 | 5.05 | 5.0105 | 534.2657 | 1.6 | 1.627 | 1.6135 | 175.6997 | 0.32886
2020-02-14 | 14:15 | 4.871 | 4.95 | 4.9105 | 523.6028 | 1.567 | 1.592 | 1.5795 | 171.9973 | 0.32849
2020-02-14 | 14:15 | 4.796 | 4.856 | 4.826 | 514.5926 | 1.543 | 1.562 | 1.5525 | 169.0572 | 0.32853
2020-02-14 | 14:15 | 4.759 | 4.784 | 4.7715 | 508.7813 | 1.531 | 1.54 | 1.5355 | 167.206 | 0.32864
2020-02-14 | 14:15 | 4.758 | 4.789 | 4.7735 | 508.9946 | 1.531 | 1.543 | 1.537 | 167.3693 | 0.32882
2020-02-14 | 14:15 | 4.802 | 4.855 | 4.8285 | 514.8592 | 1.548 | 1.568 | 1.558 | 169.6561 | 0.32952
2020-02-14 | 14:15 | 4.866 | 4.888 | 4.877 | 520.0307 | 1.571 | 1.579 | 1.575 | 171.5073 | 0.3298
2020-02-14 | 14:15 | 4.89 | 4.907 | 4.8985 | 522.3232 | 1.581 | 1.587 | 1.584 | 172.4873 | 0.33023
2020-02-14 | 14:15 | 4.914 | 4.954 | 4.934 | 526.1086 | 1.59 | 1.604 | 1.597 | 173.9029 | 0.33055
2020-02-14 | 14:15 | 4.969 | 5.05 | 5.0095 | 534.1591 | 1.609 | 1.639 | 1.624 | 176.843 | 0.33107
2020-02-14 | 14:15 | 5.076 | 5.181 | 5.1285 | 546.848 | 1.649 | 1.685 | 1.667 | 181.5255 | 0.33195
2020-02-14 | 14:15 | 5.199 | 5.258 | 5.2285 | 557.5109 | 1.693 | 1.711 | 1.702 | 185.3367 | 0.33244
2020-02-14 | 14:16 | 5.269 | 5.305 | 5.287 | 563.7487 | 1.715 | 1.726 | 1.7205 | 187.3513 | 0.33233
2020-02-14 | 14:16 | 5.304 | 5.309 | 5.3065 | 565.828 | 1.726 | 1.728 | 1.727 | 188.0591 | 0.33236
2020-02-14 | 14:16 | 5.258 | 5.3 | 5.279 | 562.8957 | 1.71 | 1.725 | 1.7175 | 187.0246 | 0.33225
2020-02-14 | 14:16 | 5.179 | 5.242 | 5.2105 | 555.5916 | 1.683 | 1.704 | 1.6935 | 184.4111 | 0.33192
2020-02-14 | 14:16 | 5.137 | 5.168 | 5.1525 | 549.4071 | 1.668 | 1.678 | 1.673 | 182.1788 | 0.33159
2020-02-14 | 14:16 | 5.095 | 5.132 | 5.1135 | 545.2485 | 1.653 | 1.666 | 1.6595 | 180.7088 | 0.33142
2020-02-14 | 14:16 | 5.015 | 5.079 | 5.047 | 538.1577 | 1.626 | 1.649 | 1.6375 | 178.3131 | 0.33134
2020-02-14 | 14:16 | 4.963 | 5 | 4.9815 | 531.1735 | 1.608 | 1.621 | 1.6145 | 175.8086 | 0.33098
2020-02-14 | 14:16 | 4.9 | 4.952 | 4.926 | 525.2556 | 1.585 | 1.605 | 1.595 | 173.6851 | 0.33067
2020-02-14 | 14:16 | 4.799 | 4.883 | 4.841 | 516.1921 | 1.55 | 1.58 | 1.565 | 170.4183 | 0.33015
2020-02-14 | 14:16 | 4.659 | 4.773 | 4.716 | 502.8634 | 1.503 | 1.542 | 1.5225 | 165.7904 | 0.32969
2020-02-14 | 14:17 | 4.558 | 4.631 | 4.5945 | 489.908 | 1.47 | 1.493 | 1.4815 | 161.3257 | 0.3293
2020-02-14 | 14:17 | 4.554 | 4.577 | 4.5655 | 486.8157 | 1.469 | 1.477 | 1.473 | 160.4001 | 0.32949
2020-02-14 | 14:17 | 4.586 | 4.626 | 4.606 | 491.1342 | 1.48 | 1.494 | 1.487 | 161.9246 | 0.3297
2020-02-14 | 14:17 | 4.642 | 4.76 | 4.701 | 501.264 | 1.5 | 1.542 | 1.521 | 165.627 | 0.33042
2020-02-14 | 14:17 | 4.8 | 4.958 | 4.879 | 520.244 | 1.556 | 1.611 | 1.5835 | 172.4329 | 0.33145
2020-02-14 | 14:17 | 4.987 | 5.07 | 5.0285 | 536.185 | 1.621 | 1.65 | 1.6355 | 178.0953 | 0.33215
2020-02-14 | 14:17 | 5.088 | 5.151 | 5.1195 | 545.8883 | 1.656 | 1.679 | 1.6675 | 181.5799 | 0.33263
2020-02-14 | 14:17 | 5.165 | 5.198 | 5.1815 | 552.4993 | 1.682 | 1.692 | 1.687 | 183.7033 | 0.3325
2020-02-14 | 14:17 | 5.205 | 5.25 | 5.2275 | 557.4043 | 1.695 | 1.709 | 1.702 | 185.3367 | 0.3325
2020-02-14 | 14:17 | 5.264 | 5.314 | 5.289 | 563.962 | 1.714 | 1.731 | 1.7225 | 187.5691 | 0.33259
2020-02-14 | 14:17 | 5.326 | 5.382 | 5.354 | 570.8929 | 1.734 | 1.754 | 1.744 | 189.9103 | 0.33265
2020-02-14 | 14:17 | 5.398 | 5.463 | 5.4305 | 579.05 | 1.759 | 1.783 | 1.771 | 192.8504 | 0.33305
2020-02-14 | 14:17 | 5.478 | 5.526 | 5.502 | 586.674 | 1.788 | 1.802 | 1.795 | 195.4638 | 0.33317
2020-02-14 | 14:18 | 5.535 | 5.569 | 5.552 | 592.0054 | 1.805 | 1.818 | 1.8115 | 197.2606 | 0.33321
2020-02-14 | 14:18 | 5.577 | 5.607 | 5.592 | 596.2706 | 1.82 | 1.831 | 1.8255 | 198.7851 | 0.33338
2020-02-14 | 14:18 | 5.615 | 5.646 | 5.6305 | 600.3758 | 1.833 | 1.843 | 1.838 | 200.1463 | 0.33337
2020-02-14 | 14:18 | 5.652 | 5.671 | 5.6615 | 603.6813 | 1.843 | 1.85 | 1.8465 | 201.0719 | 0.33308
2020-02-14 | 14:18 | 5.676 | 5.69 | 5.683 | 605.9739 | 1.852 | 1.856 | 1.854 | 201.8886 | 0.33316
2020-02-14 | 14:18 | 5.693 | 5.706 | 5.6995 | 607.7333 | 1.858 | 1.861 | 1.8595 | 202.4875 | 0.33318
2020-02-14 | 14:18 | 5.707 | 5.715 | 5.711 | 608.9595 | 1.863 | 1.864 | 1.8635 | 202.923 | 0.33323
2020-02-14 | 14:18 | 5.717 | 5.721 | 5.719 | 609.8125 | 1.865 | 1.867 | 1.866 | 203.1953 | 0.33321
2020-02-14 | 14:18 | 5.723 | 5.728 | 5.7255 | 610.5056 | 1.868 | 1.869 | 1.8685 | 203.4675 | 0.33328
2020-02-14 | 14:18 | 5.731 | 5.739 | 5.735 | 611.5186 | 1.869 | 1.871 | 1.87 | 203.6308 | 0.33299
2020-02-14 | 14:18 | 5.741 | 5.747 | 5.744 | 612.4783 | 1.872 | 1.874 | 1.873 | 203.9575 | 0.333
2020-02-14 | 14:19 | 5.748 | 5.751 | 5.7495 | 613.0647 | 1.873 | 1.875 | 1.874 | 204.0664 | 0.33286
2020-02-14 | 14:19 | 5.753 | 5.757 | 5.755 | 613.6512 | 1.875 | 1.876 | 1.8755 | 204.2298 | 0.33281
2020-02-14 | 14:19 | 5.754 | 5.757 | 5.7555 | 613.7045 | 1.875 | 1.877 | 1.876 | 204.2842 | 0.33287
2020-02-14 | 14:19 | 5.755 | 5.757 | 5.756 | 613.7578 | 1.875 | 1.876 | 1.8755 | 204.2298 | 0.33275
2020-02-14 | 14:19 | 5.751 | 5.757 | 5.754 | 613.5445 | 1.874 | 1.875 | 1.8745 | 204.1209 | 0.33269
2020-02-14 | 14:19 | 5.745 | 5.75 | 5.7475 | 612.8515 | 1.87 | 1.874 | 1.872 | 203.8486 | 0.33262
2020-02-14 | 14:19 | 5.743 | 5.745 | 5.744 | 612.4783 | 1.87 | 1.872 | 1.871 | 203.7397 | 0.33265
2020-02-14 | 14:19 | 5.744 | 5.745 | 5.7445 | 612.5316 | 1.87 | 1.871 | 1.8705 | 203.6853 | 0.33253
2020-02-14 | 14:19 | 5.745 | 5.747 | 5.746 | 612.6915 | 1.871 | 1.872 | 1.8715 | 203.7942 | 0.33262
2020-02-14 | 14:19 | 5.744 | 5.747 | 5.7455 | 612.6382 | 1.872 | 1.873 | 1.8725 | 203.9031 | 0.33283
2020-02-14 | 14:19 | 5.743 | 5.744 | 5.7435 | 612.4249 | 1.87 | 1.873 | 1.8715 | 203.7942 | 0.33277
2020-02-14 | 14:19 | 5.735 | 5.741 | 5.738 | 611.8385 | 1.868 | 1.87 | 1.869 | 203.522 | 0.33264
2020-02-14 | 14:19 | 5.727 | 5.734 | 5.7305 | 611.0388 | 1.865 | 1.868 | 1.8665 | 203.2497 | 0.33263
2020-02-14 | 14:20 | 5.719 | 5.725 | 5.722 | 610.1324 | 1.859 | 1.864 | 1.8615 | 202.7053 | 0.33223
2020-02-14 | 14:20 | 5.711 | 5.717 | 5.714 | 609.2794 | 1.847 | 1.858 | 1.8525 | 201.7252 | 0.33109
2020-02-14 | 14:20 | 5.706 | 5.709 | 5.7075 | 608.5863 | 1.841 | 1.845 | 1.843 | 200.6907 | 0.32977
2020-02-14 | 14:20 | 5.697 | 5.704 | 5.7005 | 607.8399 | 1.836 | 1.841 | 1.8385 | 200.2007 | 0.32936
2020-02-14 | 14:20 | 5.686 | 5.693 | 5.6895 | 606.667 | 1.833 | 1.838 | 1.8355 | 199.874 | 0.32946
2020-02-14 | 14:20 | 5.677 | 5.684 | 5.6805 | 605.7073 | 1.834 | 1.836 | 1.835 | 199.8196 | 0.32989
2020-02-14 | 14:20 | 5.663 | 5.675 | 5.669 | 604.4811 | 1.828 | 1.833 | 1.8305 | 199.3296 | 0.32975
2020-02-14 | 14:20 | 5.645 | 5.661 | 5.653 | 602.775 | 1.812 | 1.826 | 1.819 | 198.0773 | 0.32861
2020-02-14 | 14:20 | 5.628 | 5.641 | 5.6345 | 600.8024 | 1.791 | 1.806 | 1.7985 | 195.845 | 0.32597
2020-02-14 | 14:20 | 5.618 | 5.626 | 5.622 | 599.4695 | 1.784 | 1.791 | 1.7875 | 194.6471 | 0.3247
2020-02-14 | 14:20 | 5.604 | 5.614 | 5.609 | 598.0833 | 1.781 | 1.785 | 1.783 | 194.1571 | 0.32463
2020-02-14 | 14:21 | 5.589 | 5.602 | 5.5955 | 596.6438 | 1.776 | 1.779 | 1.7775 | 193.5582 | 0.32441
2020-02-14 | 14:21 | 5.569 | 5.585 | 5.577 | 594.6712 | 1.766 | 1.774 | 1.77 | 192.7415 | 0.32411
2020-02-14 | 14:21 | 5.544 | 5.563 | 5.5535 | 592.1654 | 1.725 | 1.762 | 1.7435 | 189.8558 | 0.32061
2020-02-14 | 14:21 | 5.526 | 5.539 | 5.5325 | 589.9262 | 1.73 | 1.744 | 1.737 | 189.148 | 0.32063
2020-02-14 | 14:21 | 5.507 | 5.524 | 5.5155 | 588.1135 | 1.736 | 1.745 | 1.7405 | 189.5291 | 0.32227
2020-02-14 | 14:21 | 5.481 | 5.501 | 5.491 | 585.5011 | 1.732 | 1.736 | 1.734 | 188.8213 | 0.3225
2020-02-14 | 14:21 | 5.453 | 5.476 | 5.4645 | 582.6754 | 1.72 | 1.731 | 1.7255 | 187.8957 | 0.32247
2020-02-14 | 14:21 | 5.42 | 5.448 | 5.434 | 579.4232 | 1.717 | 1.722 | 1.7195 | 187.2424 | 0.32315
2020-02-14 | 14:21 | 5.388 | 5.413 | 5.4005 | 575.8511 | 1.711 | 1.715 | 1.713 | 186.5346 | 0.32393
2020-02-14 | 14:21 | 5.358 | 5.382 | 5.37 | 572.5989 | 1.705 | 1.71 | 1.7075 | 185.9357 | 0.32472
2020-02-14 | 14:21 | 5.325 | 5.351 | 5.338 | 569.1868 | 1.693 | 1.703 | 1.698 | 184.9012 | 0.32485
2020-02-14 | 14:21 | 5.288 | 5.318 | 5.303 | 565.4548 | 1.681 | 1.691 | 1.686 | 183.5944 | 0.32468
2020-02-14 | 14:21 | 5.255 | 5.281 | 5.268 | 561.7227 | 1.671 | 1.678 | 1.6745 | 182.3422 | 0.32461
2020-02-14 | 14:22 | 5.224 | 5.25 | 5.237 | 558.4172 | 1.661 | 1.669 | 1.665 | 181.3077 | 0.32468
2020-02-14 | 14:22 | 5.19 | 5.216 | 5.203 | 554.7918 | 1.65 | 1.659 | 1.6545 | 180.1643 | 0.32474
2020-02-14 | 14:22 | 5.17 | 5.186 | 5.178 | 552.1261 | 1.643 | 1.649 | 1.646 | 179.2387 | 0.32463
2020-02-14 | 14:22 | 5.138 | 5.164 | 5.151 | 549.2471 | 1.637 | 1.644 | 1.6405 | 178.6398 | 0.32524
2020-02-14 | 14:22 | 5.102 | 5.131 | 5.1165 | 545.5684 | 1.624 | 1.633 | 1.6285 | 177.3331 | 0.32504
2020-02-14 | 14:22 | 5.067 | 5.094 | 5.0805 | 541.7298 | 1.615 | 1.621 | 1.618 | 176.1897 | 0.32524
2020-02-14 | 14:22 | 5.022 | 5.058 | 5.04 | 537.4113 | 1.604 | 1.614 | 1.609 | 175.2096 | 0.32603
2020-02-14 | 14:22 | 4.982 | 5.014 | 4.998 | 532.9329 | 1.59 | 1.601 | 1.5955 | 173.7396 | 0.32601
2020-02-14 | 14:22 | 4.946 | 4.974 | 4.96 | 528.8809 | 1.57 | 1.587 | 1.5785 | 171.8884 | 0.325
2020-02-14 | 14:22 | 4.901 | 4.939 | 4.92 | 524.6158 | 1.548 | 1.571 | 1.5595 | 169.8194 | 0.3237
2020-02-14 | 14:22 | 4.853 | 4.892 | 4.8725 | 519.5509 | 1.529 | 1.542 | 1.5355 | 167.206 | 0.32183
2020-02-14 | 14:22 | 4.8 | 4.845 | 4.8225 | 514.2194 | 1.51 | 1.527 | 1.5185 | 165.3548 | 0.32156
2020-02-14 | 14:23 | 4.735 | 4.787 | 4.761 | 507.6617 | 1.488 | 1.506 | 1.497 | 163.0136 | 0.32111
2020-02-14 | 14:23 | 4.699 | 4.725 | 4.712 | 502.4369 | 1.478 | 1.486 | 1.482 | 161.3802 | 0.32119
2020-02-14 | 14:23 | 4.696 | 4.699 | 4.6975 | 500.8908 | 1.478 | 1.479 | 1.4785 | 160.999 | 0.32143
2020-02-14 | 14:23 | 4.674 | 4.695 | 4.6845 | 499.5046 | 1.47 | 1.477 | 1.4735 | 160.4546 | 0.32123
2020-02-14 | 14:23 | 4.652 | 4.667 | 4.6595 | 496.8389 | 1.463 | 1.469 | 1.466 | 159.6379 | 0.32131
2020-02-14 | 14:23 | 4.649 | 4.654 | 4.6515 | 495.9858 | 1.463 | 1.465 | 1.464 | 159.4201 | 0.32142
2020-02-14 | 14:23 | 4.66 | 4.708 | 4.684 | 499.4513 | 1.466 | 1.486 | 1.476 | 160.7268 | 0.32181
2020-02-14 | 14:23 | 4.724 | 4.782 | 4.753 | 506.8087 | 1.491 | 1.511 | 1.501 | 163.4491 | 0.32251
2020-02-14 | 14:23 | 4.791 | 4.819 | 4.805 | 512.3534 | 1.514 | 1.522 | 1.518 | 165.3003 | 0.32263
2020-02-14 | 14:23 | 4.823 | 4.836 | 4.8295 | 514.9658 | 1.522 | 1.529 | 1.5255 | 166.117 | 0.32258
2020-02-14 | 14:23 | 4.841 | 4.862 | 4.8515 | 517.3117 | 1.531 | 1.539 | 1.535 | 167.1515 | 0.32312
2020-02-14 | 14:24 | 4.869 | 4.884 | 4.8765 | 519.9774 | 1.54 | 1.545 | 1.5425 | 167.9682 | 0.32303
2020-02-14 | 14:24 | 4.887 | 4.896 | 4.8915 | 521.5768 | 1.546 | 1.548 | 1.547 | 168.4582 | 0.32298
2020-02-14 | 14:24 | 4.896 | 4.911 | 4.9035 | 522.8564 | 1.55 | 1.555 | 1.5525 | 169.0572 | 0.32333
2020-02-14 | 14:24 | 4.915 | 4.924 | 4.9195 | 524.5625 | 1.558 | 1.559 | 1.5585 | 169.7105 | 0.32353
2020-02-14 | 14:24 | 4.908 | 4.922 | 4.915 | 524.0826 | 1.554 | 1.558 | 1.556 | 169.4383 | 0.3233
2020-02-14 | 14:24 | 4.892 | 4.905 | 4.8985 | 522.3232 | 1.541 | 1.553 | 1.547 | 168.4582 | 0.32252
2020-02-14 | 14:24 | 4.891 | 4.902 | 4.8965 | 522.11 | 1.541 | 1.543 | 1.542 | 167.9138 | 0.32161
2020-02-14 | 14:24 | 4.904 | 4.928 | 4.916 | 524.1893 | 1.542 | 1.551 | 1.5465 | 168.4038 | 0.32127
2020-02-14 | 14:24 | 4.939 | 4.987 | 4.963 | 529.2008 | 1.554 | 1.568 | 1.561 | 169.9828 | 0.32121
2020-02-14 | 14:24 | 5.002 | 5.044 | 5.023 | 535.5986 | 1.577 | 1.598 | 1.5875 | 172.8684 | 0.32276
2020-02-14 | 14:24 | 5.052 | 5.076 | 5.064 | 539.9704 | 1.602 | 1.61 | 1.606 | 174.883 | 0.32388
2020-02-14 | 14:24 | 5.08 | 5.096 | 5.088 | 542.5295 | 1.61 | 1.625 | 1.6175 | 176.1352 | 0.32466
2020-02-14 | 14:24 | 5.1 | 5.128 | 5.114 | 545.3018 | 1.626 | 1.637 | 1.6315 | 177.6597 | 0.3258
2020-02-14 | 14:25 | 5.137 | 5.178 | 5.1575 | 549.9402 | 1.64 | 1.654 | 1.647 | 179.3476 | 0.32612
2020-02-14 | 14:25 | 5.189 | 5.219 | 5.204 | 554.8985 | 1.66 | 1.679 | 1.6695 | 181.7977 | 0.32762
2020-02-14 | 14:25 | 5.224 | 5.242 | 5.233 | 557.9907 | 1.682 | 1.687 | 1.6845 | 183.4311 | 0.32874
2020-02-14 | 14:25 | 5.244 | 5.25 | 5.247 | 559.4835 | 1.684 | 1.689 | 1.6865 | 183.6489 | 0.32825
2020-02-14 | 14:25 | 5.251 | 5.254 | 5.2525 | 560.07 | 1.683 | 1.69 | 1.6865 | 183.6489 | 0.3279
2020-02-14 | 14:25 | 5.255 | 5.268 | 5.2615 | 561.0297 | 1.675 | 1.681 | 1.678 | 182.7233 | 0.32569
2020-02-14 | 14:25 | 5.271 | 5.277 | 5.274 | 562.3625 | 1.684 | 1.688 | 1.686 | 183.5944 | 0.32647
2020-02-14 | 14:25 | 5.277 | 5.285 | 5.281 | 563.1089 | 1.688 | 1.692 | 1.69 | 184.03 | 0.32681
2020-02-14 | 14:25 | 5.289 | 5.301 | 5.295 | 564.6017 | 1.693 | 1.697 | 1.695 | 184.5745 | 0.32691
2020-02-14 | 14:25 | 5.301 | 5.304 | 5.3025 | 565.4015 | 1.696 | 1.698 | 1.697 | 184.7923 | 0.32683
2020-02-14 | 14:25 | 5.299 | 5.303 | 5.301 | 565.2415 | 1.696 | 1.697 | 1.6965 | 184.7378 | 0.32683
2020-02-14 | 14:26 | 5.284 | 5.298 | 5.291 | 564.1752 | 1.69 | 1.695 | 1.6925 | 184.3023 | 0.32668
2020-02-14 | 14:26 | 5.263 | 5.281 | 5.272 | 562.1493 | 1.683 | 1.689 | 1.686 | 183.5944 | 0.32659
2020-02-14 | 14:26 | 5.235 | 5.256 | 5.2455 | 559.3236 | 1.677 | 1.681 | 1.679 | 182.8322 | 0.32688
2020-02-14 | 14:26 | 5.215 | 5.23 | 5.2225 | 556.8711 | 1.673 | 1.677 | 1.675 | 182.3966 | 0.32754
2020-02-14 | 14:26 | 5.204 | 5.212 | 5.208 | 555.325 | 1.67 | 1.672 | 1.671 | 181.961 | 0.32767
2020-02-14 | 14:26 | 5.195 | 5.202 | 5.1985 | 554.312 | 1.665 | 1.67 | 1.6675 | 181.5799 | 0.32758
2020-02-14 | 14:26 | 5.187 | 5.191 | 5.189 | 553.299 | 1.666 | 1.667 | 1.6665 | 181.471 | 0.32798
2020-02-14 | 14:26 | 5.183 | 5.186 | 5.1845 | 552.8192 | 1.666 | 1.668 | 1.667 | 181.5255 | 0.32836
2020-02-14 | 14:26 | 5.179 | 5.184 | 5.1815 | 552.4993 | 1.665 | 1.667 | 1.666 | 181.4166 | 0.32836
2020-02-14 | 14:26 | 5.17 | 5.179 | 5.1745 | 551.7529 | 1.662 | 1.666 | 1.664 | 181.1988 | 0.32841
2020-02-14 | 14:26 | 5.162 | 5.169 | 5.1655 | 550.7933 | 1.661 | 1.662 | 1.6615 | 180.9266 | 0.32848
2020-02-14 | 14:26 | 5.152 | 5.162 | 5.157 | 549.8869 | 1.658 | 1.661 | 1.6595 | 180.7088 | 0.32863
2020-02-14 | 14:26 | 5.146 | 5.15 | 5.148 | 548.9272 | 1.656 | 1.657 | 1.6565 | 180.3821 | 0.32861
2020-02-14 | 14:27 | 5.141 | 5.147 | 5.144 | 548.5007 | 1.654 | 1.657 | 1.6555 | 180.2732 | 0.32867
2020-02-14 | 14:27 | 5.131 | 5.139 | 5.135 | 547.5411 | 1.65 | 1.654 | 1.652 | 179.8921 | 0.32855
2020-02-14 | 14:27 | 5.12 | 5.129 | 5.1245 | 546.4215 | 1.638 | 1.65 | 1.644 | 179.0209 | 0.32762
2020-02-14 | 14:27 | 5.109 | 5.118 | 5.1135 | 545.2485 | 1.639 | 1.643 | 1.641 | 178.6942 | 0.32773
2020-02-14 | 14:27 | 5.091 | 5.106 | 5.0985 | 543.6491 | 1.633 | 1.639 | 1.636 | 178.1498 | 0.32769
2020-02-14 | 14:27 | 5.085 | 5.089 | 5.087 | 542.4229 | 1.633 | 1.636 | 1.6345 | 177.9864 | 0.32813
2020-02-14 | 14:27 | 5.092 | 5.117 | 5.1045 | 544.2889 | 1.637 | 1.646 | 1.6415 | 178.7487 | 0.32841
2020-02-14 | 14:27 | 5.12 | 5.142 | 5.131 | 547.1145 | 1.647 | 1.654 | 1.6505 | 179.7287 | 0.3285
2020-02-14 | 14:27 | 5.148 | 5.164 | 5.156 | 549.7803 | 1.656 | 1.662 | 1.659 | 180.6543 | 0.32859
2020-02-14 | 14:27 | 5.166 | 5.183 | 5.1745 | 551.7529 | 1.664 | 1.669 | 1.6665 | 181.471 | 0.3289
2020-02-14 | 14:27 | 5.187 | 5.208 | 5.1975 | 554.2054 | 1.672 | 1.677 | 1.6745 | 182.3422 | 0.32902
2020-02-14 | 14:28 | 5.212 | 5.226 | 5.219 | 556.4979 | 1.679 | 1.686 | 1.6825 | 183.2133 | 0.32923
2020-02-14 | 14:28 | 5.229 | 5.253 | 5.241 | 558.8438 | 1.686 | 1.693 | 1.6895 | 183.9756 | 0.32921
2020-02-14 | 14:28 | 5.259 | 5.281 | 5.27 | 561.936 | 1.695 | 1.702 | 1.6985 | 184.9556 | 0.32914
2020-02-14 | 14:28 | 5.286 | 5.296 | 5.291 | 564.1752 | 1.705 | 1.708 | 1.7065 | 185.8268 | 0.32938
2020-02-14 | 14:28 | 5.298 | 5.305 | 5.3015 | 565.2948 | 1.708 | 1.711 | 1.7095 | 186.1534 | 0.3293
2020-02-14 | 14:28 | 5.307 | 5.317 | 5.312 | 566.4144 | 1.711 | 1.712 | 1.7115 | 186.3712 | 0.32904
2020-02-14 | 14:28 | 5.321 | 5.332 | 5.3265 | 567.9606 | 1.709 | 1.711 | 1.71 | 186.2079 | 0.32785
2020-02-14 | 14:28 | 5.336 | 5.352 | 5.344 | 569.8266 | 1.705 | 1.71 | 1.7075 | 185.9357 | 0.3263
2020-02-14 | 14:28 | 5.357 | 5.373 | 5.365 | 572.0658 | 1.706 | 1.712 | 1.709 | 186.099 | 0.32531
2020-02-14 | 14:28 | 5.375 | 5.392 | 5.3835 | 574.0384 | 1.707 | 1.718 | 1.7125 | 186.4801 | 0.32486
2020-02-14 | 14:28 | 5.396 | 5.407 | 5.4015 | 575.9577 | 1.72 | 1.723 | 1.7215 | 187.4602 | 0.32548
2020-02-14 | 14:28 | 5.408 | 5.414 | 5.411 | 576.9707 | 1.724 | 1.727 | 1.7255 | 187.8957 | 0.32566
2020-02-14 | 14:28 | 5.416 | 5.425 | 5.4205 | 577.9837 | 1.725 | 1.73 | 1.7275 | 188.1135 | 0.32547
2020-02-14 | 14:29 | 5.427 | 5.433 | 5.43 | 578.9967 | 1.721 | 1.728 | 1.7245 | 187.7868 | 0.32433
2020-02-14 | 14:29 | 5.433 | 5.434 | 5.4335 | 579.3699 | 1.724 | 1.728 | 1.726 | 187.9502 | 0.3244
2020-02-14 | 14:29 | 5.43 | 5.433 | 5.4315 | 579.1566 | 1.722 | 1.727 | 1.7245 | 187.7868 | 0.32424
2020-02-14 | 14:29 | 5.429 | 5.431 | 5.43 | 578.9967 | 1.724 | 1.727 | 1.7255 | 187.8957 | 0.32452
2020-02-14 | 14:29 | 5.427 | 5.429 | 5.428 | 578.7834 | 1.722 | 1.728 | 1.725 | 187.8413 | 0.32455
2020-02-14 | 14:29 | 5.424 | 5.428 | 5.426 | 578.5702 | 1.72 | 1.725 | 1.7225 | 187.5691 | 0.32419
2020-02-14 | 14:29 | 5.42 | 5.425 | 5.4225 | 578.197 | 1.726 | 1.727 | 1.7265 | 188.0046 | 0.32516
2020-02-14 | 14:29 | 5.414 | 5.42 | 5.417 | 577.6105 | 1.723 | 1.726 | 1.7245 | 187.7868 | 0.32511
2020-02-14 | 14:29 | 5.405 | 5.411 | 5.408 | 576.6508 | 1.729 | 1.737 | 1.733 | 188.7124 | 0.32726
2020-02-14 | 14:29 | 5.403 | 5.408 | 5.4055 | 576.3843 | 1.729 | 1.736 | 1.7325 | 188.658 | 0.32731
2020-02-14 | 14:29 | 5.399 | 5.405 | 5.402 | 576.0111 | 1.73 | 1.738 | 1.734 | 188.8213 | 0.32781
2020-02-14 | 14:30 | 5.397 | 5.399 | 5.398 | 575.5845 | 1.74 | 1.755 | 1.7475 | 190.2914 | 0.33061
2020-02-14 | 14:30 | 5.397 | 5.401 | 5.399 | 575.6912 | 1.758 | 1.762 | 1.76 | 191.6526 | 0.33291
2020-02-14 | 14:30 | 5.391 | 5.396 | 5.3935 | 575.1047 | 1.758 | 1.761 | 1.7595 | 191.5981 | 0.33315
2020-02-14 | 14:30 | 5.383 | 5.389 | 5.386 | 574.305 | 1.755 | 1.758 | 1.7565 | 191.2714 | 0.33305
2020-02-14 | 14:30 | 5.382 | 5.383 | 5.3825 | 573.9318 | 1.755 | 1.756 | 1.7555 | 191.1625 | 0.33308
2020-02-14 | 14:30 | 5.375 | 5.381 | 5.378 | 573.452 | 1.753 | 1.756 | 1.7545 | 191.0536 | 0.33316
Average |  |  |  |  | 543.5545 |  |  |  | 168.0224 | 0.31022
[5 empty] | W/m^2 |  |  |  | W/m^2

## Error Propagation (top pyra)
Error Propagation - top Pyranometer
Akwatia
uncertainty | relative error [+/- %] | gross value [W/m^2] | divisor | u [W/m^2] | n |  | average irradiance value [W/m2]
calibration uncertainty | 1.92 | #REF! | 2 | #REF! | ∞ | 1: Akwatia | #REF! | #REF! | #REF!
directional error | 0.2 | #REF! | 1.7321 | #REF! | ∞ | 2: Accra | #REF! | 647.6788 | #REF!
temperature dependence of the sensitivity | 1 | #REF! | 1.7321 | #REF! | ∞ | 3: Kumasi | #REF! | 543.5545 | #REF!
cable sensitivity | 0.02 | #REF! | 1.7321 | #REF! | ∞
analog to digital conversion | 0.05 | #REF! | 1.7321 | #REF! | ∞
overall significance level (alpha) | 2
combined standard uncertainty | 2.1747 |  | "+/-" | #REF! | W/m^2
t_value (student distribution) |  |  |  | 2.236
expanded standard uncertainty | 4.8626 |  | +/- | #REF! | W/m^2
Accra
uncertainty | relative error [+/- %] | gross value [W/m^2] | divisor | u [W/m^2] | n
calibration unvertainty | 1.92 | 12.4354 | 2 | 6.2177 | ∞
directional error | 0.2 | 1.2954 | 1.7321 | 0.74788 | ∞
temperature dependence of the sensitivity | 1 | 6.4768 | 1.7321 | 3.7394 | ∞
cable sensitivity | 0.02 | 0.12954 | 1.7321 | 0.074788 | ∞
analog to digital conversion | 0.05 | 0.32384 | 1.7321 | 0.18697 | ∞
overall significance level (alpha) | 2
combined standard uncertainty | 2.1747 |  | "+/-" | 7.2968 | W/m^2
t_value (student distribution) |  |  |  | 2.236
expanded standard uncertainty | 4.8626 |  | +/- | 16.3156 | W/m^2
Kumasi
uncertainty | relative error [+/- %] | gross value [W/m^2] | divisor | u [W/m^2] | n
calibration unvertainty | 1.92 | 10.4362 | 2 | 5.2181 | ∞
directional error | 0.2 | #REF! | 1.7321 | #REF! | ∞
temperature dependence of the sensitivity | 1 | #REF! | 1.7321 | #REF! | ∞
cable sensitivity | 0.02 | #REF! | 1.7321 | #REF! | ∞
analog to digital conversion | 0.05 | #REF! | 1.7321 | #REF! | ∞
overall significance level (alpha) | 2
combined standard uncertainty | 2.1747 |  | "+/-" | #REF! | W/m^2
t_value (student distribution) |  |  |  | 2.236
expanded standard uncertainty | 4.8626 |  | +/- | #REF! | W/m^2
Average
combined standard uncertainty |  |  |  | #REF! | W/m^2
expanded standard uncertainty |  |  |  | #REF! | W/m^2

## Error Propagation (bottom pyra)
Error Propagation - bottom Pyranometer
Akwatia
uncertainty | relative error [+/- %] | gross value [W/m^2] | divisor | u [W/m^2] | n |  | average irradiance value [W/m2]
calibration uncertainty | 1.92 | #REF! | 2 | #REF! | ∞ | 1: Akwatia | #REF! | #REF! | #REF!
directional error | 0.2 | #REF! | 1.7321 | #REF! | ∞ | 2: Accra | #REF! | 173.882 | #REF!
temperature dependence of the sensitivity | 1 | #REF! | 1.7321 | #REF! | ∞ | 3: Kumasi | #REF! | 168.0224 | #REF!
cable sensitivity | 0.02 | #REF! | 1.7321 | #REF! | ∞
bottom inclination | 0.2 | #REF! | 1.7321 | #REF! | ∞
analog to digital conversion | 0.05 | #REF! | 1.7321 | #REF! | ∞
overall significance level (alpha) | 2
combined standard uncertainty | 2.1839 |  | "+/-" | #REF! | W/m^2
t_value (student distribution) |  |  |  | 2.236
expanded standard uncertainty | 4.8831 |  | +/- | #REF! | W/m^2
Accra
uncertainty | relative error [+/- %] | gross value [W/m^2] | divisor | u [W/m^2] | n
calibration uncertainty | 1.92 | 3.3385 | 2 | 1.6693 | ∞
directional error | 0.2 | 0.34776 | 1.7321 | 0.20078 | ∞
temperature dependence of the sensitivity | 1 | 1.7388 | 1.7321 | 1.0039 | ∞
cable sensitivity | 0.02 | 0.034776 | 1.7321 | 0.020078 | ∞
bottom inclination | 0.2 | 0.34776 | 1.7321 | 0.20078 | ∞
analog to digital conversion | 0.05 | 0.086941 | 1.7321 | 0.050195 | ∞
overall significance level (alpha) | 2
combined standard uncertainty | 2.1839 |  | "+/-" | 1.9692 | W/m^2
t_value (student distribution) |  |  |  | 2.236
expanded standard uncertainty | 4.8831 |  | +/- | 4.4032 | W/m^2
Kumasi
uncertainty | relative error [+/- %] | gross value [W/m^2] | divisor | u [W/m^2] | n
calibration uncertainty | 1.92 | 3.226 | 2 | 1.613 | ∞
directional error | 0.2 | 0.33604 | 1.7321 | 0.19402 | ∞
temperature dependence of the sensitivity | 1 | 1.6802 | 1.7321 | 0.97008 | ∞
cable sensitivity | 0.02 | 0.033604 | 1.7321 | 0.019402 | ∞
bottom inclination | 0.2 | 0.33604 | 1.7321 | 0.19402 | ∞
analog to digital conversion | 0.05 | 0.084011 | 1.7321 | 0.048504 | ∞
overall significance level (alpha) | 2
combined standard uncertainty | 2.1839 |  | "+/-" | 1.9029 | W/m^2
t_value (student distribution) |  |  |  | 2.236
expanded standard uncertainty | 4.8831 |  | +/- | 4.2548 | W/m^2
Average
combined standard uncertainty |  |  |  | #REF! | W/m^2
expanded standard uncertainty |  |  |  | #REF! | W/m^2

## Albedo Comparision
Albedo
time [min.] | Akwatia | Accra | Kumasi | Average
0:00 | 0.20994 | 0.26526 | 0.2835 | 0.2529 | Measurement Interval: | 0:00 | s
0:00 | 0.20993 | 0.26532 | 0.28359 | 0.25294
0:00 | 0.20979 | 0.26529 | 0.28357 | 0.25288
0:00 | 0.2098 | 0.26536 | 0.2836 | 0.25292
0:00 | 0.20977 | 0.26552 | 0.2836 | 0.25296
0:00 | 0.20973 | 0.2656 | 0.28359 | 0.25297
0:00 | 0.20983 | 0.26552 | 0.28356 | 0.25297
0:01 | 0.2098 | 0.26565 | 0.28343 | 0.25296
0:01 | 0.20971 | 0.2656 | 0.28347 | 0.25293
0:01 | 0.20965 | 0.26559 | 0.28354 | 0.25293
0:01 | 0.20971 | 0.26563 | 0.28327 | 0.25287
0:01 | 0.20963 | 0.26565 | 0.28356 | 0.25295
0:01 | 0.20948 | 0.26578 | 0.28342 | 0.25289
0:01 | 0.20951 | 0.26583 | 0.28359 | 0.25298
0:01 | 0.2091 | 0.26592 | 0.28348 | 0.25283
0:01 | 0.20685 | 0.26595 | 0.28334 | 0.25205
0:01 | 0.1985 | 0.26605 | 0.28328 | 0.24928
0:01 | 0.18582 | 0.266 | 0.28326 | 0.24503
0:01 | 0.18505 | 0.26599 | 0.28328 | 0.24477
0:02 | 0.19237 | 0.26606 | 0.28326 | 0.24723
0:02 | 0.18923 | 0.26604 | 0.28325 | 0.24617
0:02 | 0.18266 | 0.266 | 0.28328 | 0.24398
0:02 | 0.18284 | 0.26614 | 0.28324 | 0.24407
0:02 | 0.18319 | 0.26618 | 0.28307 | 0.24414
0:02 | 0.18337 | 0.26618 | 0.28316 | 0.24423
0:02 | 0.18343 | 0.26611 | 0.28295 | 0.24416
0:02 | 0.18365 | 0.26623 | 0.283 | 0.24429
0:02 | 0.18392 | 0.26641 | 0.283 | 0.24444
0:02 | 0.1842 | 0.26616 | 0.28305 | 0.24447
0:02 | 0.18421 | 0.26627 | 0.28309 | 0.24453
0:03 | 0.18544 | 0.26646 | 0.28303 | 0.24498
0:03 | 0.19144 | 0.2666 | 0.28306 | 0.24703
0:03 | 0.20267 | 0.26661 | 0.28311 | 0.2508
0:03 | 0.20152 | 0.2667 | 0.28316 | 0.25046
0:03 | 0.20819 | 0.26666 | 0.28295 | 0.2526
0:03 | 0.21076 | 0.26694 | 0.28305 | 0.25358
0:03 | 0.21045 | 0.26696 | 0.28271 | 0.25338
0:03 | 0.21018 | 0.26699 | 0.2828 | 0.25332
0:03 | 0.20994 | 0.267 | 0.28248 | 0.25314
0:03 | 0.2099 | 0.26707 | 0.28243 | 0.25313
0:03 | 0.20986 | 0.26678 | 0.28053 | 0.25239
0:03 | 0.20977 | 0.26667 | 0.27927 | 0.2519
0:04 | 0.20966 | 0.26673 | 0.28033 | 0.25224
0:04 | 0.20952 | 0.2668 | 0.28237 | 0.25289
0:04 | 0.20944 | 0.26686 | 0.2826 | 0.25297
0:04 | 0.20914 | 0.2669 | 0.2827 | 0.25291
0:04 | 0.20124 | 0.26696 | 0.28263 | 0.25028
0:04 | 0.18293 | 0.267 | 0.28253 | 0.24415
0:04 | 0.18233 | 0.2671 | 0.2825 | 0.24398
0:04 | 0.18271 | 0.26709 | 0.28235 | 0.24405
0:04 | 0.18293 | 0.26705 | 0.28209 | 0.24402
0:04 | 0.18327 | 0.26697 | 0.28236 | 0.2442
0:04 | 0.18309 | 0.26692 | 0.28244 | 0.24415
0:04 | 0.1828 | 0.26699 | 0.28227 | 0.24402
0:05 | 0.18294 | 0.26722 | 0.28238 | 0.24418
0:05 | 0.18295 | 0.26707 | 0.28258 | 0.2442
0:05 | 0.18289 | 0.26722 | 0.28256 | 0.24422
0:05 | 0.18297 | 0.26712 | 0.28258 | 0.24423
0:05 | 0.18313 | 0.26712 | 0.28259 | 0.24428
0:05 | 0.18309 | 0.26718 | 0.28245 | 0.24424
0:05 | 0.18317 | 0.26703 | 0.28254 | 0.24425
0:05 | 0.18311 | 0.267 | 0.28258 | 0.24423
0:05 | 0.18324 | 0.26716 | 0.28238 | 0.24426
0:05 | 0.18303 | 0.267 | 0.28226 | 0.2441
0:05 | 0.18308 | 0.26715 | 0.27987 | 0.24336
0:05 | 0.18296 | 0.26726 | 0.27867 | 0.24296
0:06 | 0.18316 | 0.26717 | 0.27867 | 0.243
0:06 | 0.1837 | 0.26728 | 0.27858 | 0.24319
0:06 | 0.18383 | 0.26725 | 0.27853 | 0.2432
0:06 | 0.18385 | 0.26716 | 0.27942 | 0.24348
0:06 | 0.18379 | 0.26729 | 0.27918 | 0.24342
0:06 | 0.18394 | 0.26723 | 0.27914 | 0.24344
0:06 | 0.18422 | 0.26721 | 0.27934 | 0.24359
0:06 | 0.18419 | 0.26708 | 0.28083 | 0.24403
0:06 | 0.18463 | 0.26708 | 0.28167 | 0.24446
0:06 | 0.18467 | 0.26712 | 0.28162 | 0.24447
0:06 | 0.18465 | 0.26707 | 0.28251 | 0.24474
0:06 | 0.18436 | 0.26716 | 0.28288 | 0.2448
0:07 | 0.18447 | 0.26716 | 0.28295 | 0.24486
0:07 | 0.18474 | 0.26701 | 0.28166 | 0.24447
0:07 | 0.18507 | 0.26704 | 0.28048 | 0.2442
0:07 | 0.18454 | 0.26704 | 0.28017 | 0.24392
0:07 | 0.18466 | 0.26693 | 0.28009 | 0.2439
0:07 | 0.18434 | 0.26689 | 0.2801 | 0.24378
0:07 | 0.18441 | 0.26675 | 0.28014 | 0.24377
0:07 | 0.1844 | 0.26666 | 0.27933 | 0.24346
0:07 | 0.18449 | 0.26705 | 0.27889 | 0.24348
0:07 | 0.18446 | 0.26701 | 0.27881 | 0.24343
0:07 | 0.18453 | 0.26717 | 0.27831 | 0.24334
0:07 | 0.18427 | 0.26725 | 0.27655 | 0.24269
0:08 | 0.18436 | 0.26733 | 0.27506 | 0.24225
0:08 | 0.18426 | 0.26727 | 0.27672 | 0.24275
0:08 | 0.18403 | 0.26742 | 0.27794 | 0.24313
0:08 | 0.18409 | 0.26748 | 0.27804 | 0.2432
0:08 | 0.18395 | 0.26752 | 0.27799 | 0.24315
0:08 | 0.18386 | 0.26744 | 0.27804 | 0.24311
0:08 | 0.18354 | 0.26745 | 0.27811 | 0.24303
0:08 | 0.18334 | 0.26739 | 0.27824 | 0.24299
0:08 | 0.18297 | 0.26746 | 0.27869 | 0.24304
0:08 | 0.18281 | 0.26745 | 0.27977 | 0.24334
0:08 | 0.18249 | 0.26744 | 0.28024 | 0.24339
0:08 | 0.18225 | 0.2675 | 0.28004 | 0.24327
0:09 | 0.18208 | 0.26751 | 0.27952 | 0.24304
0:09 | 0.1821 | 0.26762 | 0.27939 | 0.24303
0:09 | 0.18192 | 0.26762 | 0.27926 | 0.24293
0:09 | 0.18223 | 0.26763 | 0.27922 | 0.24303
0:09 | 0.18221 | 0.26751 | 0.27936 | 0.24302
0:09 | 0.18197 | 0.26755 | 0.27902 | 0.24285
0:09 | 0.18235 | 0.26749 | 0.27907 | 0.24297
0:09 | 0.1823 | 0.26739 | 0.27886 | 0.24285
0:09 | 0.18251 | 0.26744 | 0.27881 | 0.24292
0:09 | 0.18279 | 0.2676 | 0.27892 | 0.2431
0:09 | 0.18262 | 0.26755 | 0.27961 | 0.24326
0:09 | 0.18284 | 0.26761 | 0.28075 | 0.24373
0:10 | 0.18275 | 0.2676 | 0.28117 | 0.24384
0:10 | 0.18272 | 0.26747 | 0.28128 | 0.24382
0:10 | 0.18273 | 0.26749 | 0.28113 | 0.24378
0:10 | 0.18311 | 0.26745 | 0.28106 | 0.24388
0:10 | 0.18324 | 0.26753 | 0.28095 | 0.24391
0:10 | 0.18338 | 0.26763 | 0.28106 | 0.24402
0:10 | 0.18327 | 0.26732 | 0.28106 | 0.24388
0:10 | 0.18347 | 0.26743 | 0.28085 | 0.24392
0:10 | 0.1841 | 0.26751 | 0.281 | 0.2442
0:10 | 0.18563 | 0.2673 | 0.28077 | 0.24456
0:10 | 0.1984 | 0.26724 | 0.28086 | 0.24883
0:10 | 0.20971 | 0.26725 | 0.281 | 0.25266
0:11 | 0.20894 | 0.26692 | 0.281 | 0.25229
0:11 | 0.19928 | 0.2658 | 0.28144 | 0.24884
0:11 | 0.18489 | 0.26563 | 0.28145 | 0.24399
0:11 | 0.1832 | 0.26545 | 0.28178 | 0.24348
0:11 | 0.18324 | 0.26539 | 0.28192 | 0.24352
0:11 | 0.18324 | 0.26548 | 0.28219 | 0.24364
0:11 | 0.18336 | 0.26448 | 0.28217 | 0.24334
0:11 | 0.18348 | 0.26384 | 0.28237 | 0.24323
0:11 | 0.18304 | 0.26497 | 0.28215 | 0.24338
0:11 | 0.1829 | 0.26553 | 0.28228 | 0.24357
0:11 | 0.1826 | 0.26533 | 0.28242 | 0.24345
0:11 | 0.18256 | 0.26529 | 0.2824 | 0.24341
0:12 | 0.18235 | 0.26525 | 0.28263 | 0.24341
0:12 | 0.18215 | 0.26522 | 0.28265 | 0.24334
0:12 | 0.18206 | 0.26515 | 0.28256 | 0.24326
0:12 | 0.18206 | 0.26504 | 0.2826 | 0.24323
0:12 | 0.18202 | 0.26492 | 0.28243 | 0.24312
0:12 | 0.18203 | 0.26497 | 0.28241 | 0.24313
0:12 | 0.18234 | 0.26484 | 0.28241 | 0.24319
0:12 | 0.18277 | 0.26482 | 0.28221 | 0.24327
0:12 | 0.183 | 0.26518 | 0.28229 | 0.24349
0:12 | 0.18302 | 0.26502 | 0.2825 | 0.24351
0:12 | 0.18308 | 0.26501 | 0.28247 | 0.24352
0:12 | 0.18305 | 0.26511 | 0.2824 | 0.24352
0:12 | 0.18292 | 0.26508 | 0.28258 | 0.24353
0:13 | 0.18269 | 0.26525 | 0.28231 | 0.24342
0:13 | 0.18253 | 0.26514 | 0.28259 | 0.24342
0:13 | 0.18262 | 0.26507 | 0.28247 | 0.24339
0:13 | 0.18255 | 0.26502 | 0.28247 | 0.24334
0:13 | 0.18231 | 0.26506 | 0.28235 | 0.24324
0:13 | 0.18233 | 0.26498 | 0.28248 | 0.24326
0:13 | 0.18207 | 0.2651 | 0.28243 | 0.2432
0:13 | 0.18214 | 0.26507 | 0.28248 | 0.24323
0:13 | 0.18206 | 0.26474 | 0.28243 | 0.24308
0:13 | 0.18159 | 0.26454 | 0.28237 | 0.24283
0:13 | 0.18211 | 0.26466 | 0.28226 | 0.24301
0:13 | 0.1818 | 0.26467 | 0.28224 | 0.2429
0:14 | 0.18153 | 0.26466 | 0.2822 | 0.24279
0:14 | 0.18155 | 0.26474 | 0.28217 | 0.24282
0:14 | 0.18179 | 0.26481 | 0.28213 | 0.24291
0:14 | 0.18177 | 0.26417 | 0.2822 | 0.24272
0:14 | 0.18173 | 0.26391 | 0.28221 | 0.24262
0:14 | 0.18164 | 0.26404 | 0.28223 | 0.24264
0:14 | 0.18162 | 0.26437 | 0.28224 | 0.24274
0:14 | 0.1815 | 0.26435 | 0.2823 | 0.24272
0:14 | 0.18166 | 0.26428 | 0.2823 | 0.24275
0:14 | 0.18162 | 0.26514 | 0.28226 | 0.24301
0:14 | 0.18169 | 0.26637 | 0.28232 | 0.24346
0:14 | 0.1816 | 0.26644 | 0.28235 | 0.24346
0:15 | 0.18179 | 0.26645 | 0.28243 | 0.24356
0:15 | 0.18165 | 0.26621 | 0.28239 | 0.24341
0:15 | 0.18165 | 0.26636 | 0.2825 | 0.2435
0:15 | 0.18175 | 0.26642 | 0.28231 | 0.24349
0:15 | 0.18192 | 0.26658 | 0.28228 | 0.2436
0:15 | 0.1818 | 0.26647 | 0.28231 | 0.24353
0:15 | 0.18205 | 0.26661 | 0.28238 | 0.24368
0:15 | 0.18208 | 0.26658 | 0.28229 | 0.24365
0:15 | 0.18206 | 0.26661 | 0.28235 | 0.24367
0:15 | 0.1819 | 0.26645 | 0.28236 | 0.24357
0:15 | 0.18204 | 0.26652 | 0.28237 | 0.24364
0:15 | 0.18204 | 0.26646 | 0.28234 | 0.24361
0:16 | 0.18208 | 0.26648 | 0.28233 | 0.24363
0:16 | 0.18251 | 0.26643 | 0.2825 | 0.24381
0:16 | 0.18233 | 0.26624 | 0.28239 | 0.24365
0:16 | 0.18286 | 0.2664 | 0.28248 | 0.24392
0:16 | 0.18287 | 0.26624 | 0.28242 | 0.24384
0:16 | 0.18359 | 0.26549 | 0.28266 | 0.24391
0:16 | 0.18401 | 0.26429 | 0.2825 | 0.2436
0:16 | 0.18444 | 0.26408 | 0.28257 | 0.2437
0:16 | 0.18659 | 0.26286 | 0.28256 | 0.244
0:16 | 0.1866 | 0.26273 | 0.2828 | 0.24404
0:16 | 0.18732 | 0.26468 | 0.28286 | 0.24495
0:16 | 0.19428 | 0.26506 | 0.28284 | 0.24739
0:17 | 0.19032 | 0.26539 | 0.2828 | 0.24617
0:17 | 0.19196 | 0.26557 | 0.28291 | 0.24681
0:17 | 0.19395 | 0.2654 | 0.28296 | 0.24744
0:17 | 0.1942 | 0.26367 | 0.28285 | 0.24691
0:17 | 0.1968 | 0.26139 | 0.28283 | 0.24701
0:17 | 0.19385 | 0.26196 | 0.28285 | 0.24622
0:17 | 0.19906 | 0.26237 | 0.28286 | 0.24809
0:17 | 0.19124 | 0.26257 | 0.2829 | 0.24557
0:17 | 0.18592 | 0.26262 | 0.28286 | 0.2438
0:17 | 0.1872 | 0.26233 | 0.2829 | 0.24414
0:17 | 0.18941 | 0.26227 | 0.28284 | 0.24484
0:17 | 0.19887 | 0.26229 | 0.2828 | 0.24799
0:18 | 0.20681 | 0.26292 | 0.28284 | 0.25085
0:18 | 0.21086 | 0.26277 | 0.28304 | 0.25222
0:18 | 0.21033 | 0.26269 | 0.28289 | 0.25197
0:18 | 0.21036 | 0.26241 | 0.28276 | 0.25184
0:18 | 0.21035 | 0.26182 | 0.28289 | 0.25169
0:18 | 0.21027 | 0.26163 | 0.28286 | 0.25159
0:18 | 0.20993 | 0.26162 | 0.28279 | 0.25145
0:18 | 0.20969 | 0.2619 | 0.28284 | 0.25148
0:18 | 0.20879 | 0.26208 | 0.28308 | 0.25131
0:18 | 0.20621 | 0.26228 | 0.28286 | 0.25045
0:18 | 0.20957 | 0.26263 | 0.28294 | 0.25172
0:18 | 0.20961 | 0.26318 | 0.28299 | 0.25193
0:19 | 0.20935 | 0.26344 | 0.28315 | 0.25198
0:19 | 0.20965 | 0.26332 | 0.28295 | 0.25197
0:19 | 0.20936 | 0.26341 | 0.28304 | 0.25194
0:19 | 0.20912 | 0.26353 | 0.28276 | 0.2518
0:19 | 0.20894 | 0.26382 | 0.28289 | 0.25188
0:19 | 0.20904 | 0.26388 | 0.28298 | 0.25197
0:19 | 0.20871 | 0.26397 | 0.283 | 0.25189
0:19 | 0.20879 | 0.26384 | 0.28308 | 0.2519
0:19 | 0.20863 | 0.2641 | 0.28321 | 0.25198
0:19 | 0.20873 | 0.26433 | 0.28321 | 0.25209
0:19 | 0.20871 | 0.26402 | 0.28324 | 0.25199
0:19 | 0.20869 | 0.26377 | 0.28309 | 0.25185
0:20 | 0.20848 | 0.26374 | 0.28306 | 0.25176
0:20 | 0.20849 | 0.26392 | 0.28309 | 0.25184
0:20 | 0.20872 | 0.26404 | 0.283 | 0.25192
0:20 | 0.209 | 0.26427 | 0.28274 | 0.25201
0:20 | 0.20908 | 0.26379 | 0.28254 | 0.2518
0:20 | 0.20919 | 0.26388 | 0.2825 | 0.25186
0:20 | 0.20914 | 0.26453 | 0.28243 | 0.25203
0:20 | 0.20919 | 0.26612 | 0.28252 | 0.25261
0:20 | 0.20899 | 0.26643 | 0.28238 | 0.2526
0:20 | 0.20901 | 0.26632 | 0.28248 | 0.2526
0:20 | 0.20915 | 0.26632 | 0.28214 | 0.25254
0:20 | 0.20919 | 0.26627 | 0.2821 | 0.25252
0:21 | 0.20932 | 0.2664 | 0.28187 | 0.25253
0:21 | 0.20899 | 0.26637 | 0.28225 | 0.25254
0:21 | 0.20867 | 0.2665 | 0.28238 | 0.25252
0:21 | 0.20874 | 0.2664 | 0.28229 | 0.25248
0:21 | 0.20867 | 0.26663 | 0.28249 | 0.2526
0:21 | 0.20871 | 0.26652 | 0.2826 | 0.25261
0:21 | 0.20877 | 0.26645 | 0.28267 | 0.25263
0:21 | 0.20879 | 0.26656 | 0.28281 | 0.25272
0:21 | 0.20875 | 0.26649 | 0.28298 | 0.25274
0:21 | 0.20879 | 0.26662 | 0.28289 | 0.25277
0:21 | 0.20896 | 0.26658 | 0.28287 | 0.2528
0:21 | 0.20901 | 0.26648 | 0.28311 | 0.25287
0:22 | 0.20902 | 0.2662 | 0.28301 | 0.25274
0:22 | 0.20932 | 0.2663 | 0.28296 | 0.25286
0:22 | 0.20921 | 0.26619 | 0.28292 | 0.25277
0:22 | 0.20923 | 0.26614 | 0.28298 | 0.25278
0:22 | 0.2093 | 0.26614 | 0.2829 | 0.25278
0:22 | 0.20923 | 0.266 | 0.28297 | 0.25273
0:22 | 0.20928 | 0.26608 | 0.28288 | 0.25275
0:22 | 0.20938 | 0.26575 | 0.28285 | 0.25266
0:22 | 0.20957 | 0.26573 | 0.28299 | 0.25277
0:22 | 0.20979 | 0.26562 | 0.28291 | 0.25277
0:22 | 0.20984 | 0.26563 | 0.2828 | 0.25276
0:22 | 0.20998 | 0.26565 | 0.28289 | 0.25284
0:23 | 0.21012 | 0.26549 | 0.28269 | 0.25277
0:23 | 0.20987 | 0.26549 | 0.28277 | 0.25271
0:23 | 0.2099 | 0.26526 | 0.28287 | 0.25268
0:23 | 0.21001 | 0.2647 | 0.28285 | 0.25252
0:23 | 0.2104 | 0.26449 | 0.28264 | 0.25251
0:23 | 0.21057 | 0.26475 | 0.28275 | 0.25269
0:23 | 0.2107 | 0.2647 | 0.28269 | 0.2527
0:23 | 0.21063 | 0.2646 | 0.28273 | 0.25265
0:23 | 0.21079 | 0.26402 | 0.28276 | 0.25252
0:23 | 0.2112 | 0.26271 | 0.28305 | 0.25232
0:23 | 0.21168 | 0.26389 | 0.28304 | 0.25287
0:23 | 0.21128 | 0.26407 | 0.28309 | 0.25282
0:24 | 0.21125 | 0.26411 | 0.2832 | 0.25285
0:24 | 0.21063 | 0.26395 | 0.2832 | 0.2526
0:24 | 0.2107 | 0.2643 | 0.28315 | 0.25272
0:24 | 0.21094 | 0.26516 | 0.28325 | 0.25312
0:24 | 0.21127 | 0.2651 | 0.28322 | 0.2532
0:24 | 0.21267 | 0.26525 | 0.28311 | 0.25368
0:24 | 0.21214 | 0.2653 | 0.28319 | 0.25354
0:24 | 0.21188 | 0.26554 | 0.28315 | 0.25353
0:24 | 0.21024 | 0.26562 | 0.28336 | 0.25307
0:24 | 0.20046 | 0.26533 | 0.28328 | 0.24969
0:24 | 0.19067 | 0.26538 | 0.28335 | 0.24647
0:24 | 0.19036 | 0.26539 | 0.2833 | 0.24635
0:25 | 0.18964 | 0.2654 | 0.28319 | 0.24608
0:25 | 0.1901 | 0.26542 | 0.28334 | 0.24629
0:25 | 0.19002 | 0.26543 | 0.28328 | 0.24624
0:25 | 0.19014 | 0.26546 | 0.28331 | 0.2463
0:25 | 0.18999 | 0.26542 | 0.28333 | 0.24624
0:25 | 0.19012 | 0.26554 | 0.28339 | 0.24635
0:25 | 0.19019 | 0.26541 | 0.2833 | 0.2463
0:25 | 0.19016 | 0.26544 | 0.28323 | 0.24628
0:25 | 0.19024 | 0.26546 | 0.28338 | 0.24636
0:25 | 0.1913 | 0.26557 | 0.28327 | 0.24671
0:25 | 0.19494 | 0.26557 | 0.28324 | 0.24792
0:25 | 0.19757 | 0.26566 | 0.28322 | 0.24882
0:26 | 0.19493 | 0.26592 | 0.28326 | 0.24803
0:26 | 0.19138 | 0.26605 | 0.2833 | 0.24691
0:26 | 0.19039 | 0.26612 | 0.28328 | 0.24659
0:26 | 0.19073 | 0.26601 | 0.28325 | 0.24666
0:26 | 0.19094 | 0.26592 | 0.28308 | 0.24664
0:26 | 0.19048 | 0.26593 | 0.28328 | 0.24656
0:26 | 0.19064 | 0.26608 | 0.28324 | 0.24666
0:26 | 0.19022 | 0.26583 | 0.28315 | 0.2464
0:26 | 0.19053 | 0.26592 | 0.28311 | 0.24652
0:26 | 0.19015 | 0.26591 | 0.28307 | 0.24638
0:26 | 0.19025 | 0.26601 | 0.2831 | 0.24646
0:26 | 0.19032 | 0.26614 | 0.28314 | 0.24654
0:27 | 0.19038 | 0.26605 | 0.28298 | 0.24647
0:27 | 0.19027 | 0.26605 | 0.28304 | 0.24645
0:27 | 0.18987 | 0.26613 | 0.2831 | 0.24637
0:27 | 0.19004 | 0.26602 | 0.28317 | 0.24641
0:27 | 0.19013 | 0.26587 | 0.28317 | 0.24639
0:27 | 0.19004 | 0.26601 | 0.28319 | 0.24641
0:27 | 0.18999 | 0.26603 | 0.28328 | 0.24643
0:27 | 0.19012 | 0.26611 | 0.28322 | 0.24648
0:27 | 0.19044 | 0.26602 | 0.2832 | 0.24655
0:27 | 0.19038 | 0.2658 | 0.28317 | 0.24645
0:27 | 0.19053 | 0.26592 | 0.28311 | 0.24652
0:27 | 0.19061 | 0.26603 | 0.28324 | 0.24663
0:28 | 0.19081 | 0.26613 | 0.28318 | 0.24671
0:28 | 0.19044 | 0.26633 | 0.28318 | 0.24665
0:28 | 0.19032 | 0.26627 | 0.28334 | 0.24665
0:28 | 0.19062 | 0.26636 | 0.28317 | 0.24672
0:28 | 0.19082 | 0.26616 | 0.28308 | 0.24668
0:28 | 0.19036 | 0.26627 | 0.28323 | 0.24662
0:28 | 0.19035 | 0.26643 | 0.28335 | 0.24671
0:28 | 0.1903 | 0.26632 | 0.28325 | 0.24662
0:28 | 0.19034 | 0.26637 | 0.28331 | 0.24668
0:28 | 0.19012 | 0.26618 | 0.28341 | 0.24657
0:28 | 0.19008 | 0.2663 | 0.2833 | 0.24656
0:28 | 0.18985 | 0.26619 | 0.28195 | 0.246
0:29 | 0.19004 | 0.2662 | 0.28299 | 0.24641
0:29 | 0.18999 | 0.26628 | 0.28376 | 0.24668
0:29 | 0.1898 | 0.26632 | 0.28354 | 0.24656
0:29 | 0.1901 | 0.26634 | 0.28335 | 0.2466
0:29 | 0.19016 | 0.26622 | 0.2835 | 0.24662
0:29 | 0.1903 | 0.26621 | 0.28353 | 0.24668
0:29 | 0.19007 | 0.26625 | 0.28325 | 0.24653
0:29 | 0.19006 | 0.26637 | 0.2831 | 0.24651
0:29 | 0.1898 | 0.26641 | 0.28296 | 0.24639
0:29 | 0.18973 | 0.26638 | 0.28292 | 0.24634
0:29 | 0.18989 | 0.26632 | 0.28281 | 0.24634
0:29 | 0.18954 | 0.26627 | 0.28252 | 0.24611
0:30 | 0.18974 | 0.26615 | 0.2825 | 0.24613
0:30 | 0.1894 | 0.2662 | 0.28275 | 0.24612
0:30 | 0.18909 | 0.26617 | 0.2827 | 0.24599
0:30 | 0.18913 | 0.26637 | 0.28273 | 0.24608
0:30 | 0.18906 | 0.26645 | 0.28282 | 0.24611
0:30 | 0.18891 | 0.26632 | 0.28275 | 0.24599
0:30 | 0.18905 | 0.26629 | 0.28275 | 0.24603
0:30 | 0.18885 | 0.26645 | 0.28287 | 0.24606
0:30 | 0.18867 | 0.26636 | 0.28286 | 0.24596
0:30 | 0.18849 | 0.26632 | 0.28309 | 0.24597
0:30 | 0.18799 | 0.26633 | 0.28367 | 0.246
0:30 | 0.18775 | 0.26627 | 0.28404 | 0.24602
0:31 | 0.18763 | 0.26615 | 0.28418 | 0.24599
0:31 | 0.18757 | 0.26635 | 0.28443 | 0.24612
0:31 | 0.18719 | 0.2663 | 0.2845 | 0.24599
0:31 | 0.18729 | 0.26642 | 0.28461 | 0.24611
0:31 | 0.18728 | 0.2663 | 0.28446 | 0.24602
0:31 | 0.18826 | 0.2664 | 0.28448 | 0.24638
0:31 | 0.18895 | 0.26652 | 0.28461 | 0.24669
0:31 | 0.18993 | 0.26666 | 0.28462 | 0.24707
0:31 | 0.19181 | 0.26674 | 0.28463 | 0.24773
0:31 | 0.19395 | 0.26665 | 0.28467 | 0.24842
0:31 | 0.19499 | 0.26663 | 0.28447 | 0.2487
0:32 | 0.19691 | 0.26666 | 0.28454 | 0.24937
0:32 | 0.20039 | 0.26662 | 0.28458 | 0.25053
0:32 | 0.20285 | 0.2667 | 0.28446 | 0.25134
0:32 | 0.20469 | 0.26648 | 0.28449 | 0.25189
0:32 | 0.19439 | 0.26647 | 0.28467 | 0.24851
0:32 | 0.18319 | 0.26646 | 0.28463 | 0.24476
0:32 | 0.19161 | 0.26664 | 0.28475 | 0.24767
0:32 | 0.19532 | 0.26665 | 0.28458 | 0.24885
0:32 | 0.20274 | 0.26673 | 0.2847 | 0.25139
0:32 | 0.20806 | 0.26661 | 0.28457 | 0.25308
0:32 | 0.20596 | 0.26651 | 0.28487 | 0.25245
0:32 | 0.20428 | 0.26653 | 0.28492 | 0.25191
0:33 | 0.20506 | 0.26661 | 0.28502 | 0.25223
0:33 | 0.20655 | 0.26671 | 0.28501 | 0.25276
0:33 | 0.20671 | 0.26677 | 0.28515 | 0.25288
0:33 | 0.20657 | 0.26684 | 0.28516 | 0.25286
0:33 | 0.20647 | 0.26681 | 0.28493 | 0.25274
0:33 | 0.20637 | 0.2667 | 0.28491 | 0.25266
0:33 | 0.20627 | 0.26668 | 0.28477 | 0.25257
0:33 | 0.20578 | 0.26676 | 0.28443 | 0.25233
0:33 | 0.20283 | 0.26669 | 0.28433 | 0.25128
0:33 | 0.2036 | 0.2666 | 0.28432 | 0.25151
0:33 | 0.20581 | 0.26681 | 0.28411 | 0.25224
0:33 | 0.20613 | 0.26677 | 0.28441 | 0.25243
0:34 | 0.20521 | 0.26682 | 0.28411 | 0.25205
0:34 | 0.20509 | 0.26676 | 0.28398 | 0.25194
0:34 | 0.2048 | 0.2668 | 0.28387 | 0.25182
0:34 | 0.20481 | 0.2668 | 0.28414 | 0.25191
0:34 | 0.2046 | 0.2668 | 0.28444 | 0.25195
0:34 | 0.2048 | 0.26672 | 0.2844 | 0.25197
0:34 | 0.20504 | 0.26679 | 0.28465 | 0.25216
0:34 | 0.20499 | 0.26676 | 0.2846 | 0.25212
0:34 | 0.205 | 0.26676 | 0.28446 | 0.25207
0:34 | 0.20498 | 0.26695 | 0.28471 | 0.25221
0:34 | 0.20484 | 0.26701 | 0.28445 | 0.2521
0:34 | 0.20487 | 0.26717 | 0.28475 | 0.25226
0:35 | 0.20489 | 0.26717 | 0.285 | 0.25235
0:35 | 0.20525 | 0.26708 | 0.28459 | 0.25231
0:35 | 0.20559 | 0.2673 | 0.28496 | 0.25262
0:35 | 0.20555 | 0.26714 | 0.28493 | 0.25254
0:35 | 0.20546 | 0.26722 | 0.28521 | 0.25263
0:35 | 0.20548 | 0.26703 | 0.28504 | 0.25252
0:35 | 0.20543 | 0.26695 | 0.28483 | 0.2524
0:35 | 0.2055 | 0.26691 | 0.28513 | 0.25251
0:35 | 0.20557 | 0.26674 | 0.28563 | 0.25264
0:35 | 0.20551 | 0.26692 | 0.28617 | 0.25287
0:35 | 0.2057 | 0.26693 | 0.28646 | 0.25303
0:35 | 0.20559 | 0.26682 | 0.28651 | 0.25297
0:36 | 0.20544 | 0.26686 | 0.28659 | 0.25296
0:36 | 0.2057 | 0.26682 | 0.28679 | 0.2531
0:36 | 0.20578 | 0.26675 | 0.28674 | 0.25309
0:36 | 0.20525 | 0.26679 | 0.28665 | 0.25289
0:36 | 0.2047 | 0.26681 | 0.28667 | 0.25273
0:36 | 0.20407 | 0.26685 | 0.28676 | 0.25256
0:36 | 0.20322 | 0.26669 | 0.28686 | 0.25225
0:36 | 0.20382 | 0.26672 | 0.28656 | 0.25237
0:36 | 0.20345 | 0.26667 | 0.28646 | 0.25219
0:36 | 0.20464 | 0.26679 | 0.2865 | 0.25265
0:36 | 0.20539 | 0.26678 | 0.28653 | 0.2529
0:36 | 0.20462 | 0.26663 | 0.28674 | 0.25266
0:37 | 0.20364 | 0.26668 | 0.28674 | 0.25235
0:37 | 0.20154 | 0.26659 | 0.28692 | 0.25169
0:37 | 0.20063 | 0.26664 | 0.28753 | 0.2516
0:37 | 0.19243 | 0.26684 | 0.28777 | 0.24901
0:37 | 0.20162 | 0.26693 | 0.28811 | 0.25222
0:37 | 0.20786 | 0.26715 | 0.28846 | 0.25449
0:37 | 0.19453 | 0.26706 | 0.28858 | 0.25006
0:37 | 0.18741 | 0.26706 | 0.28864 | 0.2477
0:37 | 0.20492 | 0.26712 | 0.28877 | 0.2536
0:37 | 0.20601 | 0.26702 | 0.28856 | 0.25387
0:37 | 0.20532 | 0.26729 | 0.28895 | 0.25385
0:37 | 0.20559 | 0.26713 | 0.28899 | 0.2539
0:38 | 0.20661 | 0.26723 | 0.28926 | 0.25437
0:38 | 0.20661 | 0.26721 | 0.28923 | 0.25435
0:38 | 0.20634 | 0.267 | 0.29019 | 0.25451
0:38 | 0.20639 | 0.26697 | 0.29013 | 0.2545
0:38 | 0.20606 | 0.26698 | 0.29044 | 0.25449
0:38 | 0.20605 | 0.26698 | 0.29075 | 0.25459
0:38 | 0.20587 | 0.26694 | 0.29076 | 0.25452
0:38 | 0.20584 | 0.26716 | 0.29074 | 0.25458
0:38 | 0.20591 | 0.26712 | 0.29092 | 0.25465
0:38 | 0.20576 | 0.26729 | 0.29122 | 0.25476
0:38 | 0.20592 | 0.26717 | 0.291 | 0.25469
0:38 | 0.20592 | 0.26723 | 0.29152 | 0.25489
0:39 | 0.20594 | 0.26714 | 0.29176 | 0.25495
0:39 | 0.20596 | 0.26732 | 0.29229 | 0.25519
0:39 | 0.20607 | 0.26737 | 0.29294 | 0.25546
0:39 | 0.20614 | 0.26735 | 0.29362 | 0.2557
0:39 | 0.2062 | 0.26732 | 0.29393 | 0.25582
0:39 | 0.20617 | 0.26705 | 0.29423 | 0.25582
0:39 | 0.20642 | 0.26718 | 0.29443 | 0.25601
0:39 | 0.20654 | 0.26727 | 0.29491 | 0.25624
0:39 | 0.20671 | 0.2672 | 0.29564 | 0.25652
0:39 | 0.20625 | 0.26718 | 0.29563 | 0.25635
0:39 | 0.20214 | 0.26703 | 0.29598 | 0.25505
0:39 | 0.19076 | 0.26717 | 0.29607 | 0.25133
0:40 | 0.18118 | 0.26722 | 0.29683 | 0.24841
0:40 | 0.18355 | 0.26718 | 0.29695 | 0.24922
0:40 | 0.18285 | 0.26735 | 0.29708 | 0.2491
0:40 | 0.184 | 0.2674 | 0.29673 | 0.24938
0:40 | 0.18454 | 0.26743 | 0.29701 | 0.24966
0:40 | 0.18442 | 0.26747 | 0.29718 | 0.24969
0:40 | 0.18565 | 0.26722 | 0.29766 | 0.25018
0:40 | 0.18838 | 0.26742 | 0.29801 | 0.25127
0:40 | 0.18803 | 0.26735 | 0.29813 | 0.25117
0:40 | 0.19145 | 0.26752 | 0.29842 | 0.25247
0:40 | 0.19085 | 0.26756 | 0.29917 | 0.25253
0:40 | 0.18471 | 0.2676 | 0.29966 | 0.25066
0:41 | 0.18897 | 0.26764 | 0.29957 | 0.25206
0:41 | 0.20171 | 0.26759 | 0.30013 | 0.25648
0:41 | 0.20326 | 0.26752 | 0.30038 | 0.25705
0:41 | 0.19977 | 0.26752 | 0.30144 | 0.25624
0:41 | 0.19418 | 0.2676 | 0.30268 | 0.25482
0:41 | 0.18547 | 0.26762 | 0.30376 | 0.25228
0:41 | 0.17906 | 0.26774 | 0.30433 | 0.25038
0:41 | 0.17947 | 0.26777 | 0.30452 | 0.25059
0:41 | 0.17869 | 0.26781 | 0.30486 | 0.25045
0:41 | 0.17874 | 0.26778 | 0.30742 | 0.25131
0:41 | 0.17929 | 0.26785 | 0.30802 | 0.25172
0:41 | 0.17938 | 0.26783 | 0.30954 | 0.25225
0:42 | 0.17965 | 0.26793 | 0.31449 | 0.25402
0:42 | 0.18006 | 0.26786 | 0.31825 | 0.25539
0:42 | 0.18029 | 0.26768 | 0.3187 | 0.25556
0:42 | 0.18077 | 0.26785 | 0.31843 | 0.25568
0:42 | 0.18196 | 0.26784 | 0.31837 | 0.25606
0:42 | 0.18253 | 0.26798 | 0.31864 | 0.25639
0:42 | 0.18761 | 0.26803 | 0.31847 | 0.25804
0:42 | 0.1985 | 0.2679 | 0.31855 | 0.26165
0:42 | 0.19295 | 0.26793 | 0.31827 | 0.25972
0:42 | 0.18277 | 0.26794 | 0.31849 | 0.2564
0:42 | 0.18392 | 0.26794 | 0.31838 | 0.25675
0:42 | 0.1843 | 0.26797 | 0.31835 | 0.25687
0:43 | 0.18414 | 0.26558 | 0.31859 | 0.25611
0:43 | 0.18386 | 0.2607 | 0.31845 | 0.25434
0:43 | 0.18398 | 0.26775 | 0.31859 | 0.25677
0:43 | 0.1834 | 0.26786 | 0.31844 | 0.25657
0:43 | 0.18364 | 0.26783 | 0.31842 | 0.25663
0:43 | 0.18359 | 0.26787 | 0.31855 | 0.25667
0:43 | 0.18392 | 0.26782 | 0.3187 | 0.25681
0:43 | 0.19137 | 0.26772 | 0.31852 | 0.2592
0:43 | 0.18833 | 0.26769 | 0.31869 | 0.25824
0:43 | 0.1825 | 0.26768 | 0.31861 | 0.25626
0:43 | 0.18348 | 0.26783 | 0.31858 | 0.25663
0:43 | 0.18683 | 0.2678 | 0.31847 | 0.2577
0:44 | 0.19411 | 0.26777 | 0.31867 | 0.26018
0:44 | 0.20399 | 0.26765 | 0.31879 | 0.26348
0:44 | 0.18909 | 0.26741 | 0.31869 | 0.2584
0:44 | 0.1847 | 0.26746 | 0.31872 | 0.25696
0:44 | 0.18334 | 0.26763 | 0.31864 | 0.25654
0:44 | 0.18418 | 0.26769 | 0.31876 | 0.25687
0:44 | 0.18489 | 0.26771 | 0.31891 | 0.25717
0:44 | 0.18527 | 0.26758 | 0.31889 | 0.25725
0:44 | 0.18553 | 0.26752 | 0.31894 | 0.25733
0:44 | 0.18563 | 0.26757 | 0.31876 | 0.25732
0:44 | 0.19003 | 0.26769 | 0.31885 | 0.25886
0:44 | 0.20019 | 0.26755 | 0.31867 | 0.26214
0:45 | 0.2004 | 0.26749 | 0.31891 | 0.26227
0:45 | 0.20214 | 0.26749 | 0.31884 | 0.26282
0:45 | 0.18893 | 0.26732 | 0.31897 | 0.25841
0:45 | 0.18374 | 0.26731 | 0.31896 | 0.25667
0:45 | 0.18461 | 0.26732 | 0.31905 | 0.25699
0:45 | 0.18483 | 0.26725 | 0.31901 | 0.25703
0:45 | 0.18635 | 0.26731 | 0.3191 | 0.25759
0:45 | 0.18774 | 0.26731 | 0.31934 | 0.25813
0:45 | 0.18627 | 0.26736 | 0.3194 | 0.25768
0:45 | 0.19304 | 0.26739 | 0.31956 | 0.26
0:45 | 0.20098 | 0.26736 | 0.31982 | 0.26272
0:45 | 0.19339 | 0.26741 | 0.31959 | 0.26013
0:46 | 0.19078 | 0.26715 | 0.3198 | 0.25924
0:46 | 0.18365 | 0.26718 | 0.3199 | 0.25691
0:46 | 0.1885 | 0.26707 | 0.31981 | 0.25846
0:46 | 0.1985 | 0.26718 | 0.31995 | 0.26188
0:46 | 0.21046 | 0.26717 | 0.31999 | 0.26587
0:46 | 0.20719 | 0.26728 | 0.32001 | 0.26483
0:46 | 0.20605 | 0.26745 | 0.32027 | 0.26459
0:46 | 0.20616 | 0.26769 | 0.32051 | 0.26479
0:46 | 0.20606 | 0.26765 | 0.3208 | 0.26484
0:46 | 0.20584 | 0.26755 | 0.32068 | 0.26469
0:46 | 0.20572 | 0.2675 | 0.32087 | 0.2647
0:46 | 0.20554 | 0.2675 | 0.32125 | 0.26476
0:47 | 0.20535 | 0.26746 | 0.32161 | 0.26481
0:47 | 0.20513 | 0.2674 | 0.32176 | 0.26477
0:47 | 0.20504 | 0.26731 | 0.32204 | 0.2648
0:47 | 0.20497 | 0.26738 | 0.32216 | 0.26484
0:47 | 0.20484 | 0.26739 | 0.32204 | 0.26476
0:47 | 0.2047 | 0.2673 | 0.32211 | 0.2647
0:47 | 0.20463 | 0.26731 | 0.3224 | 0.26478
0:47 | 0.20451 | 0.26727 | 0.32229 | 0.26469
0:47 | 0.20464 | 0.26721 | 0.32235 | 0.26473
0:47 | 0.20463 | 0.26717 | 0.32246 | 0.26475
0:47 | 0.20468 | 0.26687 | 0.32245 | 0.26466
0:47 | 0.20454 | 0.26686 | 0.32233 | 0.26458
0:48 | 0.20445 | 0.26681 | 0.32235 | 0.26453
0:48 | 0.20452 | 0.26677 | 0.32237 | 0.26455
0:48 | 0.20459 | 0.26689 | 0.32265 | 0.26471
0:48 | 0.20453 | 0.26682 | 0.3224 | 0.26458
0:48 | 0.20458 | 0.26669 | 0.32263 | 0.26463
0:48 | 0.20452 | 0.26664 | 0.32236 | 0.26451
0:48 | 0.20436 | 0.26639 | 0.32262 | 0.26446
0:48 | 0.20441 | 0.26655 | 0.32252 | 0.26449
0:48 | 0.20424 | 0.26652 | 0.32249 | 0.26442
0:48 | 0.20432 | 0.2666 | 0.32241 | 0.26444
0:48 | 0.20439 | 0.26674 | 0.3225 | 0.26454
0:48 | 0.2042 | 0.26664 | 0.32245 | 0.26443
0:49 | 0.20426 | 0.26672 | 0.32259 | 0.26452
0:49 | 0.20427 | 0.26675 | 0.32256 | 0.26453
0:49 | 0.20422 | 0.26685 | 0.32264 | 0.26457
0:49 | 0.20425 | 0.2667 | 0.32261 | 0.26452
0:49 | 0.2043 | 0.26674 | 0.32265 | 0.26457
0:49 | 0.20419 | 0.26669 | 0.32271 | 0.26453
0:49 | 0.20413 | 0.26686 | 0.32268 | 0.26456
0:49 | 0.20408 | 0.26694 | 0.32263 | 0.26455
0:49 | 0.20409 | 0.26702 | 0.32259 | 0.26457
0:49 | 0.2042 | 0.26683 | 0.32272 | 0.26458
0:49 | 0.20418 | 0.2669 | 0.32276 | 0.26461
0:49 | 0.20418 | 0.26689 | 0.32273 | 0.2646
0:50 | 0.20424 | 0.26688 | 0.32269 | 0.2646
0:50 | 0.20417 | 0.26684 | 0.32244 | 0.26448
0:50 | 0.20417 | 0.26676 | 0.32255 | 0.26449
0:50 | 0.20421 | 0.26683 | 0.32244 | 0.26449
0:50 | 0.20418 | 0.26665 | 0.32242 | 0.26442
0:50 | 0.20417 | 0.2665 | 0.32222 | 0.2643
0:50 | 0.20416 | 0.26666 | 0.32251 | 0.26444
0:50 | 0.20419 | 0.26676 | 0.32226 | 0.2644
0:50 | 0.20425 | 0.26656 | 0.32231 | 0.26437
0:50 | 0.2042 | 0.2668 | 0.32216 | 0.26439
0:50 | 0.2042 | 0.26694 | 0.322 | 0.26438
0:50 | 0.20418 | 0.26683 | 0.32211 | 0.26437
0:50 | 0.20419 | 0.26686 | 0.32214 | 0.2644
0:51 | 0.20425 | 0.26683 | 0.32221 | 0.26443
0:51 | 0.20422 | 0.26691 | 0.32199 | 0.26438
0:51 | 0.20425 | 0.26695 | 0.32208 | 0.26443
0:51 | 0.20437 | 0.26693 | 0.32166 | 0.26432
0:51 | 0.20432 | 0.26683 | 0.32147 | 0.26421
0:51 | 0.20433 | 0.2667 | 0.321 | 0.26401
0:51 | 0.20428 | 0.26689 | 0.32095 | 0.26404
0:51 | 0.20426 | 0.26681 | 0.32031 | 0.26379
0:51 | 0.20428 | 0.26661 | 0.31975 | 0.26355
0:51 | 0.20425 | 0.26668 | 0.31949 | 0.26347
0:51 | 0.20422 | 0.26665 | 0.31926 | 0.26338
0:51 | 0.20412 | 0.26675 | 0.31858 | 0.26315
0:52 | 0.2041 | 0.26644 | 0.31865 | 0.26306
0:52 | 0.20403 | 0.26667 | 0.31841 | 0.26304
0:52 | 0.20411 | 0.26691 | 0.31798 | 0.263
0:52 | 0.20413 | 0.26704 | 0.31798 | 0.26305
0:52 | 0.20414 | 0.26714 | 0.31775 | 0.26301
0:52 | 0.20421 | 0.26729 | 0.31763 | 0.26304
0:52 | 0.20419 | 0.26722 | 0.31746 | 0.26296
0:52 | 0.20422 | 0.26713 | 0.31737 | 0.26291
0:52 | 0.20367 | 0.26733 | 0.31735 | 0.26278
0:52 | 0.2018 | 0.26735 | 0.31732 | 0.26216
0:52 | 0.20276 | 0.26727 | 0.31729 | 0.26244
0:52 | 0.2037 | 0.26721 | 0.31716 | 0.26269
0:53 | 0.2038 | 0.26729 | 0.31708 | 0.26272
0:53 | 0.2039 | 0.26744 | 0.31699 | 0.26278
0:53 | 0.20405 | 0.26724 | 0.3171 | 0.2628
0:53 | 0.20416 | 0.26717 | 0.31711 | 0.26281
0:53 | 0.20428 | 0.26718 | 0.31706 | 0.26284
0:53 | 0.20433 | 0.26724 | 0.31745 | 0.26301
0:53 | 0.20433 | 0.26747 | 0.31747 | 0.26309
0:53 | 0.20425 | 0.26718 | 0.31722 | 0.26288
0:53 | 0.20433 | 0.26721 | 0.3176 | 0.26305
0:53 | 0.20436 | 0.26751 | 0.31791 | 0.26326
0:53 | 0.20452 | 0.26738 | 0.31784 | 0.26325
0:53 | 0.2045 | 0.26689 | 0.31774 | 0.26304
0:54 | 0.20448 | 0.26733 | 0.31768 | 0.26316
0:54 | 0.20443 | 0.26754 | 0.31749 | 0.26315
0:54 | 0.20448 | 0.26749 | 0.31738 | 0.26312
0:54 | 0.20451 | 0.26759 | 0.31714 | 0.26308
0:54 | 0.20474 | 0.2677 | 0.31694 | 0.26312
0:54 | 0.20465 | 0.26762 | 0.31683 | 0.26303
0:54 | 0.20477 | 0.2677 | 0.31657 | 0.26302
0:54 | 0.20482 | 0.26756 | 0.31664 | 0.26301
0:54 | 0.20476 | 0.26757 | 0.31674 | 0.26302
0:54 | 0.20484 | 0.26775 | 0.31682 | 0.26314
0:54 | 0.20483 | 0.26788 | 0.31673 | 0.26314
0:54 | 0.20492 | 0.26768 | 0.31694 | 0.26318
0:55 | 0.20491 | 0.26781 | 0.31738 | 0.26336
0:55 | 0.20493 | 0.26769 | 0.31727 | 0.26329
0:55 | 0.20498 | 0.26775 | 0.31749 | 0.2634
0:55 | 0.20499 | 0.26785 | 0.31793 | 0.26359
0:55 | 0.20522 | 0.26783 | 0.31819 | 0.26374
0:55 | 0.20528 | 0.26795 | 0.31856 | 0.26393
0:55 | 0.2054 | 0.26807 | 0.31887 | 0.26411
0:55 | 0.20546 | 0.26789 | 0.31944 | 0.26426
0:55 | 0.20543 | 0.26806 | 0.31972 | 0.2644
0:55 | 0.20529 | 0.26799 | 0.32029 | 0.26452
0:55 | 0.20541 | 0.26809 | 0.32061 | 0.2647
0:55 | 0.20533 | 0.26829 | 0.32114 | 0.26492
0:56 | 0.20548 | 0.26805 | 0.32115 | 0.26489
0:56 | 0.20548 | 0.26807 | 0.32146 | 0.26501
0:56 | 0.20541 | 0.26823 | 0.3218 | 0.26515
0:56 | 0.20562 | 0.26829 | 0.32169 | 0.2652
0:56 | 0.20559 | 0.26829 | 0.32206 | 0.26531
0:56 | 0.20565 | 0.2682 | 0.32156 | 0.26513
0:56 | 0.20567 | 0.26817 | 0.32112 | 0.26499
0:56 | 0.20569 | 0.2679 | 0.32058 | 0.26472
0:56 | 0.20562 | 0.26813 | 0.32027 | 0.26467
0:56 | 0.20561 | 0.26821 | 0.32015 | 0.26466
0:56 | 0.2057 | 0.26823 | 0.31979 | 0.26457
0:56 | 0.20564 | 0.26831 | 0.31987 | 0.26461
0:57 | 0.20576 | 0.2684 | 0.31988 | 0.26468
0:57 | 0.20573 | 0.26841 | 0.31997 | 0.26471
0:57 | 0.20574 | 0.26855 | 0.31999 | 0.26476
0:57 | 0.20584 | 0.26848 | 0.3202 | 0.26484
0:57 | 0.20564 | 0.26851 | 0.32015 | 0.26477
0:57 | 0.20556 | 0.26847 | 0.32027 | 0.26476
0:57 | 0.20564 | 0.26854 | 0.32043 | 0.26487
0:57 | 0.20552 | 0.26855 | 0.32003 | 0.2647
0:57 | 0.20553 | 0.26867 | 0.32027 | 0.26482
0:57 | 0.20552 | 0.26866 | 0.32043 | 0.26487
0:57 | 0.20546 | 0.26845 | 0.32023 | 0.26471
0:57 | 0.20545 | 0.26846 | 0.32008 | 0.26466
0:58 | 0.20537 | 0.2683 | 0.32004 | 0.26457
0:58 | 0.2056 | 0.26833 | 0.31987 | 0.2646
0:58 | 0.20559 | 0.26837 | 0.31968 | 0.26454
0:58 | 0.20558 | 0.26858 | 0.31932 | 0.26449
0:58 | 0.20546 | 0.26853 | 0.3192 | 0.2644
0:58 | 0.20553 | 0.26856 | 0.31918 | 0.26442
0:58 | 0.20545 | 0.2687 | 0.31899 | 0.26438
0:58 | 0.20549 | 0.26864 | 0.31882 | 0.26432
0:58 | 0.20555 | 0.2686 | 0.31868 | 0.26428
0:58 | 0.20552 | 0.26866 | 0.31847 | 0.26422
0:58 | 0.20551 | 0.26866 | 0.31874 | 0.2643
0:58 | 0.20546 | 0.2687 | 0.3185 | 0.26422
0:59 | 0.20538 | 0.269 | 0.31849 | 0.26429
0:59 | 0.20553 | 0.26899 | 0.31856 | 0.26436
0:59 | 0.20547 | 0.26901 | 0.31858 | 0.26435
0:59 | 0.20549 | 0.26899 | 0.31892 | 0.26447
0:59 | 0.20543 | 0.26883 | 0.31911 | 0.26446
0:59 | 0.20545 | 0.26878 | 0.31874 | 0.26432
0:59 | 0.20549 | 0.26875 | 0.31855 | 0.26427
0:59 | 0.20553 | 0.26865 | 0.31852 | 0.26423
0:59 | 0.20554 | 0.26857 | 0.3185 | 0.2642
0:59 | 0.20561 | 0.26849 | 0.31886 | 0.26432
0:59 | 0.20543 | 0.26848 | 0.31921 | 0.26437
0:59 | 0.20543 | 0.26864 | 0.31934 | 0.26447
1:00 | 0.20546 | 0.26865 | 0.31954 | 0.26455
1:00 | 0.20545 | 0.26856 | 0.32016 | 0.26472
1:00 | 0.20544 | 0.26872 | 0.32049 | 0.26488
1:00 | 0.20555 | 0.26882 | 0.32034 | 0.2649
1:00 | 0.20545 | 0.26897 | 0.32051 | 0.26498
1:00 | 0.20547 | 0.26894 | 0.32103 | 0.26515
1:00 | 0.20541 | 0.26913 | 0.32109 | 0.26521
1:00 | 0.2055 | 0.26916 | 0.32104 | 0.26523
1:00 | 0.20554 | 0.26908 | 0.32087 | 0.26517
1:00 | 0.2055 | 0.26917 | 0.32086 | 0.26518
1:00 | 0.20542 | 0.26915 | 0.32129 | 0.26529
1:00 | 0.2056 | 0.26943 | 0.32112 | 0.26539
1:01 | 0.20552 | 0.26928 | 0.3213 | 0.26537
1:01 | 0.20574 | 0.26935 | 0.32133 | 0.26547
1:01 | 0.20577 | 0.2695 | 0.32148 | 0.26558
1:01 | 0.20576 | 0.26953 | 0.32162 | 0.26564
1:01 | 0.2059 | 0.26938 | 0.32138 | 0.26556
1:01 | 0.20606 | 0.26955 | 0.32107 | 0.26556
1:01 | 0.2062 | 0.26931 | 0.32128 | 0.2656
1:01 | 0.20634 | 0.2695 | 0.32148 | 0.26578
1:01 | 0.20636 | 0.26957 | 0.3218 | 0.26591
1:01 | 0.20645 | 0.26954 | 0.32169 | 0.26589
1:01 | 0.20657 | 0.26933 | 0.32201 | 0.26597
1:01 | 0.20654 | 0.26942 | 0.32188 | 0.26594
1:02 | 0.20668 | 0.26941 | 0.32193 | 0.26601
1:02 | 0.20684 | 0.26932 | 0.32189 | 0.26602
1:02 | 0.20678 | 0.26944 | 0.32193 | 0.26605
1:02 | 0.20698 | 0.26921 | 0.32186 | 0.26602
1:02 | 0.20705 | 0.26912 | 0.32178 | 0.26598
1:02 | 0.20697 | 0.26904 | 0.32184 | 0.26595
1:02 | 0.20715 | 0.2691 | 0.32172 | 0.26599
1:02 | 0.20715 | 0.26909 | 0.32182 | 0.26602
1:02 | 0.20705 | 0.26903 | 0.32184 | 0.26597
1:02 | 0.20638 | 0.26907 | 0.32208 | 0.26584
1:02 | 0.20429 | 0.26886 | 0.32178 | 0.26498
1:02 | 0.19909 | 0.26896 | 0.32187 | 0.2633
1:03 | 0.19082 | 0.26907 | 0.32182 | 0.26057
1:03 | 0.18546 | 0.26889 | 0.322 | 0.25878
1:03 | 0.18507 | 0.26885 | 0.32222 | 0.25871
1:03 | 0.18514 | 0.26881 | 0.32228 | 0.25874
1:03 | 0.18532 | 0.26867 | 0.32221 | 0.25873
1:03 | 0.18555 | 0.26864 | 0.32242 | 0.25887
1:03 | 0.18543 | 0.2684 | 0.32258 | 0.2588
1:03 | 0.18527 | 0.26836 | 0.32254 | 0.25873
1:03 | 0.18546 | 0.26854 | 0.32252 | 0.25884
1:03 | 0.18588 | 0.26862 | 0.32259 | 0.25903
1:03 | 0.18696 | 0.26845 | 0.32248 | 0.2593
1:03 | 0.18676 | 0.26837 | 0.32258 | 0.25923
1:04 | 0.18671 | 0.26855 | 0.32257 | 0.25928
1:04 | 0.18676 | 0.26861 | 0.32243 | 0.25927
1:04 | 0.19895 | 0.26866 | 0.32237 | 0.26333
1:04 | 0.20809 | 0.26887 | 0.3227 | 0.26655
1:04 | 0.20393 | 0.26851 | 0.32268 | 0.26504
1:04 | 0.20299 | 0.26862 | 0.32278 | 0.2648
1:04 | 0.20845 | 0.26866 | 0.32294 | 0.26668
1:04 | 0.20782 | 0.26868 | 0.32294 | 0.26648
1:04 | 0.20763 | 0.26869 | 0.32321 | 0.26651
1:04 | 0.20745 | 0.26871 | 0.32334 | 0.2665
1:04 | 0.20743 | 0.26862 | 0.32332 | 0.26646
1:04 | 0.20715 | 0.26862 | 0.32325 | 0.26634
1:05 | 0.20709 | 0.26859 | 0.32329 | 0.26633
1:05 | 0.20693 | 0.26857 | 0.32343 | 0.26631
1:05 | 0.20692 | 0.26848 | 0.3234 | 0.26627
1:05 | 0.20681 | 0.26858 | 0.32344 | 0.26628
1:05 | 0.20666 | 0.26865 | 0.32341 | 0.26624
1:05 | 0.20659 | 0.26851 | 0.32333 | 0.26614
1:05 | 0.2064 | 0.26872 | 0.3236 | 0.26624
1:05 | 0.2064 | 0.26856 | 0.32349 | 0.26615
1:05 | 0.20625 | 0.26849 | 0.32358 | 0.26611
1:05 | 0.20629 | 0.26863 | 0.32377 | 0.26623
1:05 | 0.20616 | 0.26862 | 0.3238 | 0.26619
1:05 | 0.20613 | 0.26869 | 0.32393 | 0.26625
1:06 | 0.20596 | 0.26864 | 0.32413 | 0.26624
1:06 | 0.20579 | 0.26858 | 0.32407 | 0.26615
1:06 | 0.20571 | 0.26857 | 0.32385 | 0.26604
1:06 | 0.20556 | 0.26874 | 0.32413 | 0.26614
1:06 | 0.20549 | 0.26881 | 0.32403 | 0.26611
1:06 | 0.20536 | 0.26845 | 0.32391 | 0.26591
1:06 | 0.20517 | 0.26859 | 0.32388 | 0.26588
1:06 | 0.20519 | 0.26868 | 0.32412 | 0.266
1:06 | 0.20504 | 0.26858 | 0.32387 | 0.26583
1:06 | 0.2049 | 0.26853 | 0.32371 | 0.26571
1:06 | 0.20487 | 0.26861 | 0.32365 | 0.26571
1:06 | 0.20469 | 0.26859 | 0.32345 | 0.26558
1:07 | 0.2045 | 0.26852 | 0.32354 | 0.26552
1:07 | 0.20459 | 0.26859 | 0.32345 | 0.26554
1:07 | 0.20447 | 0.26858 | 0.32312 | 0.26539
1:07 | 0.20437 | 0.26863 | 0.3232 | 0.2654
1:07 | 0.20447 | 0.26863 | 0.32319 | 0.26543
1:07 | 0.20421 | 0.26856 | 0.32326 | 0.26534
1:07 | 0.20415 | 0.26847 | 0.32317 | 0.26526
1:07 | 0.2041 | 0.26859 | 0.32324 | 0.26531
1:07 | 0.20407 | 0.26837 | 0.32355 | 0.26533
1:07 | 0.20394 | 0.26847 | 0.32361 | 0.26534
1:07 | 0.20386 | 0.26856 | 0.32336 | 0.26526
1:07 | 0.20389 | 0.26837 | 0.32329 | 0.26518
1:08 | 0.20389 | 0.26821 | 0.32314 | 0.26508
1:08 | 0.20384 | 0.26808 | 0.32297 | 0.26496
1:08 | 0.20391 | 0.268 | 0.32311 | 0.26501
1:08 | 0.20376 | 0.2681 | 0.32286 | 0.26491
1:08 | 0.20369 | 0.26753 | 0.32275 | 0.26466
1:08 | 0.20365 | 0.26709 | 0.32283 | 0.26453
1:08 | 0.20367 | 0.26648 | 0.32284 | 0.26433
1:08 | 0.20374 | 0.26673 | 0.32268 | 0.26438
1:08 | 0.20376 | 0.26697 | 0.32271 | 0.26448
1:08 | 0.20371 | 0.26677 | 0.32254 | 0.26434
1:08 | 0.20373 | 0.26694 | 0.32264 | 0.26444
1:08 | 0.20356 | 0.26754 | 0.32264 | 0.26458
1:09 | 0.20344 | 0.26753 | 0.32247 | 0.26448
1:09 | 0.20352 | 0.26754 | 0.32229 | 0.26445
1:09 | 0.20354 | 0.26756 | 0.32235 | 0.26448
1:09 | 0.2035 | 0.26757 | 0.32237 | 0.26448
1:09 | 0.20374 | 0.26771 | 0.32216 | 0.26454
1:09 | 0.20359 | 0.26769 | 0.32208 | 0.26445
1:09 | 0.20366 | 0.26767 | 0.3219 | 0.26441
1:09 | 0.20379 | 0.26782 | 0.32145 | 0.26435
1:09 | 0.20386 | 0.26781 | 0.32141 | 0.26436
1:09 | 0.2038 | 0.26767 | 0.32103 | 0.26417
1:09 | 0.20387 | 0.26767 | 0.32051 | 0.26402
1:09 | 0.20385 | 0.26777 | 0.32046 | 0.26403
1:10 | 0.20387 | 0.26768 | 0.32024 | 0.26393
1:10 | 0.20391 | 0.26786 | 0.32008 | 0.26395
1:10 | 0.20391 | 0.26791 | 0.31975 | 0.26386
1:10 | 0.20399 | 0.26797 | 0.31988 | 0.26395
1:10 | 0.20416 | 0.26803 | 0.31965 | 0.26395
1:10 | 0.20407 | 0.26825 | 0.31989 | 0.26407
1:10 | 0.20408 | 0.26828 | 0.31968 | 0.26401
1:10 | 0.20408 | 0.26797 | 0.31938 | 0.26381
1:10 | 0.20421 | 0.26784 | 0.31914 | 0.26373
1:10 | 0.20426 | 0.26775 | 0.31885 | 0.26362
1:10 | 0.20419 | 0.26786 | 0.31862 | 0.26356
1:10 | 0.20409 | 0.26795 | 0.31823 | 0.26342
1:11 | 0.20405 | 0.26783 | 0.31755 | 0.26315
1:11 | 0.2041 | 0.26795 | 0.31748 | 0.26318
1:11 | 0.20425 | 0.26804 | 0.31731 | 0.2632
1:11 | 0.20425 | 0.26795 | 0.31715 | 0.26312
1:11 | 0.20402 | 0.26781 | 0.31706 | 0.26296
1:11 | 0.20404 | 0.26802 | 0.31671 | 0.26292
1:11 | 0.20409 | 0.26784 | 0.31728 | 0.26307
1:11 | 0.20408 | 0.26793 | 0.31714 | 0.26305
1:11 | 0.20429 | 0.26804 | 0.31709 | 0.26314
1:11 | 0.20412 | 0.26809 | 0.3168 | 0.26301
1:11 | 0.20416 | 0.26798 | 0.31697 | 0.26304
1:11 | 0.20428 | 0.268 | 0.31737 | 0.26322
1:12 | 0.20416 | 0.26805 | 0.31722 | 0.26314
1:12 | 0.2044 | 0.26794 | 0.31708 | 0.26314
1:12 | 0.20429 | 0.26797 | 0.31693 | 0.26306
1:12 | 0.20417 | 0.26808 | 0.31672 | 0.26299
1:12 | 0.20421 | 0.2679 | 0.31703 | 0.26305
1:12 | 0.20417 | 0.2679 | 0.31678 | 0.26295
1:12 | 0.20401 | 0.2679 | 0.31711 | 0.26301
1:12 | 0.20432 | 0.26814 | 0.31638 | 0.26294
1:12 | 0.20417 | 0.26806 | 0.31686 | 0.26303
1:12 | 0.20414 | 0.26807 | 0.31659 | 0.26293
1:12 | 0.20426 | 0.26806 | 0.31639 | 0.2629
1:12 | 0.20428 | 0.26802 | 0.3162 | 0.26283
1:13 | 0.20419 | 0.26811 | 0.31636 | 0.26289
1:13 | 0.20414 | 0.26829 | 0.31641 | 0.26295
1:13 | 0.20413 | 0.26834 | 0.31627 | 0.26291
1:13 | 0.20404 | 0.2684 | 0.31626 | 0.2629
1:13 | 0.20406 | 0.26836 | 0.31626 | 0.26289
1:13 | 0.20408 | 0.26828 | 0.31641 | 0.26292
1:13 | 0.20411 | 0.26855 | 0.31629 | 0.26298
1:13 | 0.20406 | 0.26848 | 0.31626 | 0.26293
1:13 | 0.2041 | 0.26857 | 0.31607 | 0.26291
1:13 | 0.2039 | 0.26856 | 0.31612 | 0.26286
1:13 | 0.20392 | 0.26858 | 0.31609 | 0.26286
1:13 | 0.20378 | 0.26866 | 0.31655 | 0.263
1:14 | 0.20387 | 0.26863 | 0.31642 | 0.26297
1:14 | 0.20387 | 0.26852 | 0.31644 | 0.26294
1:14 | 0.20392 | 0.2686 | 0.31639 | 0.26297
1:14 | 0.20393 | 0.26857 | 0.31636 | 0.26295
1:14 | 0.20397 | 0.26843 | 0.31626 | 0.26289
1:14 | 0.20401 | 0.26865 | 0.3166 | 0.26309
1:14 | 0.20404 | 0.26859 | 0.31642 | 0.26302
1:14 | 0.20401 | 0.26854 | 0.31659 | 0.26305
1:14 | 0.20412 | 0.26871 | 0.31664 | 0.26316
1:14 | 0.20408 | 0.26843 | 0.31676 | 0.26309
1:14 | 0.20406 | 0.26851 | 0.31675 | 0.26311
1:14 | 0.20408 | 0.26846 | 0.31671 | 0.26308
1:15 | 0.20409 | 0.26857 | 0.3168 | 0.26315
1:15 | 0.20409 | 0.26871 | 0.31712 | 0.2633
1:15 | 0.20422 | 0.26867 | 0.31655 | 0.26315
1:15 | 0.20438 | 0.2688 | 0.31662 | 0.26327
1:15 | 0.20451 | 0.26871 | 0.31664 | 0.26329
1:15 | 0.20454 | 0.26888 | 0.31647 | 0.2633
1:15 | 0.2046 | 0.26895 | 0.31652 | 0.26336
1:15 | 0.20453 | 0.26878 | 0.3163 | 0.26321
1:15 | 0.20452 | 0.2689 | 0.31658 | 0.26333
1:15 | 0.20468 | 0.26902 | 0.31663 | 0.26344
1:15 | 0.20468 | 0.26898 | 0.31659 | 0.26341
1:15 | 0.20472 | 0.26914 | 0.31716 | 0.26367
1:16 | 0.20469 | 0.26911 | 0.31685 | 0.26355
1:16 | 0.20475 | 0.26915 | 0.317 | 0.26363
1:16 | 0.20483 | 0.2691 | 0.31702 | 0.26365
1:16 | 0.20487 | 0.26905 | 0.317 | 0.26364
1:16 | 0.2049 | 0.26911 | 0.31683 | 0.26362
1:16 | 0.20486 | 0.26923 | 0.31672 | 0.2636
1:16 | 0.20496 | 0.26937 | 0.31654 | 0.26362
1:16 | 0.20495 | 0.26936 | 0.31666 | 0.26366
1:16 | 0.20513 | 0.26941 | 0.31633 | 0.26362
1:16 | 0.20517 | 0.26939 | 0.31623 | 0.2636
1:16 | 0.2054 | 0.26941 | 0.31647 | 0.26376
1:16 | 0.20543 | 0.26956 | 0.31615 | 0.26371
1:17 | 0.20545 | 0.26957 | 0.31642 | 0.26381
1:17 | 0.20539 | 0.26955 | 0.31642 | 0.26378
1:17 | 0.20537 | 0.26942 | 0.31651 | 0.26377
1:17 | 0.20529 | 0.26959 | 0.31632 | 0.26373
1:17 | 0.20552 | 0.26948 | 0.31628 | 0.26376
1:17 | 0.20569 | 0.26962 | 0.31611 | 0.26381
1:17 | 0.2058 | 0.26972 | 0.31654 | 0.26402
1:17 | 0.20577 | 0.26964 | 0.31606 | 0.26382
1:17 | 0.2057 | 0.26968 | 0.3163 | 0.26389
1:17 | 0.20578 | 0.26973 | 0.31622 | 0.26391
1:17 | 0.20569 | 0.26966 | 0.31626 | 0.26387
1:17 | 0.2057 | 0.26958 | 0.31617 | 0.26382
1:18 | 0.20566 | 0.26973 | 0.31662 | 0.264
1:18 | 0.20575 | 0.26957 | 0.31653 | 0.26395
1:18 | 0.20558 | 0.26914 | 0.31626 | 0.26366
1:18 | 0.20558 | 0.26227 | 0.31712 | 0.26166
1:18 | 0.2056 | 0.26922 | 0.317 | 0.26394
1:18 | 0.20573 | 0.26953 | 0.31704 | 0.2641
1:18 | 0.2057 | 0.26959 | 0.31725 | 0.26418
1:18 | 0.20588 | 0.26967 | 0.31738 | 0.26431
1:18 | 0.20611 | 0.26981 | 0.31787 | 0.2646
1:18 | 0.20623 | 0.2699 | 0.31797 | 0.2647
1:18 | 0.20646 | 0.26979 | 0.31803 | 0.26476
1:18 | 0.20658 | 0.26996 | 0.31797 | 0.26484
1:19 | 0.20694 | 0.26982 | 0.31831 | 0.26502
1:19 | 0.20685 | 0.26992 | 0.31858 | 0.26512
1:19 | 0.20681 | 0.2697 | 0.31877 | 0.26509
1:19 | 0.20627 | 0.26986 | 0.31909 | 0.26508
1:19 | 0.20406 | 0.26986 | 0.31907 | 0.26433
1:19 | 0.19958 | 0.26969 | 0.31924 | 0.26284
1:19 | 0.20527 | 0.26989 | 0.31982 | 0.26499
1:19 | 0.19612 | 0.2696 | 0.31983 | 0.26185
1:19 | 0.19541 | 0.26962 | 0.32028 | 0.26177
1:19 | 0.20142 | 0.26971 | 0.32057 | 0.2639
1:19 | 0.19665 | 0.26965 | 0.32088 | 0.26239
1:19 | 0.19807 | 0.26974 | 0.32079 | 0.26287
1:20 | 0.18727 | 0.26975 | 0.32141 | 0.25947
1:20 | 0.18103 | 0.26976 | 0.32152 | 0.25744
1:20 | 0.18276 | 0.26969 | 0.32172 | 0.25805
1:20 | 0.1821 | 0.26969 | 0.3224 | 0.25806
1:20 | 0.18215 | 0.26979 | 0.32248 | 0.25814
1:20 | 0.18254 | 0.26956 | 0.32322 | 0.25844
1:20 | 0.18246 | 0.2693 | 0.32315 | 0.2583
1:20 | 0.18229 | 0.26951 | 0.32368 | 0.25849
1:20 | 0.18221 | 0.26963 | 0.32393 | 0.25859
1:20 | 0.18192 | 0.26961 | 0.32398 | 0.2585
1:20 | 0.18214 | 0.2691 | 0.32409 | 0.25844
1:20 | 0.18219 | 0.26942 | 0.32459 | 0.25873
1:21 | 0.18279 | 0.26956 | 0.32433 | 0.25889
1:21 | 0.1831 | 0.26955 | 0.32496 | 0.25921
1:21 | 0.18314 | 0.2696 | 0.32486 | 0.2592
1:21 | 0.18363 | 0.26954 | 0.32511 | 0.25942
1:21 | 0.18389 | 0.26943 | 0.3253 | 0.25954
1:21 | 0.18428 | 0.26962 | 0.32513 | 0.25968
1:21 | 0.18464 | 0.26971 | 0.3254 | 0.25992
1:21 | 0.18481 | 0.26973 | 0.32552 | 0.26002
1:21 | 0.18467 | 0.26984 | 0.32538 | 0.25997
1:21 | 0.18475 | 0.26981 | 0.32571 | 0.26009
1:21 | 0.18431 | 0.2696 | 0.32597 | 0.25996
1:21 | 0.18446 | 0.26975 | 0.32607 | 0.2601
1:22 | 0.18813 | 0.26958 | 0.32623 | 0.26131
1:22 | 0.20225 | 0.26968 | 0.32637 | 0.2661
1:22 | 0.20903 | 0.26979 | 0.32632 | 0.26838
1:22 | 0.20812 | 0.26988 | 0.32616 | 0.26805
1:22 | 0.20707 | 0.27015 | 0.32618 | 0.2678
1:22 | 0.20667 | 0.27014 | 0.3261 | 0.26764
1:22 | 0.20656 | 0.27022 | 0.32603 | 0.2676
1:22 | 0.20618 | 0.27008 | 0.32611 | 0.26746
1:22 | 0.20607 | 0.26996 | 0.3258 | 0.26728
1:22 | 0.20599 | 0.27009 | 0.3255 | 0.26719
1:22 | 0.20575 | 0.27002 | 0.32537 | 0.26704
1:22 | 0.20549 | 0.2701 | 0.3253 | 0.26696
1:23 | 0.20524 | 0.27014 | 0.32529 | 0.26689
1:23 | 0.20486 | 0.27007 | 0.32501 | 0.26665
1:23 | 0.20366 | 0.27014 | 0.32491 | 0.26624
1:23 | 0.20578 | 0.27004 | 0.32444 | 0.26675
1:23 | 0.20563 | 0.27006 | 0.32404 | 0.26658
1:23 | 0.20533 | 0.27007 | 0.32353 | 0.26631
1:23 | 0.20536 | 0.27005 | 0.32261 | 0.26601
1:23 | 0.20535 | 0.27007 | 0.32224 | 0.26589
1:23 | 0.205 | 0.27 | 0.32171 | 0.26557
1:23 | 0.20456 | 0.27001 | 0.3217 | 0.26542
1:23 | 0.2018 | 0.27014 | 0.32164 | 0.26453
1:23 | 0.20066 | 0.27011 | 0.32167 | 0.26415
1:24 | 0.20106 | 0.27006 | 0.32159 | 0.26424
1:24 | 0.20228 | 0.27002 | 0.32168 | 0.26466
1:24 | 0.19976 | 0.27009 | 0.3214 | 0.26375
1:24 | 0.20164 | 0.27018 | 0.32185 | 0.26456
1:24 | 0.20382 | 0.26995 | 0.32203 | 0.26527
1:24 | 0.20518 | 0.26998 | 0.32207 | 0.26574
1:24 | 0.20275 | 0.27 | 0.32232 | 0.26502
1:24 | 0.20472 | 0.27003 | 0.32273 | 0.26583
1:24 | 0.20511 | 0.27004 | 0.32287 | 0.26601
1:24 | 0.2048 | 0.27013 | 0.32325 | 0.26606
1:24 | 0.20488 | 0.2701 | 0.32376 | 0.26625
1:24 | 0.20494 | 0.2698 | 0.32415 | 0.2663
1:25 | 0.20476 | 0.26982 | 0.32449 | 0.26636
1:25 | 0.20465 | 0.26995 | 0.32502 | 0.26654
1:25 | 0.20463 | 0.27006 | 0.32535 | 0.26668
1:25 | 0.20475 | 0.26989 | 0.32569 | 0.26678
1:25 | 0.20481 | 0.27004 | 0.32634 | 0.26707
1:25 | 0.2047 | 0.27012 | 0.32659 | 0.26714
1:25 | 0.20468 | 0.27023 | 0.32679 | 0.26723
1:25 | 0.20462 | 0.27028 | 0.32709 | 0.26733
1:25 | 0.20481 | 0.27047 | 0.32743 | 0.26757
1:25 | 0.20484 | 0.27043 | 0.32739 | 0.26755
1:25 | 0.2048 | 0.27035 | 0.32755 | 0.26757
1:25 | 0.20492 | 0.27032 | 0.32766 | 0.26763
1:26 | 0.20493 | 0.27036 | 0.32739 | 0.26756
1:26 | 0.20502 | 0.27039 | 0.3275 | 0.26763
1:26 | 0.20497 | 0.27038 | 0.32763 | 0.26766
1:26 | 0.20501 | 0.27044 | 0.32759 | 0.26768
1:26 | 0.20504 | 0.27036 | 0.32755 | 0.26765
1:26 | 0.2051 | 0.26998 | 0.32763 | 0.26757
1:26 | 0.20509 | 0.26993 | 0.32776 | 0.26759
1:26 | 0.20494 | 0.26977 | 0.32807 | 0.2676
1:26 | 0.20487 | 0.27002 | 0.32806 | 0.26765
1:26 | 0.20491 | 0.27007 | 0.32829 | 0.26776
1:26 | 0.20511 | 0.26992 | 0.32836 | 0.2678
1:26 | 0.20496 | 0.27002 | 0.32842 | 0.2678
1:27 | 0.20502 | 0.2701 | 0.32833 | 0.26781
1:27 | 0.20519 | 0.27007 | 0.32853 | 0.26793
1:27 | 0.20531 | 0.26999 | 0.32859 | 0.26796
1:27 | 0.2053 | 0.27002 | 0.3285 | 0.26794
1:27 | 0.2052 | 0.2702 | 0.32849 | 0.26796
1:27 | 0.20519 | 0.27028 | 0.32834 | 0.26794
1:27 | 0.20527 | 0.27028 | 0.32799 | 0.26785
1:27 | 0.20534 | 0.27002 | 0.32811 | 0.26783
1:27 | 0.20535 | 0.27 | 0.32815 | 0.26784
1:27 | 0.20548 | 0.26988 | 0.32817 | 0.26784
1:27 | 0.20548 | 0.26989 | 0.32815 | 0.26784
1:27 | 0.20546 | 0.26991 | 0.32818 | 0.26785
1:28 | 0.20552 | 0.26968 | 0.32814 | 0.26778
1:28 | 0.20539 | 0.26989 | 0.328 | 0.26776
1:28 | 0.2054 | 0.26973 | 0.32815 | 0.26776
1:28 | 0.20528 | 0.26978 | 0.32789 | 0.26765
1:28 | 0.20449 | 0.26976 | 0.32772 | 0.26732
1:28 | 0.20494 | 0.26982 | 0.32783 | 0.26753
1:28 | 0.20556 | 0.26999 | 0.32804 | 0.26786
1:28 | 0.20517 | 0.26979 | 0.32782 | 0.26759
1:28 | 0.20514 | 0.26982 | 0.32785 | 0.26761
1:28 | 0.20525 | 0.26972 | 0.32801 | 0.26766
1:28 | 0.20533 | 0.26972 | 0.32783 | 0.26763
1:28 | 0.20539 | 0.26974 | 0.32781 | 0.26765
1:29 | 0.2055 | 0.26962 | 0.3276 | 0.26757
1:29 | 0.2055 | 0.26978 | 0.3275 | 0.26759
1:29 | 0.20529 | 0.26993 | 0.32711 | 0.26744
1:29 | 0.2053 | 0.26977 | 0.32694 | 0.26733
1:29 | 0.2053 | 0.26979 | 0.32666 | 0.26725
1:29 | 0.20532 | 0.26984 | 0.32684 | 0.26733
1:29 | 0.20514 | 0.27017 | 0.3264 | 0.26724
1:29 | 0.205 | 0.26988 | 0.32622 | 0.26703
1:29 | 0.20505 | 0.26978 | 0.32568 | 0.26684
1:29 | 0.20509 | 0.26978 | 0.32533 | 0.26674
1:29 | 0.20515 | 0.26965 | 0.32495 | 0.26658
1:29 | 0.20524 | 0.26985 | 0.32476 | 0.26662
1:30 | 0.20498 | 0.26993 | 0.32442 | 0.26644
1:30 | 0.20489 | 0.26981 | 0.32452 | 0.26641
1:30 | 0.20506 | 0.26994 | 0.32442 | 0.26648
1:30 | 0.20503 | 0.27037 | 0.32527 | 0.26689
1:30 | 0.20481 | 0.2702 | 0.32581 | 0.26694
1:30 | 0.20499 | 0.27018 | 0.32643 | 0.2672
1:30 | 0.20497 | 0.27014 | 0.32687 | 0.26733
1:30 | 0.20509 | 0.27006 | 0.32786 | 0.26767
1:30 | 0.20507 | 0.27006 | 0.32818 | 0.26777
1:30 | 0.20519 | 0.26994 | 0.3281 | 0.26774
1:30 | 0.20518 | 0.27015 | 0.32843 | 0.26792
1:30 | 0.20483 | 0.27003 | 0.32873 | 0.26787
1:31 | 0.20504 | 0.27028 | 0.32902 | 0.26812
1:31 | 0.20521 | 0.27009 | 0.32904 | 0.26811
1:31 | 0.20535 | 0.2702 | 0.32948 | 0.26834
1:31 | 0.20536 | 0.27037 | 0.32955 | 0.26843
1:31 | 0.20529 | 0.27019 | 0.32966 | 0.26838
1:31 | 0.20536 | 0.27026 | 0.3298 | 0.26848
1:31 | 0.20538 | 0.27013 | 0.32979 | 0.26843
1:31 | 0.20538 | 0.27002 | 0.3301 | 0.2685
1:31 | 0.20533 | 0.27015 | 0.33014 | 0.26854
1:31 | 0.20527 | 0.27001 | 0.3303 | 0.26853
1:31 | 0.20532 | 0.26985 | 0.33038 | 0.26851
1:31 | 0.20539 | 0.26951 | 0.3305 | 0.26847
1:32 | 0.20554 | 0.26952 | 0.33054 | 0.26853
1:32 | 0.20553 | 0.26969 | 0.33041 | 0.26854
1:32 | 0.20563 | 0.26969 | 0.33058 | 0.26863
1:32 | 0.20574 | 0.26963 | 0.33052 | 0.26863
1:32 | 0.20559 | 0.26999 | 0.33032 | 0.26863
1:32 | 0.2056 | 0.2699 | 0.33013 | 0.26854
1:32 | 0.20544 | 0.27012 | 0.33012 | 0.26856
1:32 | 0.20544 | 0.27028 | 0.32989 | 0.26854
1:32 | 0.20567 | 0.27029 | 0.32996 | 0.26864
1:32 | 0.2057 | 0.27038 | 0.32988 | 0.26866
1:32 | 0.20586 | 0.27054 | 0.32976 | 0.26872
1:32 | 0.20594 | 0.27055 | 0.32968 | 0.26872
1:33 | 0.20598 | 0.27072 | 0.32948 | 0.26873
1:33 | 0.20602 | 0.27073 | 0.32945 | 0.26873
1:33 | 0.20609 | 0.27087 | 0.32943 | 0.2688
1:33 | 0.20602 | 0.27071 | 0.32935 | 0.26869
1:33 | 0.20606 | 0.27068 | 0.32929 | 0.26868
1:33 | 0.20601 | 0.27065 | 0.32936 | 0.26867
1:33 | 0.20593 | 0.27063 | 0.32919 | 0.26858
1:33 | 0.20599 | 0.27064 | 0.32935 | 0.26866
1:33 | 0.20604 | 0.27074 | 0.32923 | 0.26867
1:33 | 0.20613 | 0.27082 | 0.32914 | 0.2687
1:33 | 0.20603 | 0.27079 | 0.3291 | 0.26864
1:33 | 0.20538 | 0.27088 | 0.32904 | 0.26844
1:34 | 0.20348 | 0.271 | 0.32897 | 0.26782
1:34 | 0.20562 | 0.27103 | 0.32894 | 0.26853
1:34 | 0.20646 | 0.27087 | 0.32918 | 0.26884
1:34 | 0.20643 | 0.27056 | 0.3293 | 0.26876
1:34 | 0.20648 | 0.27069 | 0.32939 | 0.26885
1:34 | 0.20658 | 0.27053 | 0.3295 | 0.26887
1:34 | 0.20655 | 0.27059 | 0.32937 | 0.26884
1:34 | 0.20651 | 0.27058 | 0.3294 | 0.26883
1:34 | 0.20647 | 0.27016 | 0.32922 | 0.26861
1:34 | 0.20634 | 0.27003 | 0.32919 | 0.26852
1:34 | 0.20637 | 0.27015 | 0.32911 | 0.26854
1:34 | 0.2063 | 0.27021 | 0.32903 | 0.26851
1:35 | 0.20636 | 0.27051 | 0.329 | 0.26863
1:35 | 0.20625 | 0.27059 | 0.32879 | 0.26855
1:35 | 0.2061 | 0.27051 | 0.32865 | 0.26842
1:35 | 0.206 | 0.27045 | 0.32865 | 0.26837
1:35 | 0.20595 | 0.27054 | 0.3287 | 0.2684
1:35 | 0.20588 | 0.27058 | 0.32863 | 0.26837
1:35 | 0.20588 | 0.27064 | 0.32866 | 0.26839
1:35 | 0.20597 | 0.27056 | 0.32883 | 0.26845
1:35 | 0.20591 | 0.27052 | 0.32861 | 0.26835
1:35 | 0.20593 | 0.27053 | 0.32855 | 0.26834
1:35 | 0.20593 | 0.27087 | 0.32829 | 0.26836
1:35 | 0.2059 | 0.27094 | 0.32796 | 0.26827
1:36 | 0.20582 | 0.2707 | 0.32757 | 0.26803
1:36 | 0.20597 | 0.27073 | 0.32707 | 0.26792
1:36 | 0.20586 | 0.27095 | 0.32668 | 0.26783
1:36 | 0.20587 | 0.27085 | 0.32665 | 0.26779
1:36 | 0.20578 | 0.27063 | 0.32683 | 0.26774
1:36 | 0.2059 | 0.27086 | 0.32668 | 0.26781
1:36 | 0.20576 | 0.27072 | 0.32683 | 0.26777
1:36 | 0.20576 | 0.27079 | 0.32665 | 0.26773
1:36 | 0.20575 | 0.27091 | 0.32664 | 0.26777
1:36 | 0.20552 | 0.27088 | 0.32681 | 0.26774
1:36 | 0.20549 | 0.27087 | 0.32748 | 0.26795
1:36 | 0.20552 | 0.27112 | 0.3277 | 0.26811
1:37 | 0.20548 | 0.27068 | 0.32809 | 0.26808
1:37 | 0.20563 | 0.27058 | 0.32821 | 0.26814
1:37 | 0.20562 | 0.27069 | 0.32812 | 0.26814
1:37 | 0.2057 | 0.27094 | 0.32828 | 0.26831
1:37 | 0.20564 | 0.27113 | 0.32825 | 0.26834
1:37 | 0.20565 | 0.2711 | 0.32864 | 0.26846
1:37 | 0.20567 | 0.27081 | 0.32871 | 0.2684
1:37 | 0.20581 | 0.2707 | 0.32884 | 0.26845
1:37 | 0.20573 | 0.27047 | 0.32884 | 0.26835
1:37 | 0.20573 | 0.27042 | 0.32898 | 0.26838
1:37 | 0.20578 | 0.27026 | 0.32929 | 0.26844
1:37 | 0.20589 | 0.27013 | 0.32929 | 0.26844
1:38 | 0.20577 | 0.27051 | 0.32951 | 0.2686
1:38 | 0.20582 | 0.27045 | 0.32992 | 0.26873
1:38 | 0.20579 | 0.27087 | 0.33017 | 0.26894
1:38 | 0.20555 | 0.27102 | 0.33032 | 0.26896
1:38 | 0.20565 | 0.27099 | 0.33035 | 0.269
1:38 | 0.20567 | 0.2707 | 0.33074 | 0.26904
1:38 | 0.20577 | 0.27088 | 0.33094 | 0.26919
1:38 | 0.20574 | 0.27088 | 0.33074 | 0.26912
1:38 | 0.20568 | 0.27096 | 0.33068 | 0.26911
1:38 | 0.20565 | 0.27089 | 0.3306 | 0.26904
1:38 | 0.20558 | 0.27099 | 0.33062 | 0.26906
1:38 | 0.20561 | 0.27091 | 0.33071 | 0.26908
1:39 | 0.20556 | 0.27083 | 0.33063 | 0.26901
1:39 | 0.20551 | 0.27076 | 0.33073 | 0.269
1:39 | 0.20542 | 0.2711 | 0.33083 | 0.26912
1:39 | 0.20543 | 0.27095 | 0.33071 | 0.26903
1:39 | 0.20541 | 0.27128 | 0.3308 | 0.26916
1:39 | 0.20544 | 0.27144 | 0.33065 | 0.26918
1:39 | 0.20541 | 0.2715 | 0.3308 | 0.26924
1:39 | 0.20546 | 0.27134 | 0.33089 | 0.26923
1:39 | 0.20544 | 0.27148 | 0.33072 | 0.26922
1:39 | 0.2053 | 0.27146 | 0.33083 | 0.26919
1:39 | 0.20537 | 0.27157 | 0.33065 | 0.2692
1:39 | 0.20537 | 0.27164 | 0.33056 | 0.26919
1:40 | 0.20536 | 0.27198 | 0.33086 | 0.2694
1:40 | 0.20539 | 0.27196 | 0.33062 | 0.26932
1:40 | 0.2053 | 0.27232 | 0.33041 | 0.26934
1:40 | 0.20526 | 0.27237 | 0.33049 | 0.26937
1:40 | 0.20525 | 0.27248 | 0.33038 | 0.26937
1:40 | 0.20526 | 0.27229 | 0.33035 | 0.2693
1:40 | 0.20532 | 0.27211 | 0.33042 | 0.26928
1:40 | 0.20524 | 0.2723 | 0.33046 | 0.26933
1:40 | 0.20531 | 0.2723 | 0.33029 | 0.2693
1:40 | 0.20538 | 0.27214 | 0.33019 | 0.26924
1:40 | 0.20543 | 0.2722 | 0.33012 | 0.26925
1:40 | 0.20536 | 0.27212 | 0.32993 | 0.26914
1:41 | 0.20535 | 0.27227 | 0.33008 | 0.26923
1:41 | 0.20538 | 0.27221 | 0.3298 | 0.26913
1:41 | 0.2053 | 0.27222 | 0.32982 | 0.26911
1:41 | 0.20548 | 0.27227 | 0.32978 | 0.26917
1:41 | 0.2055 | 0.27204 | 0.32942 | 0.26899
1:41 | 0.20542 | 0.27226 | 0.32914 | 0.26894
1:41 | 0.2056 | 0.27193 | 0.3289 | 0.26881
1:41 | 0.20552 | 0.27193 | 0.32935 | 0.26893
1:41 | 0.20573 | 0.27205 | 0.3295 | 0.26909
1:41 | 0.20581 | 0.27186 | 0.33004 | 0.26924
1:41 | 0.20578 | 0.27172 | 0.3303 | 0.26927
1:41 | 0.20575 | 0.27185 | 0.33006 | 0.26922
1:42 | 0.20577 | 0.27182 | 0.32982 | 0.26914
1:42 | 0.2058 | 0.27187 | 0.32974 | 0.26914
1:42 | 0.20574 | 0.27193 | 0.32993 | 0.2692
1:42 | 0.20573 | 0.27214 | 0.32985 | 0.26924
1:42 | 0.20572 | 0.27189 | 0.33011 | 0.26924
1:42 | 0.20586 | 0.27207 | 0.3303 | 0.26941
1:42 | 0.20585 | 0.27216 | 0.33045 | 0.26948
1:42 | 0.206 | 0.27221 | 0.33052 | 0.26958
1:42 | 0.20589 | 0.27233 | 0.33045 | 0.26956
1:42 | 0.20589 | 0.27254 | 0.3307 | 0.26971
1:42 | 0.20587 | 0.27239 | 0.33082 | 0.26969
1:42 | 0.20579 | 0.27247 | 0.33091 | 0.26972
1:43 | 0.20578 | 0.27248 | 0.33103 | 0.26976
1:43 | 0.2058 | 0.27232 | 0.33095 | 0.26969
1:43 | 0.2059 | 0.27258 | 0.33103 | 0.26984
1:43 | 0.20577 | 0.2727 | 0.33094 | 0.26981
1:43 | 0.20565 | 0.27275 | 0.33115 | 0.26985
1:43 | 0.20565 | 0.27266 | 0.33103 | 0.26978
1:43 | 0.20561 | 0.2726 | 0.33099 | 0.26973
1:43 | 0.20564 | 0.27246 | 0.33097 | 0.26969
1:43 | 0.20554 | 0.27265 | 0.33094 | 0.26971
1:43 | 0.20552 | 0.27253 | 0.33096 | 0.26967
1:43 | 0.20559 | 0.27245 | 0.33079 | 0.26961
1:43 | 0.20561 | 0.27259 | 0.33066 | 0.26962
1:44 | 0.20572 | 0.27244 | 0.3305 | 0.26955
1:44 | 0.20564 | 0.27249 | 0.33043 | 0.26952
1:44 | 0.20571 | 0.27239 | 0.33006 | 0.26939
1:44 | 0.20577 | 0.27238 | 0.3302 | 0.26945
1:44 | 0.2057 | 0.27254 | 0.32991 | 0.26938
1:44 | 0.20576 | 0.27247 | 0.33001 | 0.26941
1:44 | 0.20554 | 0.27263 | 0.32974 | 0.2693
1:44 | 0.2056 | 0.27251 | 0.32944 | 0.26918
1:44 | 0.20546 | 0.27264 | 0.32915 | 0.26909
1:44 | 0.20551 | 0.27283 | 0.32892 | 0.26909
1:44 | 0.20557 | 0.2728 | 0.32899 | 0.26912
1:44 | 0.20545 | 0.27287 | 0.32891 | 0.26908
1:45 | 0.2055 | 0.27276 | 0.32886 | 0.26904
1:45 | 0.2056 | 0.27291 | 0.32849 | 0.269
1:45 | 0.20559 | 0.27283 | 0.32853 | 0.26898
1:45 | 0.20569 | 0.27279 | 0.32864 | 0.26904
1:45 | 0.20573 | 0.27288 | 0.32882 | 0.26915
1:45 | 0.20566 | 0.27279 | 0.32952 | 0.26932
1:45 | 0.20569 | 0.27295 | 0.3298 | 0.26948
1:45 | 0.20544 | 0.27292 | 0.33023 | 0.26953
1:45 | 0.20495 | 0.27295 | 0.33055 | 0.26948
1:45 | 0.20485 | 0.27291 | 0.33107 | 0.26961
1:45 | 0.20522 | 0.27295 | 0.33195 | 0.27004
1:45 | 0.20511 | 0.27294 | 0.33244 | 0.27016
1:46 | 0.20508 | 0.27296 | 0.33233 | 0.27012
1:46 | 0.20511 | 0.27287 | 0.33236 | 0.27011
1:46 | 0.20518 | 0.27299 | 0.33225 | 0.27014
1:46 | 0.20527 | 0.27276 | 0.33192 | 0.26998
1:46 | 0.20526 | 0.27277 | 0.33159 | 0.26987
1:46 | 0.20524 | 0.27297 | 0.33142 | 0.26988
1:46 | 0.20526 | 0.27295 | 0.33134 | 0.26985
1:46 | 0.20517 | 0.27301 | 0.33098 | 0.26972
1:46 | 0.20528 | 0.27328 | 0.33067 | 0.26974
1:46 | 0.20531 | 0.27324 | 0.33015 | 0.26957
1:46 | 0.20544 | 0.27294 | 0.32969 | 0.26936
1:46 | 0.20552 | 0.27305 | 0.3293 | 0.26929
1:47 | 0.20554 | 0.2732 | 0.32949 | 0.26941
1:47 | 0.20553 | 0.27319 | 0.3297 | 0.26947
1:47 | 0.20546 | 0.27306 | 0.33042 | 0.26965
1:47 | 0.20552 | 0.27298 | 0.33145 | 0.26998
1:47 | 0.20565 | 0.27316 | 0.33215 | 0.27032
1:47 | 0.20552 | 0.27322 | 0.33263 | 0.27046
1:47 | 0.20551 | 0.27321 | 0.3325 | 0.27041
1:47 | 0.2055 | 0.273 | 0.3325 | 0.27033
1:47 | 0.20551 | 0.27307 | 0.33259 | 0.27039
1:47 | 0.20554 | 0.27309 | 0.33265 | 0.27043
1:47 | 0.2056 | 0.27301 | 0.33305 | 0.27055
1:47 | 0.2057 | 0.27317 | 0.33317 | 0.27068
1:48 | 0.2057 | 0.27313 | 0.33321 | 0.27068
1:48 | 0.20576 | 0.273 | 0.33338 | 0.27071
1:48 | 0.20574 | 0.27296 | 0.33337 | 0.27069
1:48 | 0.20568 | 0.27331 | 0.33308 | 0.27069
1:48 | 0.20584 | 0.27317 | 0.33316 | 0.27072
1:48 | 0.20585 | 0.27326 | 0.33318 | 0.27076
1:48 | 0.20594 | 0.27329 | 0.33323 | 0.27082
1:48 | 0.20607 | 0.27341 | 0.33321 | 0.2709
1:48 | 0.20594 | 0.27349 | 0.33328 | 0.2709
1:48 | 0.20586 | 0.27351 | 0.33299 | 0.27079
1:48 | 0.20583 | 0.27332 | 0.333 | 0.27072
1:48 | 0.20547 | 0.27333 | 0.33286 | 0.27055
1:49 | 0.20485 | 0.2731 | 0.33281 | 0.27025
1:49 | 0.20591 | 0.2731 | 0.33287 | 0.27063
1:49 | 0.20617 | 0.27307 | 0.33275 | 0.27066
1:49 | 0.20599 | 0.27307 | 0.33269 | 0.27058
1:49 | 0.20587 | 0.27317 | 0.33262 | 0.27055
1:49 | 0.20598 | 0.27326 | 0.33265 | 0.27063
1:49 | 0.20594 | 0.27335 | 0.33253 | 0.27061
1:49 | 0.20598 | 0.27333 | 0.33262 | 0.27064
1:49 | 0.20606 | 0.27339 | 0.33283 | 0.27076
1:49 | 0.20621 | 0.27338 | 0.33277 | 0.27079
1:49 | 0.2062 | 0.27338 | 0.33264 | 0.27074
1:49 | 0.20616 | 0.2732 | 0.33263 | 0.27066
1:50 | 0.20599 | 0.27314 | 0.33223 | 0.27045
1:50 | 0.20614 | 0.27326 | 0.33109 | 0.27016
1:50 | 0.20609 | 0.27341 | 0.32977 | 0.26976
1:50 | 0.20603 | 0.27332 | 0.32936 | 0.26957
1:50 | 0.20607 | 0.27302 | 0.32946 | 0.26952
1:50 | 0.206 | 0.2732 | 0.32989 | 0.2697
1:50 | 0.20614 | 0.27321 | 0.32975 | 0.2697
1:50 | 0.20603 | 0.27328 | 0.32861 | 0.26931
1:50 | 0.20609 | 0.27319 | 0.32597 | 0.26842
1:50 | 0.20622 | 0.27325 | 0.3247 | 0.26806
1:50 | 0.20631 | 0.27308 | 0.32463 | 0.26801
1:50 | 0.20651 | 0.27312 | 0.32441 | 0.26802
1:51 | 0.20641 | 0.27323 | 0.32411 | 0.26792
1:51 | 0.20643 | 0.27312 | 0.32061 | 0.26672
1:51 | 0.20521 | 0.27319 | 0.32063 | 0.26634
1:51 | 0.20499 | 0.27296 | 0.32227 | 0.26674
1:51 | 0.20452 | 0.27319 | 0.3225 | 0.26673
1:51 | 0.20174 | 0.27323 | 0.32247 | 0.26581
1:51 | 0.19909 | 0.27317 | 0.32315 | 0.26514
1:51 | 0.20703 | 0.27329 | 0.32393 | 0.26808
1:51 | 0.20588 | 0.27324 | 0.32472 | 0.26795
1:51 | 0.2054 | 0.27327 | 0.32485 | 0.26784
1:51 | 0.206 | 0.27319 | 0.32468 | 0.26796
1:51 | 0.20578 | 0.27303 | 0.32461 | 0.2678
1:52 | 0.20329 | 0.27341 | 0.32468 | 0.26713
1:52 | 0.20647 | 0.27333 | 0.32474 | 0.26818
1:52 | 0.20374 | 0.27326 | 0.32463 | 0.26721
1:52 | 0.20533 | 0.27463 | 0.32524 | 0.2684
1:52 | 0.20663 | 0.27591 | 0.32504 | 0.26919
1:52 | 0.20628 | 0.2736 | 0.32524 | 0.26837
1:52 | 0.20663 | 0.27334 | 0.32603 | 0.26866
1:52 | 0.20665 | 0.27345 | 0.32601 | 0.2687
1:52 | 0.20642 | 0.27331 | 0.325 | 0.26825
1:52 | 0.20639 | 0.2733 | 0.3237 | 0.2678
1:52 | 0.20626 | 0.27351 | 0.32183 | 0.2672
1:52 | 0.2063 | 0.27331 | 0.32156 | 0.26706
1:53 | 0.20662 | 0.27327 | 0.32111 | 0.267
1:53 | 0.20649 | 0.27352 | 0.32119 | 0.26707
1:53 | 0.20664 | 0.27336 | 0.32143 | 0.26714
1:53 | 0.20668 | 0.27346 | 0.32123 | 0.26712
1:53 | 0.20671 | 0.27338 | 0.32131 | 0.26713
1:53 | 0.20679 | 0.27329 | 0.32142 | 0.26717
1:53 | 0.20682 | 0.27342 | 0.32181 | 0.26735
1:53 | 0.20687 | 0.2732 | 0.32251 | 0.26753
1:53 | 0.2066 | 0.27338 | 0.32263 | 0.26753
1:53 | 0.20487 | 0.27335 | 0.32258 | 0.26693
1:53 | 0.20122 | 0.27335 | 0.32312 | 0.26589
1:54 | 0.20468 | 0.27316 | 0.32303 | 0.26696
1:54 | 0.19427 | 0.27305 | 0.32298 | 0.26343
1:54 | 0.18036 | 0.27287 | 0.32333 | 0.25886
1:54 | 0.18137 | 0.27309 | 0.32353 | 0.25933
1:54 | 0.18255 | 0.27331 | 0.3233 | 0.25972
1:54 | 0.18274 | 0.27338 | 0.32252 | 0.25955
1:54 | 0.1829 | 0.27344 | 0.32161 | 0.25931
1:54 | 0.18316 | 0.27349 | 0.32127 | 0.25931
1:54 | 0.18312 | 0.27358 | 0.32121 | 0.2593
1:54 | 0.18295 | 0.27345 | 0.32276 | 0.25972
1:54 | 0.18299 | 0.27337 | 0.32388 | 0.26008
1:54 | 0.1829 | 0.27329 | 0.32466 | 0.26028
1:55 | 0.1833 | 0.27337 | 0.3258 | 0.26082
1:55 | 0.183 | 0.27317 | 0.32612 | 0.26076
1:55 | 0.18684 | 0.27326 | 0.32762 | 0.26257
1:55 | 0.18298 | 0.27352 | 0.32874 | 0.26174
1:55 | 0.1832 | 0.27348 | 0.32825 | 0.26164
1:55 | 0.18284 | 0.27344 | 0.3279 | 0.2614
1:55 | 0.18335 | 0.27315 | 0.32569 | 0.26073
1:55 | 0.18397 | 0.27339 | 0.32647 | 0.26128
1:55 | 0.18432 | 0.2734 | 0.32681 | 0.26151
1:55 | 0.18449 | 0.27341 | 0.32691 | 0.2616
1:55 | 0.18502 | 0.27339 | 0.32683 | 0.26175
1:55 | 0.1878 | 0.27329 | 0.32683 | 0.26264
1:56 | 0.18856 | 0.27329 | 0.32668 | 0.26284
1:56 | 0.19595 | 0.27331 | 0.32659 | 0.26528
1:56 | 0.20619 | 0.27326 | 0.32688 | 0.26878
1:56 | 0.20584 | 0.27318 | 0.32754 | 0.26885
1:56 | 0.20313 | 0.27328 | 0.32767 | 0.26802
1:56 | 0.20038 | 0.27318 | 0.32758 | 0.26705
1:56 | 0.19844 | 0.27343 | 0.32798 | 0.26662
1:56 | 0.20223 | 0.27357 | 0.32836 | 0.26805
1:56 | 0.20865 | 0.27346 | 0.32836 | 0.27016
1:56 | 0.20755 | 0.27365 | 0.32841 | 0.26987
1:56 | 0.20709 | 0.2735 | 0.32848 | 0.26969
1:56 | 0.2069 | 0.27366 | 0.32863 | 0.26973
1:57 | 0.20704 | 0.27352 | 0.32861 | 0.26972
1:57 | 0.20701 | 0.27369 | 0.32867 | 0.26979
1:57 | 0.20683 | 0.27395 | 0.32855 | 0.26977
1:57 | 0.20673 | 0.27382 | 0.32762 | 0.26939
1:57 | 0.20672 | 0.27391 | 0.32773 | 0.26945
1:57 | 0.2065 | 0.27376 | 0.32769 | 0.26932
1:57 | 0.20611 | 0.27375 | 0.32813 | 0.26933
1:57 | 0.20615 | 0.27372 | 0.32841 | 0.26943
1:57 | 0.20526 | 0.2737 | 0.3285 | 0.26915
1:57 | 0.20184 | 0.27383 | 0.32859 | 0.26809
1:57 | 0.19612 | 0.27387 | 0.3289 | 0.26629
1:57 | 0.20591 | 0.27373 | 0.32902 | 0.26955
1:58 | 0.20698 | 0.27345 | 0.32923 | 0.26989
1:58 | 0.20614 | 0.27363 | 0.32921 | 0.26966
1:58 | 0.20622 | 0.27356 | 0.32914 | 0.26964
1:58 | 0.20479 | 0.27359 | 0.32938 | 0.26925
1:58 | 0.20004 | 0.27353 | 0.3293 | 0.26762
1:58 | 0.19513 | 0.27331 | 0.32904 | 0.26583
1:58 | 0.19132 | 0.27324 | 0.32785 | 0.26414
1:58 | 0.19602 | 0.27356 | 0.3263 | 0.26529
1:58 | 0.19693 | 0.27363 | 0.32531 | 0.26529
1:58 | 0.19954 | 0.27347 | 0.32486 | 0.26595
1:58 | 0.19845 | 0.27322 | 0.32548 | 0.26572
1:58 | 0.19001 | 0.27343 | 0.32566 | 0.26303
1:59 | 0.18899 | 0.2735 | 0.32547 | 0.26265
1:59 | 0.18864 | 0.27353 | 0.32433 | 0.26217
1:59 | 0.18642 | 0.27336 | 0.3244 | 0.2614
1:59 | 0.18889 | 0.27324 | 0.32424 | 0.26212
1:59 | 0.19438 | 0.27331 | 0.32452 | 0.26407
1:59 | 0.20117 | 0.27316 | 0.32455 | 0.26629
1:59 | 0.20005 | 0.27324 | 0.32419 | 0.26583
1:59 | 0.20319 | 0.2732 | 0.32516 | 0.26718
1:59 | 0.20795 | 0.27336 | 0.32511 | 0.26881
1:59 | 0.20757 | 0.27344 | 0.32726 | 0.26942
1:59 | 0.20803 | 0.2736 | 0.32731 | 0.26965
1:59 | 0.20745 | 0.27371 | 0.32781 | 0.26966
2:00 | 0.20658 | 0.27369 | 0.33061 | 0.27029
2:00 | 0.20612 | 0.27363 | 0.33291 | 0.27089
2:00 | 0.20465 | 0.27339 | 0.33315 | 0.2704
2:00 | 0.20513 | 0.27338 | 0.33305 | 0.27052
2:00 | 0.2066 | 0.27343 | 0.33308 | 0.27103
2:00 | 0.20572 | 0.27337 | 0.33316 | 0.27075