Protocol for beh_operant.ino

Program purpose: Operant conditioning for liquid reward or TTL for using external hardware (e.g. lasers for optogenetic stimulation) with wheel rotation as an operant response. Optional use of a wheel-brake, retractable spout, and tones for both tasks.

Preparation Instructions:

- Around once per week, calibrate the solenoid open duration by following protocol_helper_calibratesolenoid. Save the solenoid open duration(s) somewhere convenient in a vector format with 1 value per solenoid (e.g. 3 for 3 spouts: each vector would look like  {value, value, value}).

- Fill the solution line(s) by following protocol_helper_opensol.

- Optional: connect BNC or TTL outputs from the arduino console to the external hardware.

Arduino Setup Instructions:

- Open the program

- Set the parameters and input / output pins

- Carefully read the comments next to each variable

- Upload script to arduino by clicking the “Upload” arrow button on the top left corner.

- The arduino will now have the script running on it and it will wait until a start command is sent over serial.

Arduino Test Instructions:

Test the program at the start of each day to ensure all hardware is working and events are being recorded correctly.

- To test the program, open the serial monitor by clicking the “Serial Monitor” button on the top right corner.

Note: If you do not see this text printed in the serial monitor, see troubleshooting arduino software document

- Start the arduino program by clicking the button to send an empty string over serial.

- Watching the Serial Monitor while interacting with the hardware, check to ensure that…

  - Rotating the wheel in the intended active direction produces code 81.

  - Rotating the wheel  in the intended inactive direction produces code 71.

  - Touching the spout produces code 30 + current spout.

  - Rotation of the wheel in the active direction that reaches the fixed ratio triggers the intended sequence of events (break, tone, spout extension, liquid delivery or external TTL signal).

  - If performing progressive ratio, check to ensure that the cost of the reinforcer increases as intended.

  - Rotation in the inactive direction that reaches the fixed ratio triggers the intended sequence.

- For multi-spout systems, if the intended spout is not detecting touch,check other spouts to see if they are detecting touch. If you find that a spout other than the one rotated towards the mouse is detecting touch, then check the servo_retract_radial_degs vector to ensure the values are correct for your multi-spout head, and check to ensure the servo is working properly.

  - If liquid is emerging from the wrong spout, then check the pinSol vector to ensure that the correct pins are included for your system.

- Reset the arduino by closing and reopening the Serial Monitor, and then close Serial Monitor prior to the session.

Arduino Run Instructions:

- Open the write_serial.py program (shown here using Sublime Text 3)

- Set the COMPORT and BAUDRATE to match the arduino script

Arduino comport is shown in the tools drop down menu.

Arduino baudrate is set with “Serial.begin” call

- Set the subject name for the session in the python program

Note: See general arduino software document for more details about the functionality of the write_serial.py program.

- After the program and hardware have been tested, head-fix the animal.

- Set the spout position as shown in this diagram:

- Run the python program (Cntrl + B for Sublime Text 3).

- Values should begin to print with the start of the session, and will continue to print during recorded events.

Note: If you see an error, refer to the general arduino software document for troubleshooting tips.

- At the end of the session, the spout will retract (if using a retractable spout), the break will engage,  and the python script will print “stop read_serial”).

- Remove subject from head-fixation.

- Restart arduino prior to subsequent subjects by opening / closing the arduino Serial Monitor. Otherwise, the spout will extend for the next subject and could result in poking the subject.