## Slide 1: OSI² One build workshop Gradient Coils
Sebastian Littin
Medical Physics, University Medical Center Freiburg, Germany
Berlin, September 22th, 2023

## Slide 2: Overview
Gradient Coils in MRI
Gradient Coil Design
Gradient Coil Manufacturing

## Slide 3: Overview
Gradient Coils in MRI
Gradient Coil Design
Gradient Coil Manufacturing

## Slide 4: Gradient Coils in MRI: NMR vs MRI
What is necessary for the NMR Signal?
Sample => Signal source (1H, 13C, 17O, …)
Magnet => Polarization (T; static (0Hz))
RF coil => Excitation/ Acquisition (µT; MHz)
Shims => Field Homogenization
NMR Spectrometer
Gradient Coils => Signal Localization
Gradients enable imaging
NMR => MRI
up to ~300mT/m; up to ~5 kHz

## Slide 5: Spatially Varying Magnetic Fields - a.k.a. “Gradients”
Spatial variation of B0-field: Linearly varying magnetic field (Bz)
Local variation of the Larmor frequency
Linear relation between object and frequency space
Homogeneous slice profiles
Straight-forward Fourier Transform (FFT)
Usually three orthogonal magnetic (Bz) field gradients along spatial directions x,y and z
Extends beyond imaging region
Object
Constant Gradient

## Slide 6: How linear are our encoding fields?
All gradients are non-constant to some degree
FOV=50cm, 20cm above iso-center no gradient distortion correction
clinical scanner, XZ (coronal) slice FOV=50cm, iso-center
FOV=50cm, iso-center no gradient distortion correction
375 mm

## Slide 7: Gradient Coils in MRI: Image Encoding
Accumulated Phase:
Spatial Resolution:
Wanted:
High gradient amplitude G (mT/m, scales with current)
Fast switching/ slew rate (T/s, scales with voltage)
Efficiency vs. slew rate
Area =
Gradient Moment
Amplitude (G)
Time (t)
T
Imaging: Gradient Moment

## Slide 8: Overview
Gradient Coils in MRI
Gradient Coil Design
Gradient Coil Manufacturing

## Slide 9: Gradient Coil Design: Boundary Element Method
Derive wire paths from stream lines: Trajectories of equally spaced steady flow of current to approximate surface current density
Stream function , relates to surface current density according to:
Resulting field with points can be expressed using sensitivity matrix :
Cost function with additional constraints for defined as:

## Slide 10: Openly Available Software Repositories
Fully flexible design and post processing “CoilGen: Open Source MR Coil Layout Generator“ [1]
https://github.com/Philipp-MR/CoilGen
Educational “Thin-wire approximation” [2]
https://github.com/Sebastian-MR/ThinWire_MRIGradientCoilDesign
Python based “GradientDesignTool” (Patrick Fuchs)
https://github.com/LUMC-LowFieldMRI/GradientDesignTool
Matlab based “CoilDesign” (Gale Bringout)
https://github.com/gBringout/CoilDesign/
[3] Amrein P et al CoilGen: Open-source MR coil layout generator. Magn Reson Med. 2022; 88( 3): 1465- 1479. doi:10.1002/mrm.29294
[4] Littin S et al (2021) Methods: Of Stream Functions and Thin Wires: An Intuitive Approach to Gradient Coil Design. Front. Phys. 9:699468. doi:10.3389/fphy.2021.699468

## Slide 11: How to design a gradient coil?
Define basic properties:
  Current carrying surface(s)
    Size/ Shape/ Resolution
  Target (main (&shield))
    Size/ Shape/ Resolution
  Main field (B0) direction
    Bz || B0
    Transverse components Bx/By ?
  Additional Constraints
    Torque/ force balance
    Minimum wire distance
    …
=> Fully flexible design and postprocessing “CoilGen: Open Source MR Coil Layout Generator“ [4]
https://github.com/Philipp-MR/CoilGen
[4] Amrein P et al CoilGen: Open-source MR coil layout generator. Magn Reson Med. 2022; 88( 3): 1465- 1479. doi:10.1002/mrm.29294

## Slide 12: Boundary Element Method: CoilGen
Generate triangular surface mesh
Design on arbitrary surfaces
Optimize the stream function on boundary elements of a mesh [7,8]
Advanced post-processing options
Output: Non-overlapping, interconnected coil layout.
[7] Peeren G N. Stream function approach for determining optimal surface currents Journal of Computational Physics. 2003;191:305–321.
[8] Lemdiasov et al. A stream function method for gradient coil design Concepts in Magnetic Resonance Part B: Magnetic Resonance Engineering. 2005;26B:67–80.

## Slide 13: CoilGen: Beyond Stream Functions
Project on 2D surface, discretize
Analyze topology, define path
Interconnect in & between groups
Project back, manage overlaps

## Slide 14: CoilGen: Coil Designs for Halbach Systems
x-Gradient
y-Gradient
z-Gradient

## Slide 15: CoilGen: Coil Designs for Halbach Systems
x-Gradient
y-Gradient
z-Gradient

## Slide 16: CoilGen: 3D printed carrier (FH Dortmund)
Images courtesy of Nils Allek (FH Dortmund)

## Slide 17: Overview
Gradient Coils in MRI
Gradient Coil Design
Gradient Coil Manufacturing

## Slide 18: Gradient Coil Manufacturing
How to realize a complex wire path?
Different approaches feasible:
Printed circuit boards (PCBs)
Wind on carrier structure (3D printed or milled)
Wind from copper wire
Directly cooled hollow wire
Remove material from solid sheet material (milling machine/ laser-cutter/ water-jet)

## Slide 19: Printed Circuit Boards (PCBs)
[10] Zimmerman Cooley et al. Design and implementation of a low-cost, tabletop MRI scanner for education and research prototyping JMR Vol. 317, p. 106764
[11] OCRA Gradients, Images courtesy of M. Prier and P. Amrein
[12] Amrein et al. Optimal bi-planar gradient coil configurations for diamnond nitrogen-vacancy based diffusion-weighted NMR experiments, MagnResonMaterPhy (under review)
Established manufacturing technique
Simple, cheap and precise from experienced companies
Limited cross section of conductor
[10]
[11]
[12]

## Slide 20: 3D Printed Carrier Structures
3D printed coil carrier [16,17]
  Powder bed and inkjet head 3D printer
Individually insulated Litz wire
Vacuum infusion with epoxy resin
[17] Littin, S., et al. (2018), Development and implementation of an 84-channel matrix gradient coil. Magn. Reson. Med., 79: 1181-1191. https://doi.org/10.1002/mrm.26700
[16] Littin, S., et al. (2021), Approaching order of magnitude increase of gradient strength: Non-linear breast gradient coil for diffusion encoding, Proc. ISMRM 3096, 2021

## Slide 21: Manufacturing Take Home Message(s)
Building hardware is:
  Fun and rewarding
  Tedious and exhausting: Plan for a “second first try”…
Use established connectors, methods, etc.
Question established practices
Check and double check for ferromagnetic parts
Measure before casting

## Slide 22: Summary
Gradients enable for imaging: NMR => MRI
All gradients are non-constant to some degree
Stream function to approximate surface current density
BEM-methods mostly used
CoilGen: Pipeline for coil design established
Translation of CoilGen to Python in progress
Post processing necessary
Variety of manufacturing methods for implementing gradient coils within a research environment

## Slide 23
Acknowledgements
Philipp Amrein
Feng Jia
Ben Wilhelm
Maxim Zaitsev
All current and previous colleagues and collaboration partners
Further reading on gradient coils
[21] Hidalgo-Tobon, S.S. (2010), Theory of gradient coil design methods for magnetic resonance imaging. Concepts Magn. Reson., 36A: 223-242. https://doi.org/10.1002/cmr.a.20163
[23] Gudino, N. and Littin, S. (2023), Advancements in Gradient System Performance for Clinical and Research MRI. J Magn Reson Imaging, 57: 57-70. https://doi.org/10.1002/jmri.28421
This work is part of the A4IM project.
This project (22HLT02 A4IM) has received funding from the European Partnership on Metrology, co-financed from the European Union‘s Horizon Europe Research and Innovation Programme and by the Participating States.