## Slide 1: Shimming
Tom O’Reilly
Leiden University Medical Center

## Slide 2: Diameter optimisation
25cm DSV: B0 = 46.35 mT
ΔB0 = 3000 ppm

## Slide 3: Construction: Magnet imperfections
Hugon et al, JMR, 2010
3000 ppm
±2 %
9000 ppm
±2 degrees

## Slide 4: Construction: Torque induced errors
33000 ppm
±1.5 degrees

## Slide 5: Torque induced errors: Compensation
Augment magnet rotations
Prior to cutting of rings!

## Slide 6: Construction: Dud magnets

## Slide 7: Construction: The finished magnet!

## Slide 8: Shimming: Mapping
3D positioning system
    COSI Measure1 (open source)
Hall effect sensors
    LakeShore Cryotronics F71
    AlphaLab Inc, High precision probe
Fit spherical harmonics
    Helps reduce noise
Temperature stable environment
    !!!!!!!!!!!
Consider background field
1. https://github.com/opensourceimaging/cosi-measure

## Slide 9: Spherical harmonic decomposition

## Slide 10: Field maps
Homogeneity: Simulated Measured
250 mm DSV: 3000 ppm 24795 ppm

## Slide 11: Shimming:
Use (small) magnets to correct inhomogeneity
    Iron shims less suitable
Perform optimisation on measured field
    Cannot be pre-computed
    Be flexible in shim system design
Typically an iterative process
    Consider when designing shim system

## Slide 12: Shim calculations
Python based optimisation
    (will be) open source
Define properties/location of magnets
Read in measuerd B0 map
Calculates optimal shim configuration

## Slide 13: Algrotihm outputs

## Slide 14: Construction
Txt file: x,y,z, angle

## Slide 15: Construction

## Slide 16: Mapping: 24795 ppm to 6806 ppm

## Slide 17: Repeat process

## Slide 18: Iteration 2 and 3 – Full shimming
3889 ppm
3383 ppm

## Slide 19: General notes
Limits to how good the shim can be
    Typically equal to initial simulations
Shim is very flexible
    Strength and size of magnets
Don’t try to get it perfect first time
    3 or 4 iterations typically needed
Try to adjust number of magnets
    Decreasing returns on investment