## Slide 1: RF coil optimization 101
Tom O’Reilly

## Slide 2: Introduction
What makes a good coil?
    Tx/Rx difference?
    High/low field difference?
How do you build a good coil?
    What can you ‘optimise’?
    Practical limits
With the good comes the bad
    Scanning with ‘good coils’

## Slide 3: What makes a good transmit coil?
Homogenous
B1+ field
High SAR efficiency
(B1+ per SAR)
High Efficiency
(B1+ per input power)

## Slide 4: What makes a good HIGH FIELD transmit coil?
Homogenous
B1+ field
High SAR efficiency
(B1+ per SAR)
High Efficiency
(B1+ per input power)
At high field SAR is often limiting
Wavelength can be an issue
Large RF amps

## Slide 5: What makes a good LOW FIELD transmit coil?
Homogenous
B1+ field
High SAR efficiency
(B1+ per SAR)
High Efficiency
(B1+ per input power)
Usually OK
(virtually) no SAR
No wavelength issues

## Slide 6: What makes a good receive coil?
Unique sensitivity profiles
(parallel imaging)
High B1-
(sensitivity)

## Slide 7: What makes a good Tx/Rx coil?
Maximise B field
Coil type
    Depends on B0 direction
Coil size
    Smaller = Better
Minimise losses
Looooong story
    Different for different frequencies

## Slide 8: Q-uantify the Q-uality of the coil
Quantify high B field and low losses in 1 number
Quality factory:
    L = coil inductance, a measure of energy stored in B field
    R = resistance, a loss term that dissipates energy
SNR
    *For coils with the same geometry!
    Maximise Q

## Slide 9: Maximise B-field
Solenoid is champion
    ~3x better than other coils
Minimum size set by anatomy
    Knee coil: 15cm radius, 15cm length
How to get more magnetic field from a solenoid?
    More turns?
    but too, so Q independent of n
    Tested this, it appears to hold perfectly true
Conclusion: Stuck with what we have

## Slide 10: Minimise R
  Address the loss terms instead:
    R = total losses in the coil, many terms contribute
    Two predominant loss term categories:
      Rcoil: losses in the coil itself
      Rsample: losses due to a conductive sample in the coil
  Huge difference between low and high field
    At high field Rsample dominant
      Small coils are better, Rsample lower, SNR higher
    At low field Rcoil dominant*
      Large coils are better, Rcoil lower, SNR higher

## Slide 11: Minimise R – Cryogenic coils

## Slide 12: Optimisation - Reference coil
All experiments are done on the same coil base:
    156 mm diameter, 150 mm long solenoid
    Wound on a 3D printed PLA former

## Slide 13: Addressing Rcoil
Thicker wire
    Reduces wire resistance
Wire diameter (mm) | Q factor
0,8 | 130
1,5 | 280
r
Expected ~4x improvement

## Slide 14: Addressing Rcoil
Thicker wire
    * at DC
    * at RF
Wire diameter (mm) | Q factor
0,8 | 130
1,5 | 280

## Slide 15: Skin depth – 1.5 Tesla body coil

## Slide 16: Skin depth – 1.5 Tesla body coil

## Slide 17: Litz wire
Many thin insulated strands together
Skin effect still present
    More surface area for current
Reduced conductor mass per meter
    Higher resistance at low F
Loses effectivity at higher f
Current distribution purely due to skin effect

## Slide 18: Litz wire coil
Wire type | Q factor
Solid copper | 280
Litz | 424
Close to the limit

## Slide 19: Addressing Rsample
Maxwell says no without E
Electric field causes losses in conductive sample
  E field of coil increases if current distribution is not constant
    Long conductors should be avoided
      Can be split to reduce impact

## Slide 20: Addressing Rsample
    45 turn solenoid: 22 meters of wire
      150 m wavelength: 52 degrees of phase
    Testing number of segmentations
    Number of turns reduced
      Performance unchanged
      Sample noise ‘dominated’ coil!
Segmentations | Q unloaded | Q loaded | Ratio
1 | 305 | 91 | 3,4
3 | 427 | 177 | 2,4

## Slide 21: Signal (to-noise) improvement: Spin echo
Coil Q | Max signal [uV]
88 | 40
234 | 65
362 | 74
Coil Q | Max signal [uV]
88 | 40
234 | 65
362 | 74

## Slide 22: Why do we NOT want high Qs
Coil ringing gets long
    Exponential ramp/decay:
    Heavily distorts pulse shape
Narrow coil bandwidth limits ‘FOV’ of coil
    Problem on Tx and Rx
      Tx outside BW can damage electronics
      Rx sensitivity weighting in freq. encoding

## Slide 23: Coil ringing

## Slide 24: Coil ringing

## Slide 25: Coil ringing
100 us hard RF pulse

## Slide 26: Coil Bandwidth – Frequency dependent signal

## Slide 27: Coil bandwidth - Noise
Low Q coil
High Q coil

## Slide 28: Coil Bandwidth: Correction
÷
=

## Slide 29: Imaging performance

## Slide 30: Vetinary scanning – Post mortem
PD 3D TSE sequence
FoV: 150 × 150 × 180 mm3
Acq. Mat.: 100 × 100 × 60
TR/TE: 600 ms/18 ms
ETL: 5
No averaging
Circular k-space window
Acq. BW: 20 kHz
Acq. time: 7 min 39 sec

## Slide 31: Vetinary scanning – Post mortem
T2w 3D TSE sequence
FoV: 150 × 150 × 180 mm3
Acq. Mat.: 100 × 100 × 60
TR/TE/TEeff: 2000 ms/18 ms
ETL: 20
No averaging
Circular k-space window
Acq. BW: 20 kHz
Acq. time: 7 min 42 sec

## Slide 32: In-vivo imaging - 3D TSE
FoV: 220 × 220 × 225 mm3
Acq. Mat.: 110 × 110 × 75
TR/TE: 600 ms/15 ms
ETL: 7
Elliptical k-space mask
No averaging
Acq. BW: 20 kHz
Acq. time: 9 min 2 sec
CPR reconstruction
No denoising

## Slide 33: In-vivo imaging – 3D IR-TSE
FoV: 220 × 220 × 180 mm3
Acq. Mat.: 110 × 110 × 60
TR/TI/TE: 6000 ms/200 ms/18 ms
ETL: 110
No averaging
Acq. BW: 20 kHz
Acq. time: 6 min 12 sec
CPR reconstruction
No denoising

## Slide 34: Summary
Coil optimization is field dependent
    Coil materials plays huge role at low field
Body loss dominance can be achieved at low field
    Comes from B field, not E field
Simple correction for BW related shading
Coil optimization yields improved SNR
    2x in phantoms, 2> in-vivo >2½

## Slide 35
Thanks!
“The project (22HLT02 A4IM) has received funding from the European Partnership on Metrology, co-financed from the European Union’s Horizon Europe Research and Innovation Programme and by the Participating States.”
Funder name: European Partnership on Metrology
Funder ID: 10.13039/100019599
Grant number: 22HLT02 A4IM

## Slide 36: RF power
Coil Q | 90° pulse (Vpp) | 90° power (W)
88 | 77,7 | 15,1
234 | 45,6 | 5,2
362 | 37,4 | 3,5
Coil Q | 90° pulse (Vpp) | 90° power (W)
88 | 77,7 | 15,1
234 | 45,6 | 5,2
362 | 37,4 | 3,5