Methodologies employed for lipid mediator identification and quantitation

- Establishing the LOQ

- To establish the LOQ of the assay standards were injected in triplicates at the following concentrations: 0.025, 0.05, 0.1, 0.2, 0.4 pg

- The AUC values were calculated

- The Relative Standard Deviation (RSD) for AUC values at each concentration were then calculated

- The lowest concentration which gave an RSD value < 20% was used as the LOQ

- Lipid Mediator Identification

- The retention time for a specific lipid mediator is established in the relevant standard mix (ie samples labelled as Std Mix, AT-Mix or Iso Mix)

- The retention time of the relevant deuterium labelled internal standard is established in the relevant standard mix

- The delta for the retention time of the internal standard and mediator of interest is calculated

- The retention time of the internal standard is established in the sample

- The expected retention time for the mediator of interest is then calculated using the delta calculated in the reference mix. E.g. if the retention time of RvD5 in the Standard mix is of 13.15 min and that of its relevant deuterium label (ie d4-LTB4) in the same mix is of 13.6 min giving a delta of 0.45 min and the retention time of d4-LTB4 in the sample is 13.2, then the expected retention time of RvD5 in the sample is of 13.65 ± 0.1 min.

- If the retention time matches, then the peak is integrated and if the peak area is equivalent as that obtained for the LOQ then the mediator is deemed to be present in the sample

- Standard Curves

- Standard curves were constructed by injecting the following amounts of the lipid mediators: 0, 0.78, 1.56, 3.12, 6.25, 12.5, 25, 50, 100, 200 pg

- The AUC for signals at the expected retention time for each of the molecules is then integrated

- Linear regression for each of the mediator of interest then calculated in Prizm using the value obtained for blank samples (ie 0pg) as the intercept.

- Where standard curves for specific mediators were not constructed, (usually due to limited availability of the standard material) the mediator with the closest physical properties is used as a surrogate. E.g we use the standard curve for MaR1 to calculate the concentrations of its double bond isomer 7S, 14S-diHDPA. (Please note that standard curves for all the mediators for which quantities are reported in the PLoS One paper i.e. TXB2, PGE2, PGD2, PGF2a, RvE3, 5S,15S-diHETE, LXB4 and RvD5 were available).

- Lipid Mediator Quantitation

- The AUC values for the deuterium labeled internal standards and those for the mediator of interest are copied from Analyst into the respective tabs in Prism.

- The calculated concentrations are then copied back to Excel

- Using the calculated concentrations for the 100% control, recoveries are calculated by dividing the calculated concentration for each of the internal standards by that calculated in the 100% control. E.g if the calculated concentration of d4-LTB4 in the sample was of 150 pg and that in the 100% control was of 300pg, then the recovery would be of 50%.

- Using the estimated recovery values, the concentration of the mediator in the sample is then obtained by multiplying the recovery for the relevant deuterium label by the calculated amount of the molecules. E.g if the calculated concentration of RvD5 from the AUC values is of 1.2 pg, and the recovery was of 50%, then the calculated concentration is multiplied by 2 to give a final concentration in the sample of 2.4 pg.