## Info
Supplementary table: MSSettings | Tab | Used in
timsTOF-Pro | LC and mass spectrometer instrument setting used in generating DIA wildtype and LysoTAG HEK293 cells Lyso-IP, LysoTAG-HA-IP and Control-IP experiments | SFig2, SFig 7
timsTOF-SCP | LC and mass spectrometer instrument setting used in generating DIA wildtype HEK293 cells Lyso-IP and whole cell lysates experiments | Fig 1
Orbitrap Exploris480 | LC and mass spectrometer instrument setting used in generating DIA PBMCs single and six individuals Lyso-IP, mock-IP and whole cell lysate experiments | Fig 2
Orbitrap Astral | LC and mass spectrometer instrument setting used in generating DIA iPSC Lyso-IP and whole cell experiments | SFig 5

## timsTOF-Pro
DIA-MS instrument settings used in this study and variableDIA Isolation windows on timsTOF-Pro
MS settings: |  | LC System | EvoSep One
Scan Begin | 100 m/z | LC Column | EV1106 ENDURANCE COLUMN – 30 SPD
Scan End | 1700 m/z
Ion Polarity | Positive
Scan Mode | dia-PASEF
TIMS Settings:
Mode | Custom
1/K0 Start | 0.6
1/K0 End | 1.43
Ramp Time | 100 ms
Accu. Time | 100 ms
Duty Cycle | 100%
Ramp Rate | 9.34 Hz
MS Averaging | 1 x
Source Type | CaptiveSpray
Capillary | 1400 v 1528 nA
nanoBooster | 0.05 bar
Dry Gas | 3.0 l/min
Dry Temp | 180 oC
General
Deflection 1 Delta | 70.0 V
isCID Energy | 0.0 eV
Funnel 2 RF | 200.0 Vpp
Multipole RF | 200.0 Vpp
Collision cell
Collision Energy | 10.0 eV
Collision RF | 1500.0 Vpp
Quadrupole
Ion Energy | 5.0 eV
Low Mass | 200.0 m/z
Focus Pre TOF
Transfer Time | 60.0 µs
Pre Pulse Storage | 12.0 µs
Detection
High Sensitivity Detection - Low Sample Amount
TIMS
Δt1 (Deflection Transfer -> Capillary Exit) | minus 20.0 V
Δt1 (Deflection Transfer -> Deflection Discard) | minus 160.0 V
Δt1 (Funnel 1 In -> Deflection Transfer) | 110.0 V
Δt4 (Accumulation Trap -> Funnel 1 In) | 110.0 V
Δt5 (Accumulation Exit -> Accumulation Transfer) | 0.0 V
Δt6 (Ramp Start -> Accumulation Exit) | 55.0 V
Collision Cell In | 300.0 V
dia-PASEF Settings
Mass Range | 400.0 Da to 1201.0 Da
Mobility Range | 0.60 1/K0 to 1.43 1/K0
Cycle Time Estimate | 1.17s
Collision Energy Settings
1/K0 (V s/cm2) | Collision Energy (eV)
0.85 | 20
1.3 | 59
dia-PASEF Isolation list
#MS Type | Cycle Id | Start IM [1/K0] | End IM [1/K0] | Start Mass [m/z] | End Mass [m/z] | CE [eV]
MS1 | 0 | - | - | - | - | -
PASEF | 1 | 0.9 | 1.2 | 800 | 826 | -
PASEF | 1 | 0.6 | 0.9 | 400 | 426 | -
PASEF | 2 | 0.92 | 1.22 | 825 | 851 | -
PASEF | 2 | 0.62 | 0.92 | 425 | 451 | -
PASEF | 3 | 0.93 | 1.23 | 850 | 876 | -
PASEF | 3 | 0.63 | 0.93 | 450 | 476 | -
PASEF | 4 | 0.95 | 1.25 | 875 | 901 | -
PASEF | 4 | 0.65 | 0.95 | 475 | 501 | -
PASEF | 5 | 0.96 | 1.26 | 900 | 926 | -
PASEF | 5 | 0.66 | 0.96 | 500 | 526 | -
PASEF | 6 | 0.98 | 1.28 | 925 | 951 | -
PASEF | 6 | 0.68 | 0.98 | 525 | 551 | -
PASEF | 7 | 0.99 | 1.29 | 950 | 976 | -
PASEF | 7 | 0.69 | 0.99 | 550 | 576 | -
PASEF | 8 | 1.01 | 1.31 | 975 | 1001 | -
PASEF | 8 | 0.71 | 1.01 | 575 | 601 | -
PASEF | 9 | 1.02 | 1.32 | 1000 | 1026 | -
PASEF | 9 | 0.72 | 1.02 | 600 | 626 | -
PASEF | 10 | 1.04 | 1.34 | 1025 | 1051 | -
PASEF | 10 | 0.74 | 1.04 | 625 | 651 | -
PASEF | 11 | 1.06 | 1.36 | 1050 | 1076 | -
PASEF | 11 | 0.76 | 1.06 | 650 | 676 | -
PASEF | 12 | 1.07 | 1.37 | 1075 | 1101 | -
PASEF | 12 | 0.77 | 1.07 | 675 | 701 | -
PASEF | 13 | 1.09 | 1.39 | 1100 | 1126 | -
PASEF | 13 | 0.79 | 1.09 | 700 | 726 | -
PASEF | 14 | 1.1 | 1.4 | 1125 | 1151 | -
PASEF | 14 | 0.8 | 1.1 | 725 | 751 | -
PASEF | 15 | 1.12 | 1.42 | 1150 | 1176 | -
PASEF | 15 | 0.82 | 1.12 | 750 | 776 | -
PASEF | 16 | 1.13 | 1.43 | 1175 | 1201 | -
PASEF | 16 | 0.83 | 1.13 | 775 | 801 | -

## timsTOF-SCP
DIA-MS instrument settings used in this study and variableDIA Isolation windows on timsTOF-SCP
MS settings: |  | LC System | EvoSep One
Scan Begin | 100 m/z | LC Column | EV1106 ENDURANCE COLUMN – 30 SPD
Scan End | 1700 m/z
Ion Polarity | Positive
Scan Mode | dia-PASEF
TIMS Settings:
Mode | Custom
1/K0 Start | 0.64
1/K0 End | 1.4
Ramp Time | 100 ms
Accu. Time | 100 ms
Duty Cycle | 100%
Ramp Rate | 9.34 Hz
MS Averaging | 1 x
Source Type | CaptiveSpray
Capillary | 1500 v 1505 nA
nanoBooster | 0.20 bar
Dry Gas | 1.0 l/min
Dry Temp | 200 oC
General
Transfer
Deflection 0 Delta | 110.0 V
Funnel 0 RF | 250.0 Vpp
Multipole 0 RF | 100.0 Vpp
Deflection 1 Delta | 70.0 V
isCID Energy | 0.0 eV
Funnel 2 RF | 200.0 Vpp
Multipole RF | 500.0 Vpp
Collision cell
Collision Energy | 10.0 eV
Collision RF | 1500.0 Vpp
Quadrupole
Ion Energy | 5.0 eV
Low Mass | 200.0 m/z
Focus Pre TOF
Transfer Time | 60.0 µs
Pre Pulse Storage | 12.0 µs
Detection
High Sensitivity Detection - Low Sample Amount
TIMS
Δt1 (Deflection Transfer -> Capillary Exit) | minus 20.0 V
Δt1 (Deflection Transfer -> Deflection Discard) | minus 160.0 V
Δt1 (Funnel 1 In -> Deflection Transfer) | 110.0 V
Δt4 (Accumulation Trap -> Funnel 1 In) | 110.0 V
Δt5 (Accumulation Exit -> Accumulation Transfer) | 0.0 V
Δt6 (Ramp Start -> Accumulation Exit) | 55.0 V
Collision Cell In | 250.0 V
dia-PASEF Settings
Mass Range | 400.0 Da to 1000.0 Da
Mobility Range | 0.64 1/K0 to 1.37 1/K0
Cycle Time Estimate | 0.96 S
Collision Energy Settings
1/K0 (V s/cm2) | Collision Energy (eV)
0.6 | 20
1.6 | 59
dia-PASEF Isolation list
#MS Type | Cycle Id | Start IM [1/K0] | End IM [1/K0] | Start Mass [m/z] | End Mass [m/z] | CE [eV]
MS1 | 0 | - | - | - | - | -
PASEF | 1 | 1.01 | 1.37 | 800 | 825 | -
PASEF | 1 | 0.83 | 1.01 | 600 | 625 | -
PASEF | 1 | 0.64 | 0.83 | 400 | 425 | -
PASEF | 2 | 1.04 | 1.37 | 825 | 850 | -
PASEF | 2 | 0.85 | 1.04 | 625 | 650 | -
PASEF | 2 | 0.64 | 0.85 | 425 | 450 | -
PASEF | 3 | 1.06 | 1.37 | 850 | 875 | -
PASEF | 3 | 0.87 | 1.06 | 650 | 675 | -
PASEF | 3 | 0.64 | 0.87 | 450 | 475 | -
PASEF | 4 | 1.09 | 1.37 | 875 | 900 | -
PASEF | 4 | 0.9 | 1.09 | 675 | 700 | -
PASEF | 4 | 0.64 | 0.9 | 475 | 500 | -
PASEF | 5 | 1.11 | 1.37 | 900 | 925 | -
PASEF | 5 | 0.92 | 1.11 | 700 | 725 | -
PASEF | 5 | 0.64 | 0.92 | 500 | 525 | -
PASEF | 6 | 1.13 | 1.37 | 925 | 950 | -
PASEF | 6 | 0.94 | 1.13 | 725 | 750 | -
PASEF | 6 | 0.64 | 0.94 | 525 | 550 | -
PASEF | 7 | 1.16 | 1.37 | 950 | 975 | -
PASEF | 7 | 0.97 | 1.16 | 750 | 775 | -
PASEF | 7 | 0.64 | 0.97 | 550 | 575 | -
PASEF | 8 | 1.18 | 1.37 | 975 | 1000 | -
PASEF | 8 | 0.99 | 1.18 | 775 | 800 | -
PASEF | 8 | 0.64 | 0.99 | 575 | 600 | -

## Orbitrap Exploris 480
DIA-MS instrument settings used in this study and variableDIA Isolation windows on Orbitrap Exploris 480
Method duration | 145 min |  |  | {Initial Time} | Instrument Setup
MS Global settings: |  |  |  |  | Sampler.LowDispersionMode | Off
| Infusion mode: | Liquid Chromatrography | LC column: |  | Sampler.WashSpeed | 4.000 [µl/s]
| Expected LC peak width (s): | 20 | 2 cm pre-column: (C18, 5 μm, 100 Ao , 100 μ, 2 cm Nano-viper column # 164564, Thermo Scientific) |  | Sampler.WashVolume | 200.000 [µl]
| Advanced Peak determination: | TRUE | Analytical column: 50 cm analytical column (C18, 5 μm, 50 cm, 100 Ao Easy nano spray column # ES903, Thermo Scientific) |  | Sampler.PunctureDepth | 7.000 [mm]
| Default charge state: | 3 |  |  | Sampler.SampleHeight | 1.000 [mm]
| Internal mass calibration: | off | Note: If needed enable user defined calibratnt ion (Polysilaxolane 445.124650 or enable EasyIC option |  | Sampler.WasteSpeed | 4.000 [µl/s]
[5 empty] | Sampler.DispenseDelay | 2.000 [s]
Full scan settings: |  |  |  |  | Sampler.DispSpeed | 2.000 [µl/s]
| Orbitrap resolution: | 120000 |  |  | Sampler.DrawSpeed | 0.300 [µl/s]
| Scan range (m/z): | 375-1500 |  |  | Sampler.DrawDelay | 5.000 [s]
| RF lens(%): | 40 |  |  | Sampler.RinseBetweenReinjections | No
| AGC target: | Custom |  |  | Sampler.FlushVolume | 5.000 [µl]
| Normalized AGC target (%): | 300 |  |  | Sampler.TransVialPunctureDepth | 8.000 [mm]
| Maxiumum injection Time mode: | Custom |  |  | Sampler.TransLiquidHeight | 5.000 [mm]
| Maxiumum injection Time (ms): | 30 |  |  | Sampler.TransportVialCapacity | 99999
| Microscans: | 1 |  |  | Sampler.LastTransportVial | R1
| Data type: | Profile |  |  | Sampler.FirstTransportVial | R1
[5 empty] | Sampler.InjectMode | ulPickUp
tMS2 or DIA settings | Isolation offset: | Off |  |  | Sampler.LoopWashFactor | 2
| Collision Energy Mode: | Stepped |  |  | Sampler.PumpDevice | LoadingPump
| Collision Energy Type: | Normaized |  |  | Sampler.TempCtrl | On
| HCD Collision Energy (%): | 25, 28, 32 |  |  | Sampler.Temperature.Nominal | 7.0 [°C]
| Orbitrap reolution: | 30000 |  |  | Sampler.ReadyTempDelta | 3.0 [°C]
| Scan ragne mode: | Define m/z range |  |  | Sampler.Temperature.LowerLimit | 4.0 [°C]
| Scan Range (m/z): | 200 - 1200 | Note: Maximum of the matched fragement ions (b series and y-series) fall within this range and if needed this can be modified. |  | Sampler.Temperature.UpperLimit | 45.0 [°C]
| RF Lens (%): | 50 |  |  | PumpModule.LoadingPump.%A.Equate | %A
| AGC target: | Custom |  |  | PumpModule.LoadingPump.%B.Equate | %B
| Normalized AGC target (%): | 3000 | Note: It is recommended to fill the trap with a maximum accumulation of ions (3000% = 3E6 ions) for each of the DIA window to increase the sensitivity |  | PumpModule.LoadingPump.%C.Equate | %C
| Maxiumum injection Time mode: | Custom |  |  | PumpModule.LoadingPump.Pressure.LowerLimit | 0 [bar]
| Maxiumum injection Time (ms): | 70 |  |  | PumpModule.LoadingPump.Pressure.UpperLimit | 500 [bar]
| Micorscans: | 1 |  |  | PumpModule.LoadingPump.MaximumFlowRampUp | Infinite
| Data type: | Profile |  |  | PumpModule.LoadingPump.MaximumFlowRampDown | 10 [µl/min²]
| Polarity: | Positive |  |  | PumpModule.NC_Pump.%A.Equate | %A
| Loop control: | N |  |  | PumpModule.NC_Pump.%B.Equate | %B
| N (Number of Spectra): | 24 | We inlcude one full MS1 scan after every 24 DIA scans to accommodate maximum possible MS1 scans |  | PumpModule.NC_Pump.Pressure.LowerLimit | 0 [bar]
| Dynamic RT: | Off |  |  | PumpModule.NC_Pump.Pressure.UpperLimit | 900 [bar]
| Time Mode: | Unscheduled |  |  | PumpModule.NC_Pump.MaximumFlowRampUp | Infinite
[5 empty] | PumpModule.NC_Pump.MaximumFlowRampDown | Infinite
Scheme of vDIA windows mass list table: |  |  |  |  | ColumnOven.TempCtrl | On
| m/z | z | Isolation Window (mz) |  | ColumnOven.Temperature.Nominal | 45.0 [°C]
| 383.375 | 3 | 66.8 |  | ColumnOven.Temperature.LowerLimit | 20.0 [°C]
| 423 | 3 | 13.5 |  | ColumnOven.Temperature.UpperLimit | 75.0 [°C]
| 435 | 3 | 11.5 |  | ColumnOven.EquilibrationTime | 0.5 [min]
| 446.5 | 3 | 12.5 |  | ColumnOven.ReadyTempDelta | 4.0 [°C]
| 458 | 3 | 11.5 |  | ColumnOven.ValveLeft | 1_2
| 469 | 3 | 11.5 | 0 | Inject Preparation
| 480 | 3 | 11.5 |  | Wait | Sampler.Ready And PumpModule.LoadingPump.Ready And PumpModule.NC_Pump.Ready And ColumnOven.Ready
| 490.5 | 3 | 10.5 | 0 | Inject
| 501 | 3 | 11.5 |  | Sampler.Inject
| 512 | 3 | 11.5 | 0 | Start Run
| 523 | 3 | 11.5 |  | PumpModule.LoadingPump.LoadingPump_Pressure.AcqOn
| 533.5 | 3 | 10.5 |  | PumpModule.NC_Pump.NC_Pump_Pressure.AcqOn
| 544 | 3 | 11.5 | 0 | Run | Duration = 145.000 [min]
| 554.5 | 3 | 10.5 |  | PumpModule.LoadingPump.Flow.Nominal | 10.000 [µl/min]
| 565 | 3 | 11.5 |  | PumpModule.LoadingPump.%B.Value | 0.0 [%]
| 575.5 | 3 | 10.5 |  | PumpModule.LoadingPump.%C.Value | 0.0 [%]
| 586 | 3 | 11.5 |  | PumpModule.LoadingPump.Curve | 5
| 597.5 | 3 | 12.5 |  | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
| 609.5 | 3 | 12.5 |  | PumpModule.NC_Pump.%B.Value | 3.0 [%]
| 621.5 | 3 | 12.5 |  | PumpModule.NC_Pump.Curve | 5
| 633 | 3 | 11.5 | 3
| 645 | 3 | 13.5 |  | ColumnOven.ValveLeft | 6_1
| 657.5 | 3 | 12.5 | 7
| 670.5 | 3 | 14.5 |  | Sampler.MS_Start.On | Duration=1
| 684 | 3 | 13.5 | 12
| 697 | 3 | 13.5 |  | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
| 710.5 | 3 | 14.5 |  | PumpModule.NC_Pump.%B.Value | 7.0 [%]
| 725.5 | 3 | 16.5 |  | PumpModule.NC_Pump.Curve | 5
| 741 | 3 | 15.5 | 115
| 756.5 | 3 | 16.5 |  | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
| 773.5 | 3 | 18.5 |  | PumpModule.NC_Pump.%B.Value | 25.0 [%]
| 791 | 3 | 17.5 |  | PumpModule.NC_Pump.Curve | 5
| 808.5 | 3 | 18.5 | 129
| 827 | 3 | 19.5 |  | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
| 846.5 | 3 | 20.5 |  | PumpModule.NC_Pump.%B.Value | 37.0 [%]
| 866.5 | 3 | 20.5 |  | PumpModule.NC_Pump.Curve | 5
| 887.5 | 3 | 22.5 | 130
| 910.5 | 3 | 24.5 |  | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
| 935.5 | 3 | 26.5 |  | PumpModule.NC_Pump.%B.Value | 95.0 [%]
| 962.5 | 3 | 28.5 |  | PumpModule.NC_Pump.Curve | 5
| 992 | 3 | 31.5 | 135.3
| 1025 | 3 | 35.5 |  | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
| 1063 | 3 | 41.5 |  | PumpModule.NC_Pump.%B.Value | 95.0 [%]
| 1108.5 | 3 | 50.5 |  | PumpModule.NC_Pump.Curve | 5
| 1391.625 | 3 | 516.8 | 135.8
[5 empty] | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
[5 empty] | PumpModule.NC_Pump.%B.Value | 3.0 [%]
[5 empty] | PumpModule.NC_Pump.Curve | 5
140
[5 empty] | ColumnOven.ValveLeft | 1_2
145
[5 empty] | PumpModule.NC_Pump.Flow.Nominal | 0.250 [µl/min]
[5 empty] | PumpModule.NC_Pump.%B.Value | 3.0 [%]
[5 empty] | PumpModule.NC_Pump.Curve | 5
[5 empty] | PumpModule.LoadingPump.Flow.Nominal | 10.000 [µl/min]
[5 empty] | PumpModule.LoadingPump.%B.Value | 0.0 [%]
[5 empty] | PumpModule.LoadingPump.%C.Value | 0.0 [%]
[5 empty] | PumpModule.LoadingPump.Curve | 5
|  |  |  | 145 | Stop Run
PumpModule.LoadingPump.LoadingPump_Pressure.AcqOff
PumpModule.NC_Pump.NC_Pump_Pressure.AcqOff
End

## Orbitrap Astral
DIA-MS instrument settings used in this study and variableDIA Isolation windows on Orbitrap Astral
Global Settings |  | LC column: | ---- Overview ----
| Use Static Source Gasses | 2 cm pre-column: (C18, 5 μm, 100 Ao , 100 μ, 2 cm Nano-viper column # 164564, Thermo Scientific) | Name: New Instrument Method
| Use Ion Source Settings from Tune = Not Checked | Analytical column: Thermo, ES906 (150 um x 15 cm) Thermo Scientific) | Comment:
| Collision Gas pressure (mTorr)= 1 |  | Run time: 30.000 [min]
| Method Duration (min)= 30 |  | Instrument: Vanquish_Neo on mrc-astral
| Ion Source Type = NSI |  | Description:
| Spray Voltage = Static |  | ---- Script ----
| Spray Voltage: Positive Ion (V) = 1900 |  | initial Instrument Setup
| Spray Voltage: Negative Ion (V) = 600 |  | Neo.SamplerModule.Sampler.LoadingVolume: 3.000 [µl]
| Infusion Mode (LC)= False |  | Neo.SamplerModule.Sampler.FastLoadingWanted: Yes
| Ion Transfer Tube Temp (°C) = 290 |  | Neo.SamplerModule.Sampler.LoadingFlow: 60.000 [µl/min]
| APPI Lamp = Not in use |  | Neo.SamplerModule.Sampler.LoadingMode: FlowControl
| FAIMS Mode = Not Installed |  | Neo.SamplerModule.Sampler.TrapZebraWashCycles: 4
| Application Mode = Peptide |  | Neo.SamplerModule.Sampler.TrapZebraWashWanted: No
| Pressure Mode = Standard |  | Neo.PumpModule.Pump.FastColumnEquilibrationWanted: No
| Expected Peak Width (s) = 10 |  | Neo.PumpModule.Pump.ColumnEquilibrationFactor: 2.0
| Default Charge State = 2 |  | Neo.SamplerModule.Sampler.FastTrapWashWanted: No
| Advanced Peak Determination = True |  | Neo.SamplerModule.Sampler.TrapWashFactorStrong: Automatic
Experiment 1 |  |  | Neo.SamplerModule.Sampler.TrapWashFactorWeak: Automatic
| Experiment Name = MS |  | Neo.PumpModule.Pump.%B.Equate: "%B"
| Start Time (min) = 0 |  | Neo.PumpModule.Pump.%A.Equate: "%A"
| End Time (min) = 30 |  | Neo.PumpModule.Pump.%A_Solvent: H2O
Scan MasterScan | Detector Type = Orbitrap |  | Neo.PumpModule.Pump.%B_Solvent: ACN80
| Orbitrap Resolution = 240000 |  | Neo.SamplerModule.Sampler.SolventWeakWpName: "H2O 0.1FA"
| Scan Range (m/z) = 380-980 |  | Neo.SamplerModule.Sampler.SolventStrongWpName: "ACN 80 FA"
| Time (ms) = 5 |  | Neo.SamplerModule.Sampler.SolventWeakName: "H2O 0.1FA"
| Normalised AGC Target (%)= 500 |  | Neo.SamplerModule.Sampler.SolventStrongName: "ACN 80 FA"
| Microscans = 1 |  | Neo.SamplerModule.TempCtrl: On
| RF Lens(%) = 40 |  | Neo.SamplerModule.Temperature.Nominal: 7.0 [°C]
| Maximum Injection Time Mode = Custom |  | Neo.SamplerModule.Sampler.PunctureOffset: NoOffset
| AGC Target = Custom |  | Neo.SamplerModule.Sampler.VialBottomDetection: On
| EASY-IC = Off |  | Neo.SamplerModule.Sampler.DispenseSpeed: 5.000 [µl/s]
| DataType = Profile |  | Neo.SamplerModule.Sampler.DrawDelay: 2.0 [s]
| Polarity = Positive |  | Neo.SamplerModule.Sampler.DrawSpeed: 0.200 [µl/s]
| Source Fragmentation = False |  | Neo.SamplerModule.Sampler.AirGapEnable: Off
| Scan Description = |  | Neo.SamplerModule.Sampler.OuterNeedleWashCycleTime: Default
Neo.SamplerModule.Sampler.InjectWashWeakTime: 5.0 [s]
Neo.SamplerModule.Sampler.InjectWashWeakSpeed: 80.0 [µl/s]
Neo.SamplerModule.Sampler.InjectWashStrongSpeed: 80.0 [µl/s]
Neo.SamplerModule.Sampler.InjectWashStrongTime: 3.0 [s]
Experiment 2 |  |  | Neo.SamplerModule.Sampler.InjectWashMode: AfterDraw
| Experiment Name = DIA |  | Neo.SamplerModule.Sampler.TrapColumnFlushDirection: Backward
| Start Time (min) = 0 |  | Neo.ColumnCompartmentModule.ColumnChamber.TempCtrl: On
| End Time (min) = 30 |  | Neo.ColumnCompartmentModule.ColumnChamber.Temperature.Nominal: 30.00 [°C]
Scan DIAScan | Window Overlap (m/z) = 0 |  | Neo.SystemControllerModule.SeparationColumn1.Diameter: 150 [µm]
| Isolation Window = Custom |  | Neo.SystemControllerModule.SeparationColumn1.Length: 15.0 [cm]
| Isolation Window (m/z) = 4 |  | Neo.SystemControllerModule.SeparationColumn1.MaximumPressure: 1000 [bar]
| Multiplex Ions Enabled = False |  | Neo.SystemControllerModule.SeparationColumn1.MaximumFlow: 10.0 [µl/min]
| Window Placement Optimization = On |  | Neo.SystemControllerModule.SeparationColumn1.MaximumTemperature: 60.0 [°C]
| Number Of Scan Events = 149 |  | Neo.SystemControllerModule.SeparationColumn1.MaximumPressureChangeDown: 1000 [bar/min]
| Detector Type = Astral |  | Neo.SystemControllerModule.SeparationColumn1.MaximumPressureChangeUp: 1000 [bar/min]
| Collision Energy Type = Normalized |  | Neo.SystemControllerModule.TrapColumn1.Diameter: 300 [µm]
| HCD Collision Energies (%) = 25 |  | Neo.SystemControllerModule.TrapColumn1.Length: 0.5 [cm]
| Scan Range (m/z) = 150-2000 |  | Neo.SystemControllerModule.TrapColumn1.MaximumPressure: 1500 [bar]
| Precursor MassRange (m/z) = 380-980 |  | Neo.SystemControllerModule.TrapColumn1.MaximumFlow: 300.0 [µl/min]
| Time (ms) = 3 |  | Neo.SystemControllerModule.TrapColumn1.MaximumTemperature: 120.0 [°C]
| Normalised AGC Target (%)= 500 |  | Neo.SystemControllerModule.TrapColumn1.MaximumPressureChangeDown: 1000 [bar/min]
| Microscans = 1 |  | Neo.SystemControllerModule.TrapColumn1.MaximumPressureChangeUp: 1000 [bar/min]
| Maximum Injection Time Mode = Custom |  | Neo.SystemControllerModule.TrapColumn1.BackwardFlushSupported: Yes
| RF Lens(%) = 40 |  | Neo.SystemControllerModule.InjectionWorkflow: TrapAndElute
| AGC Target = Custom |  | Neo.SystemControllerModule.FlowRegime: NanoCap
| Scan Range Mode = Define m/z Range |  | Neo.SystemControllerModule.SystemSetup: VanquishNeo
| EASY-IC = Off |  | 0.000 [min] Equilibration
| DataType = Centroid |  | Neo.PumpModule.Pump.Flow.Nominal: 1.300 [µl/min]
| Polarity = Positive |  | Neo.PumpModule.Pump.%B.Value: 4.0 [%]
| Source Fragmentation = False |  | Neo.PumpModule.Pump.Curve: 5
| DIA Window Type = Auto |  | 0.000 [min] Inject Preparation
| DIA Window Mode = m/z Range |  | Wait Neo.SamplerModule.Sampler.Ready And Neo.PumpModule.Pump.Ready And Neo.ColumnCompartmentModule.Ready And Neo.SystemControllerModule.Ready
| Loop Control = Time |  | 0.000 [min] Inject
| Time (Sec) = 0.6 |  | Neo.SamplerModule.Sampler.Inject
| Time Mode = Start/End Time |  | 0.000 [min] Start Run
| Scan Description = |  | Neo.PumpModule.Pump.Pump_Pressure.AcqOn
| TMT = Off |  | Neo.SamplerModule.Sampler.Sampler_Pressure.AcqOn
>>>>>>>>>>>>> Mass List Table <<<<<<<<<<<<<< |  |  | 0.000 [min] Run
m/z¦ m/z Range¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 1.300 [µl/min]
382¦380.422805-384.424624¦ |  |  | Neo.PumpModule.Pump.%B.Value: 4.0 [%]
386¦384.424624-388.426443¦ |  |  | Neo.PumpModule.Pump.Curve: 5
390¦388.426443-392.428262¦ |  |  | 0.500 [min]
394¦392.428262-396.430081¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 1.300 [µl/min]
398¦396.430081-400.4319¦ |  |  | Neo.PumpModule.Pump.%B.Value: 4.0 [%]
402¦400.4319-404.433719¦ |  |  | Neo.PumpModule.Pump.Curve: 5
406¦404.433719-408.435538¦ |  |  | 0.600 [min]
410¦408.435538-412.437357¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 0.800 [µl/min]
414¦412.437357-416.439176¦ |  |  | Neo.PumpModule.Pump.%B.Value: 8.0 [%]
418¦416.439176-420.440995¦ |  |  | Neo.PumpModule.Pump.Curve: 5
422¦420.440995-424.442814¦ |  |  | 1.000 [min]
426¦424.442814-428.444633¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 0.800 [µl/min]
430¦428.444633-432.446452¦ |  |  | Neo.PumpModule.Pump.Curve: 5
434¦432.446452-436.448271¦ |  |  | 18.600 [min]
438¦436.448271-440.45009¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 0.800 [µl/min]
442¦440.45009-444.451909¦ |  |  | Neo.PumpModule.Pump.%B.Value: 22.5 [%]
446¦444.451909-448.453728¦ |  |  | Neo.PumpModule.Pump.Curve: 5
450¦448.453728-452.455547¦ |  |  | 27.700 [min]
454¦452.455547-456.457366¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 0.800 [µl/min]
458¦456.457366-460.459185¦ |  |  | Neo.PumpModule.Pump.%B.Value: 35.0 [%]
462¦460.459185-464.461004¦ |  |  | Neo.PumpModule.Pump.Curve: 5
466¦464.461004-468.462823¦ |  |  | 28.200 [min]
470¦468.462823-472.464642¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 2.000 [µl/min]
474¦472.464642-476.466461¦ |  |  | Neo.PumpModule.Pump.%B.Value: 55.0 [%]
478¦476.466461-480.46828¦ |  |  | Neo.PumpModule.Pump.Curve: 5
482¦480.46828-484.470099¦ |  |  | Neo.PumpModule.Pump.StartColumnWash
486¦484.470099-488.471918¦ |  |  | 28.800 [min]
490¦488.471918-492.473737¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 2.000 [µl/min]
494¦492.473737-496.475556¦ |  |  | Neo.PumpModule.Pump.%B.Value: 99.0 [%]
498¦496.475556-500.477375¦ |  |  | Neo.PumpModule.Pump.Curve: 5
502¦500.477375-504.479194¦ |  |  | 30.000 [min]
506¦504.479194-508.481013¦ |  |  | Neo.PumpModule.Pump.Flow.Nominal: 2.000 [µl/min]
510¦508.481013-512.482832¦ |  |  | Neo.PumpModule.Pump.%B.Value: 99.0 [%]
514¦512.482832-516.484651¦ |  |  | Neo.PumpModule.Pump.Curve: 5
518¦516.484651-520.48647¦ |  |  | 30.000 [min] Stop Run
522¦520.48647-524.488289¦ |  |  | Neo.PumpModule.Pump.Pump_Pressure.AcqOff
526¦524.488289-528.490108¦ |  |  | Neo.SamplerModule.Sampler.Sampler_Pressure.AcqOff
530¦528.490108-532.491927¦ |  |  | 30.000 [min] Post Run
534¦532.491927-536.493746¦ |  |  | Neo.PumpModule.Pump.StartColumnEquilibration
538¦536.493746-540.495565¦
542¦540.495565-544.497384¦
546¦544.497384-548.499203¦
550¦548.499203-552.501022¦
554¦552.501022-556.502841¦
558¦556.502841-560.50466¦
562¦560.50466-564.506479¦
566¦564.506479-568.508298¦
570¦568.508298-572.510117¦
574¦572.510117-576.511936¦
578¦576.511936-580.513755¦
582¦580.513755-584.515574¦
586¦584.515574-588.517393¦
590¦588.517393-592.519212¦
594¦592.519212-596.521031¦
598¦596.521031-600.52285¦
602¦600.52285-604.524669¦
606¦604.524669-608.526488¦
610¦608.526488-612.528307¦
614¦612.528307-616.530126¦
618¦616.530126-620.531945¦
622¦620.531945-624.533764¦
626¦624.533764-628.535583¦
630¦628.535583-632.537402¦
634¦632.537402-636.539221¦
638¦636.539221-640.54104¦
642¦640.54104-644.542859¦
646¦644.542859-648.544678¦
650¦648.544678-652.546497¦
654¦652.546497-656.548316¦
658¦656.548316-660.550135¦
662¦660.550135-664.551954¦
666¦664.551954-668.553773¦
670¦668.553773-672.555592¦
674¦672.555592-676.557411¦
678¦676.557411-680.55923¦
682¦680.55923-684.561049¦
686¦684.561049-688.562868¦
690¦688.562868-692.564687¦
694¦692.564687-696.566506¦
698¦696.566506-700.568325¦
702¦700.568325-704.570144¦
706¦704.570144-708.571963¦
710¦708.571963-712.573782¦
714¦712.573782-716.575601¦
718¦716.575601-720.57742¦
722¦720.57742-724.579239¦
726¦724.579239-728.581058¦
730¦728.581058-732.582877¦
734¦732.582877-736.584696¦
738¦736.584696-740.586515¦
742¦740.586515-744.588334¦
746¦744.588334-748.590153¦
750¦748.590153-752.591972¦
754¦752.591972-756.593791¦
758¦756.593791-760.59561¦
762¦760.59561-764.597429¦
766¦764.597429-768.599248¦
770¦768.599248-772.601067¦
774¦772.601067-776.602886¦
778¦776.602886-780.604705¦
782¦780.604705-784.606524¦
786¦784.606524-788.608343¦
790¦788.608343-792.610162¦
794¦792.610162-796.611981¦
798¦796.611981-800.6138¦
802¦800.6138-804.615619¦
806¦804.615619-808.617438¦
810¦808.617438-812.619257¦
814¦812.619257-816.621076¦
818¦816.621076-820.622895¦
822¦820.622895-824.624714¦
826¦824.624714-828.626533¦
830¦828.626533-832.628352¦
834¦832.628352-836.630171¦
838¦836.630171-840.63199¦
842¦840.63199-844.633809¦
846¦844.633809-848.635628¦
850¦848.635628-852.637447¦
854¦852.637447-856.639266¦
858¦856.639266-860.641085¦
862¦860.641085-864.642904¦
866¦864.642904-868.644723¦
870¦868.644723-872.646542¦
874¦872.646542-876.648361¦
878¦876.648361-880.65018¦
882¦880.65018-884.651999¦
886¦884.651999-888.653818¦
890¦888.653818-892.655637¦
894¦892.655637-896.657456¦
898¦896.657456-900.659275¦
902¦900.659275-904.661094¦
906¦904.661094-908.662913¦
910¦908.662913-912.664732¦
914¦912.664732-916.666551¦
918¦916.666551-920.66837¦
922¦920.66837-924.670189¦
926¦924.670189-928.672008¦
930¦928.672008-932.673827¦
934¦932.673827-936.675646¦
938¦936.675646-940.677465¦
942¦940.677465-944.679284¦
946¦944.679284-948.681103¦
950¦948.681103-952.682922¦
954¦952.682922-956.684741¦
958¦956.684741-960.68656¦
962¦960.68656-964.688379¦
966¦964.688379-968.690198¦
970¦968.690198-972.692017¦
974¦972.692017-976.693836¦
978¦976.693836-980.695655¦
>>>>>>>>>>>>> End Mass List Table <<<<<<<<<<<<<<